## Data
Product Number | HAC Compliant? | Air Interface | Bands | Form Factor | Standard Version | Max Power dBm | Max Avg. E Field (V/m) | Max Avg. E Field dB(V/m) | After Exclusion E field (V/m) | After Exclusion E field dB(V/m) | PMF Scaled E field (V/m) | PMF Scaled E field dB(V/m) | Rating E field | Max Avg. H Field (A/m) | Max Avg. H Field dB(A/m) | After Exclusion H field (A/m) | After Exclusion H field dB(A/m) | PMF Scaled H field (A/m) | PMF Scaled H field dB(A/m) | Rating H field | E H Ratings | Delta (dB) E-Field relative to M3 (pass) | Delta (dB) H-Field relative to M3 (pass) | OK E-Field Dominate | Amount H-Field is Dominate in dB | Corrected Delta (dB) E-Field relative to M3 (pass) | Corrected Delta (dB) H-Field relative to M3 (pass) | E/H ratio (Ohms)
1 | YES | iDEN | 800 | Candy Bar | 2007 | 28.06 | 69.97 | 36.8982 | 69.96 | 36.897 | 343.5 | 50.7185 | M3 | 0.1152 | -18.771 | 0.1048 | -19.5928 | 0.456 | -6.8207 | M4 | EM3 - HM4 | -14.1018 | -19.371 | OK |  | -0.28147 | -7.4207 | 753.2895 | Total Number of Phones in Sample = |  |  | 468
1 |  |  | 900 |  |  | 28.06 | 69.57 | 36.8484 | 69.57 | 36.8484 | 334 | 50.4749 | M3 | 0.109 | -19.2515 | 0.1018 | -19.845 | 0.445 | -7.0328 | M4 | EM3 - HM4 | -14.1516 | -19.8515 | OK |  | -0.52507 | -7.6328 | 750.5618 | Total Number of Phones Where H-field is dominate |  |  |  | 101
2 | YES | iDEN | 800 | Clam | 2007 | 28.6 | 30.44 | 29.6689 | 30.1 | 29.5713 | 147.8 | 43.3935 | M4 | 0.07575 | -22.4123 | 0.06276 | -24.0463 | 0.273 | -11.2767 | M4 | EM4 - HM4 | -21.3311 | -23.0123 | OK |  | -7.6065 | -11.8767 | 541.3919 |  |  | %of H-field Dominate Phones = |  | 21.5812
2 |  |  | 900 |  |  | 28.6 | 18.82 | 25.4924 | 17.375 | 24.7985 | 83.4 | 38.4233 | M4 | 0.06907 | -23.2142 | 0.05515 | -25.1691 | 0.241 | -12.3597 | M4 | EM4 - HM4 | -25.5076 | -23.8142 | Not OK | 1.6934 | -12.5767 | -12.9597 | 346.0581 |  | Average dB for Dominate H-Field = |  |  | 1.4761
3 | YES | iDEN | 800 | Clam | 2006 | 28.06 | 45.03 | 33.07 | 45.03 | 33.07 | 213.9 | 46.6042 | M3 | 0.08649 | -21.2607 | 0.06137 | -24.2409 | 0.286 | -10.8727 | M4 | EM3 - HM4 | -17.93 | -21.8607 | OK |  | -4.3958 | -11.4727 | 747.9021
3 |  |  | 900 |  |  | 28.06 | 30.42 | 29.6632 | 30.41 | 29.6603 | 146.9 | 43.3404 | M4 | 0.07638 | -22.3404 | 0.05668 | -24.9314 | 0.263 | -11.6009 | M4 | EM4 - HM4 | -21.3368 | -22.9404 | OK |  | -7.6596 | -12.2009 | 558.5551
4 | YES | CDMA | 850 | Clam | 2006 | 25 | 73.15 | 37.2843 | 73.15 | 37.2843 | 71 | 37.0252 | M3 | 0.1532 | -16.2948 | 0.1032 | -19.7264 | 0.096 | -20.3546 | M4 | EM3 - HM4 | -13.7157 | -16.8948 | OK |  | -13.9748 | -20.9546 | 739.5833
4 |  |  | 1900 |  |  | 25 | 61.07 | 35.7166 | 60.68 | 35.6609 | 62.5 | 35.9176 | M4 | 0.1514 | -16.3975 | 0.1407 | -17.0341 | 0.128 | -17.8558 | M4 | EM4 - HM4 | -5.2834 | -6.9975 | OK |  | -5.0824 | -8.4558 | 488.2812
5 | YES | CDMA | 850 | Clam | 2006 | 25 | 76.86 | 37.714 | 76.87 | 37.7151 | 76.1 | 37.6277 | M4 | 0.1479 | -16.6006 | 0.1051 | -19.5679 | 0.103 | -19.7433 | M4 | EM4 - HM4 | -13.286 | -17.2006 | OK |  | -13.3723 | -20.3433 | 738.835
5 |  |  | 1900 |  |  | 25 | 43.03 | 32.6754 | 43 | 32.6694 | 43 | 32.6694 | M4 | 0.1028 | -19.7601 | 0.1021 | -19.8195 | 0.098 | -20.1755 | M4 | EM4 - HM4 | -8.3246 | -10.3601 | OK |  | -8.3306 | -10.7755 | 438.7755
6 | YES | CDMA | 850 | Clam | 2006 | 25 | 113.5 | 41.0999 | 113.5 | 41.0999 | 126 | 42.0074 | M4 | 0.2116 | -13.4897 | 0.1467 | -16.6714 | 0.157 | -16.082 | M4 | EM4 - HM4 | -9.9001 | -14.0897 | OK |  | -8.9926 | -16.682 | 802.5478
6 |  |  | 1900 |  |  | 25 | 47.53 | 33.5394 | 43.1 | 32.6895 | 48.7 | 33.7506 | M4 | 0.1326 | -17.5491 | 0.1228 | -18.216 | 0.124 | -18.1316 | M4 | EM4 - HM4 | -7.4606 | -8.1491 | OK |  | -7.2494 | -8.7316 | 392.7419
7 | YES | CDMA | 850 | Candy Bar | 2006 | 25 | 124.5 | 41.9034 | 124.5 | 41.9034 | 123.3 | 41.8193 | M4 | 0.2126 | -13.4487 | 0.15 | -16.4782 | 0.147 | -16.6537 | M4 | EM4 - HM4 | -9.0966 | -14.0487 | OK |  | -9.1807 | -17.2537 | 838.7755
7 |  |  | 1900 |  |  | 25 | 57.35 | 35.1707 | 51.2 | 34.1854 | 51.2 | 34.1854 | M4 | 0.1586 | -15.9939 | 0.1385 | -17.171 | 0.133 | -17.523 | M4 | EM4 - HM4 | -5.8293 | -6.5939 | OK |  | -6.8146 | -8.123 | 384.9624
8 | YES | CDMA | 850 | Clam | 2006 | 25 | 65.42 | 36.3142 | 65.38 | 36.3089 | 69.3 | 36.8147 | M4 | 0.1569 | -16.0875 | 0.1152 | -18.771 | 0.121 | -18.3443 | M4 | EM4 - HM4 | -14.6858 | -16.6875 | OK |  | -14.1853 | -18.9443 | 572.7273
8 |  |  | 1900 |  |  | 25 | 46.75 | 33.3956 | 45.67 | 33.1926 | 47.5 | 33.5339 | M4 | 0.1148 | -18.8012 | 0.1089 | -19.2594 | 0.11 | -19.1721 | M4 | EM4 - HM4 | -7.6044 | -9.4012 | OK |  | -7.4661 | -9.7721 | 431.8182
9 | YES | CDMA | 850 | Clam | 2006 | [20 empty] | 0 | 0
9 |  |  | 1900 |  |  | 24.7 | 51.05 | 34.1599 | 50.62 | 34.0864 | 57.2 | 35.1479 | M4 | 0.1332 | -17.5099 | 0.1208 | -18.3587 | 0.122 | -18.2728 | M4 | EM4 - HM4 |  |  |  |  | 0.98801 | -0.76289 | 468.8525
10 | YES | CDMA | 850 | Clam | 2006 | 25 | 95 | 39.5545 | 95 | 39.5545 | 105.4 | 40.4568 | M4 | 0.1991 | -14.0186 | 0.143 | -16.8933 | 0.153 | -16.3062 | M4 | EM4 - HM4 | -11.4455 | -14.6186 | OK |  | -10.5432 | -16.9062 | 688.8889
10 |  |  | 1900 |  |  | 25 | 51.7 | 34.2698 | 51.7 | 34.2698 | 58.4 | 35.3283 | M4 | 0.1941 | -14.2395 | 0.1366 | -17.291 | 0.138 | -17.2024 | M4 | EM4 - HM4 | -6.7302 | -4.8395 | Not OK | 1.8907 | -5.6717 | -7.8024 | 423.1884
11 | YES | CDMA | 850 | Clam | 2006 | 25 | 105.3 | 40.4486 | 105.3 | 40.4486 | 111.6 | 40.9533 | M4 | 0.3516 | -9.079 | 0.2667 | -11.4795 | 0.28 | -11.0568 | M4 | EM4 - HM4 | -10.5514 | -9.679 | Not OK | 0.87241 | -10.0467 | -11.6568 | 398.5714
11 |  |  | 1900 |  |  | 25 | 44.96 | 33.0565 | 39.6 | 31.9539 | 41.2 | 32.2979 | M4 | 0.1273 | -17.9034 | 0.1109 | -19.1014 | 0.112 | -19.0156 | M4 | EM4 - HM4 | -7.9435 | -8.5034 | OK |  | -8.7021 | -9.6156 | 367.8571
12 | YES | CDMA | 850 | Clam | 2006 | 25 | 87.18 | 38.8083 | 87.17 | 38.8073 | 92.4 | 39.3134 | M4 | 0.1586 | -15.9939 | 0.1057 | -19.5185 | 0.111 | -19.0935 | M4 | EM4 - HM4 | -12.1917 | -16.5939 | OK |  | -11.6866 | -19.6935 | 832.4324
12 |  |  | 1900 |  |  | 25 | 54.58 | 34.7407 | 53.27 | 34.5297 | 55.4 | 34.8702 | M4 | 0.1579 | -16.0324 | 0.1495 | -16.5072 | 0.151 | -16.4205 | M4 | EM4 - HM4 | -6.2593 | -6.6324 | OK |  | -6.1298 | -7.0205 | 366.8874
13 | YES | CDMA | 850 | Clam | 2006 | [20 empty] | 0 | 0
13 |  |  | 1900 |  |  | 24.7 | 48.06 | 33.6357 | 47.98 | 33.6212 | 52.3 | 34.37 | M4 | 0.1099 | -19.18 | 0.1066 | -19.4449 | 0.097 | -20.2646 | M4 | EM4 - HM4 |  |  |  |  | 0.73436 | -1.0845 | 539.1753
14 | YES | CDMA | 850 | QWERTY Candy Bar | 2006 | 26 | 102.2 | 40.189 | 101.8 | 40.155 | 98.7 | 39.8863 | M4 | 0.1978 | -14.0755 | 0.1935 | -14.2664 | 0.18 | -14.8945 | M4 | EM4 - HM4 | -10.811 | -14.6755 | OK |  | -11.1137 | -15.4945 | 548.3333
14 |  |  | 1900 |  |  | 25 | 71.75 | 37.1164 | 65.63 | 36.342 | 67.6 | 36.5989 | M3 | 0.2121 | -13.4692 | 0.2121 | -13.4692 | 0.193 | -14.2889 | M3 | EM3 - HM3 | -3.8836 | -4.0692 | OK |  | -4.4011 | -4.8889 | 350.2591
15 | YES | CDMA | 850 | Clam | 2006 | 25 | 100.9 | 40.0778 | 100.9 | 40.0778 | 103.9 | 40.3323 | M4 | 0.19 | -14.4249 | 0.1288 | -17.8017 | 0.134 | -17.4579 | M4 | EM4 - HM4 | -10.9222 | -15.0249 | OK |  | -10.6677 | -18.0579 | 775.3731
15 |  |  | 1900 |  |  | 25 | 63.44 | 36.0473 | 61.28 | 35.7464 | 66.8 | 36.4955 | M3 | 0.1696 | -15.4115 | 0.1407 | -17.0341 | 0.128 | -17.8558 | M4 | EM3 - HM4 | -4.9527 | -6.0115 | OK |  | -4.5045 | -8.4558 | 521.875
16 | YES | CDMA | 850 | Slider | 2006 | 25 | 153.6 | 43.7278 | 145.1 | 43.2333 | 158.2 | 43.9841 | M4 | 0.2335 | -12.6343 | 0.1683 | -15.4783 | 0.175 | -15.1392 | M4 | EM4 - HM4 | -7.2722 | -13.2343 | OK |  | -7.0159 | -15.7392 | 904
16 |  |  | 1900 |  |  | 25 | 63.9 | 36.11 | 55.87 | 34.9436 | 60.9 | 35.6923 | M4 | 0.1565 | -16.1097 | 0.1099 | -19.18 | 0.1 | -20 | M4 | EM4 - HM4 | -4.89 | -6.7097 | OK |  | -5.3077 | -10.6 | 609
17 | YES | CDMA | 850 | Slider | 2006 | 25 | 170.5 | 44.6345 | 170.5 | 44.6345 | 179 | 45.0571 | M4 | 0.2577 | -11.7777 | 0.1816 | -14.8177 | 0.187 | -14.5632 | M4 | EM4 - HM4 | -6.3655 | -12.3777 | OK |  | -5.9429 | -15.1632 | 957.2193
17 |  |  | 1900 |  |  | 25 | 71.89 | 37.1334 | 62.08 | 35.859 | 65.8 | 36.3645 | M3 | 0.1731 | -15.2341 | 0.1243 | -18.1106 | 0.128 | -17.8558 | M4 | EM3 - HM4 | -3.8666 | -5.8341 | OK |  | -4.6355 | -8.4558 | 514.0625
18 | YES | CDMA | 850 | Clam | 2006 | 25 | 94.47 | 39.5059 | 94.47 | 39.5059 | 99.2 | 39.9302 | M4 | 0.1857 | -14.6238 | 0.1137 | -18.8848 | 0.121 | -18.3443 | M4 | EM4 - HM4 | -11.4941 | -15.2238 | OK |  | -11.0698 | -18.9443 | 819.8347
18 |  |  | 1900 |  |  | 25 | 51.45 | 34.2277 | 48.02 | 33.6284 | 48.5 | 33.7148 | M4 | 0.1509 | -16.4262 | 0.1067 | -19.4367 | 0.095 | -20.4455 | M4 | EM4 - HM4 | -6.7723 | -7.0262 | OK |  | -7.2852 | -11.0455 | 510.5263
19 | YES | CDMA | 850 | QWERTY Candy Bar | 2006 | 26 | 114.4 | 41.1685 | 114.4 | 41.1685 | 117.8 | 41.4229 | M4 | 0.1907 | -14.393 | 0.1366 | -17.291 | 0.138 | -17.2024 | M4 | EM4 - HM4 | -9.8315 | -14.993 | OK |  | -9.5771 | -17.8024 | 853.6232
19 |  |  | 1900 |  |  | 25 | 92.34 | 39.3078 | 76.85 | 37.7129 | 83 | 38.3816 | M3 | 0.2821 | -10.9919 | 0.2817 | -11.0043 | 0.293 | -10.6626 | M3 | EM3 - HM3 | -1.6922 | -1.5919 | Not OK | 0.10026 | -2.6184 | -1.2626 | 283.2765
20 | YES | CDMA | 850 | Slider | 2006 | 25 | 154.6 | 43.7842 | 154.6 | 43.7842 | 162.3 | 44.2064 | M4 | 0.2427 | -12.2986 | 0.168 | -15.4938 | 0.173 | -15.2391 | M4 | EM4 - HM4 | -7.2158 | -12.8986 | OK |  | -6.7936 | -15.8391 | 938.1503
20 |  |  | 1900 |  |  | 25 | 71.46 | 37.0813 | 63.3 | 36.0281 | 67.1 | 36.5345 | M3 | 0.1855 | -14.6331 | 0.1398 | -17.0899 | 0.144 | -16.8328 | M4 | EM3 - HM4 | -3.9187 | -5.2331 | OK |  | -4.4655 | -7.4328 | 465.9722
21 | YES | CDMA | 850 | Clam | 2006 | 25 | 106.9 | 40.5796 | 106.9 | 40.5796 | 112.2 | 40.9999 | M4 | 0.2077 | -13.6513 | 0.1447 | -16.7906 | 0.149 | -16.5363 | M4 | EM4 - HM4 | -10.4204 | -14.2513 | OK |  | -10.0001 | -17.1363 | 753.0201
21 |  |  | 1900 |  |  | 25 | 48.23 | 33.6663 | 48.21 | 33.6627 | 51.1 | 34.1684 | M4 | 0.158 | -16.0269 | 0.133 | -17.523 | 0.137 | -17.2656 | M4 | EM4 - HM4 | -7.3337 | -6.6269 | Not OK | 0.7068 | -6.8316 | -7.8656 | 372.9927
22 | YES | CDMA | 850 | Clam | 2006 | 25 | 77.02 | 37.7321 | 77.02 | 37.7321 | 80.9 | 38.159 | M4 | 0.1442 | -16.8207 | 0.101 | -19.9136 | 0.104 | -19.6593 | M4 | EM4 - HM4 | -13.2679 | -17.4207 | OK |  | -12.841 | -20.2593 | 777.8846
22 |  |  | 1900 |  |  | 25 | 46.43 | 33.336 | 46.13 | 33.2797 | 48.9 | 33.7862 | M4 | 0.1386 | -17.1647 | 0.133 | -17.523 | 0.137 | -17.2656 | M4 | EM4 - HM4 | -7.664 | -7.7647 | OK |  | -7.2138 | -7.8656 | 356.9343
23 | YES | CDMA | 850 | Clam | 2006 | 25 | 97.66 | 39.7943 | 94.61 | 39.5187 | 96.5 | 39.6905 | M4 | 0.171 | -15.3401 | 0.1155 | -18.7484 | 0.119 | -18.4891 | M4 | EM4 - HM4 | -11.2057 | -15.9401 | OK |  | -11.3095 | -19.0891 | 810.9244
23 |  |  | 1900 |  |  | 25 | 53.48 | 34.5638 | 52.04 | 34.3267 | 56.2 | 34.9947 | M4 | 0.1635 | -15.7296 | 0.1447 | -16.7906 | 0.149 | -16.5363 | M4 | EM4 - HM4 | -6.4362 | -6.3296 | Not OK | 0.10653 | -6.0053 | -7.1363 | 377.1812
24 | YES | CDMA | 850 | Clam | 2006 | 25 | 101.2 | 40.1036 | 101.2 | 40.1036 | 103.3 | 40.282 | M4 | 0.1816 | -14.8177 | 0.1214 | -18.3156 | 0.125 | -18.0618 | M4 | EM4 - HM4 | -10.8964 | -15.4177 | OK |  | -10.718 | -18.6618 | 826.4
24 |  |  | 1900 |  |  | 25 | 46.02 | 33.2589 | 46.02 | 33.2589 | 49.7 | 33.9271 | M4 | 0.1363 | -17.3101 | 0.1049 | -19.5845 | 0.108 | -19.3315 | M4 | EM4 - HM4 | -7.7411 | -7.9101 | OK |  | -7.0729 | -9.9315 | 460.1852
25 | YES | CDMA | 850 | Clam | 2006 | 25 | 109.7 | 40.8041 | 107.8 | 40.6524 | 110 | 40.8279 | M4 | 0.2228 | -13.0417 | 0.1631 | -15.7509 | 0.168 | -15.4938 | M4 | EM4 - HM4 | -10.1959 | -13.6417 | OK |  | -10.1721 | -16.0938 | 654.7619
25 |  |  | 1900 |  |  | 25 | 40.98 | 32.2514 | 38.61 | 31.734 | 41.7 | 32.4027 | M4 | 0.1135 | -18.9001 | 0.1039 | -19.6677 | 0.107 | -19.4123 | M4 | EM4 - HM4 | -8.7486 | -9.5001 | OK |  | -8.5973 | -10.0123 | 389.7196
26 | YES | CDMA | 850 | Clam | 2006 | 25 | 93.1 | 39.379 | 93.1 | 39.379 | 95.9 | 39.6364 | M4 | 0.1883 | -14.503 | 0.1204 | -18.3875 | 0.124 | -18.1316 | M4 | EM4 - HM4 | -11.621 | -15.103 | OK |  | -11.3636 | -18.7316 | 773.3871
26 |  |  | 1900 |  |  | 25 | 34.85 | 30.8441 | 34.3 | 30.7059 | 36.7 | 31.2933 | M4 | 0.1372 | -17.2529 | 0.098 | -20.1755 | 0.099 | -20.0873 | M4 | EM4 - HM4 | -10.1559 | -7.8529 | Not OK | 2.303 | -9.7067 | -10.6873 | 370.7071
[21 empty] | E - H |  |  |  |  | 0 | 0
27 | YES | CDMA | 850 | Clam | 2006 | 25 | 100.8 | 40.0692 | 100.6 | 40.052 | 103.6 | 40.3072 | M4 | 0.2249 | -12.9602 | 0.1602 | -15.9067 | 0.165 | -15.6503 | M4 | EM4 - HM4 | -10.9308 | -13.5602 | OK |  | -10.6928 | -16.2503 | 627.8788
27 |  |  | 1700 |  |  | 25 | 69 | 36.777 | 68.68 | 36.7366 | 72.8 | 37.2426 | M3 | 0.1972 | -14.1019 | 0.1743 | -15.1741 | 0.176 | -15.0897 | M4 | EM3 - HM4 | -4.223 | -4.7019 | OK |  | -3.7574 | -5.6897 | 413.6364
[6 empty] | 25 | 55.34 | 34.8608 | 53.36 | 34.5443 | 57.1 | 35.1327 | M4 | 0.1699 | -15.3961 | 0.1257 | -18.0133 | 0.127 | -17.9239 | M4 | EM4 - HM4 | -16.1392 | -15.9961 | Not OK | 0.14308 | -15.8673 | -18.5239 | 449.6063
28 | YES | CDMA | 850 | Clam | 2007 | 24.5 | 88.2 | 38.9094 | 88.16 | 38.9054 | 90.8 | 39.1617 | M4 | 0.1766 | -15.0602 | 0.1243 | -18.1106 | 0.128 | -17.8558 | M4 | EM4 - HM4 | -12.0906 | -15.6602 | OK |  | -11.8383 | -18.4558 | 709.375
28 |  |  | 1900 |  |  | 24.5 | 60.12 | 35.5804 | 60.09 | 35.576 | 64.3 | 36.1642 | M3 | 0.1729 | -15.2441 | 0.1525 | -16.3346 | 0.154 | -16.2496 | M4 | EM3 - HM4 | -5.4196 | -5.8441 | OK |  | -4.8358 | -6.8496 | 417.5325
[21 empty] | E - H |  |  |  |  | 0 | 0
29 | YES | CDMA | 850 | Slider | 2007 | 25 | 129.3 | 42.232 | 128.8 | 42.1983 | 132.7 | 42.4574 | M4 | 0.2289 | -12.8071 | 0.1718 | -15.2995 | 0.177 | -15.0405 | M4 | EM4 - HM4 | -8.768 | -13.4071 | OK |  | -8.5426 | -15.6405 | 749.7175
29 | [20 empty] | E - H |  |  |  |  | 0 | 0
29 |  |  | 1900 |  |  | 25 | 52.4 | 34.3866 | 47.29 | 33.4954 | 50.6 | 34.083 | M4 | 0.1361 | -17.3228 | 0.1297 | -17.7412 | 0.131 | -17.6546 | M4 | EM4 - HM4 | -6.6134 | -7.9228 | OK |  | -6.917 | -8.2546 | 386.2595
29 |  |  | 1700 |  |  | 25 | 64.47 | 36.1872 | 58.5 | 35.3431 | 62.6 | 35.9315 | M4 | 0.1448 | -16.7846 | 0.1298 | -17.7345 | 0.135 | -17.3933 | M4 | EM4 - HM4 | -4.8128 | -7.3846 | OK |  | -5.0685 | -7.9933 | 463.7037
30 | YES | CDMA | 850 | Slider | 2007 | 25 | 107.2 | 40.6039 | 107.2 | 40.6039 | 110.4 | 40.8594 | M4 | 0.1821 | -14.7938 | 0.1369 | -17.2719 | 0.141 | -17.0156 | M4 | EM4 - HM4 | -10.3961 | -15.3938 | OK |  | -10.1406 | -17.6156 | 782.9787
30 |  |  | 1700 |  |  | 25 | 64.47 | 36.1872 | 58.5 | 35.3431 | 62.6 | 35.9315 | M4 | 0.1448 | -16.7846 | 0.1298 | -17.7345 | 0.135 | -17.3933 | M4 | EM4 - HM4 | -4.8128 | -7.3846 | OK |  | -5.0685 | -7.9933 | 463.7037
30 |  |  | 1900 |  |  | 25 | 52.4 | 34.3866 | 47.29 | 33.4954 | 50.6 | 34.083 | M4 | 0.1361 | -17.3228 | 0.1297 | -17.7412 | 0.131 | -17.6546 | M4 | EM4 - HM4 | -6.6134 | -7.9228 | OK |  | -6.917 | -8.2546 | 386.2595
[21 empty] | E - H |  |  |  |  | 0 | 0
31 | YES | CDMA | 850 | Candy Bar | 2007 | 25 | 128.6 | 42.1848 | 128.58 | 42.1835 | 136.3 | 42.6899 | M4 | 0.2472 | -12.139 | 0.1828 | -14.7605 | 0.181 | -14.8464 | M4 | EM4 - HM4 | -8.8152 | -12.739 | OK |  | -8.3101 | -15.4464 | 753.0387
31 |  |  | 1700 |  |  | 25 | 89.15 | 39.0024 | 84.95 | 38.5833 | 90.9 | 39.1713 | M3 | 0.2023 | -13.8801 | 0.1894 | -14.4524 | 0.197 | -14.1107 | M3 | EM3 - HM3 | -1.9976 | -4.4801 | OK |  | -1.8287 | -4.7107 | 461.4213
31 |  |  | 1900 |  |  | 25 | 78.84 | 37.9349 | 75.43 | 37.5509 | 79.2 | 37.9745 | M3 | 0.2008 | -13.9447 | 0.1867 | -14.5771 | 0.196 | -14.1549 | M3 | EM3 - HM3 | -3.0651 | -4.5447 | OK |  | -3.0255 | -4.7549 | 404.0816
[21 empty] | E - H |  |  |  |  | 0 | 0
32 | YES | GSM | 850 | Clam | 2006 | 32.5 | 75.66 | 37.5773 | 75.66 | 37.5773 | 214.9 | 46.6447 | M3 | 0.1672 | -15.5353 | 0.1181 | -18.555 | 0.313 | -10.0891 | M4 | EM3 - HM4 | -10.9227 | -13.6353 | OK |  | -1.8553 | -8.1891 | 686.5815
32 |  | GSM | 1900 |  |  | 29.5 | 24.71 | 27.8575 | 23.86 | 27.5534 | 68 | 36.6502 | M3 | 0.05911 | -24.5668 | 0.0574 | -24.8218 | 0.123 | -18.2019 | M4 | EM3 - HM4 | -10.6425 | -12.6668 | OK |  | -1.8498 | -6.3019 | 552.8455
32 |  | WCDMA | 850 |  |  | 24 | 87.21 | 38.8113 | 87.17 | 38.8073 | 80.2 | 38.0835 | M4 | 0.1766 | -15.0602 | 0.1258 | -18.0064 | 0.117 | -18.6363 | M4 | EM4 - HM4 | -12.1887 | -15.6602 | OK |  | -12.9165 | -19.2363 | 685.4701
32 |  | WCDMA | 1900 |  |  | 24 | 36.84 | 31.3264 | 36.2 | 31.1742 | 33.3 | 30.4489 | M4 | 0.08865 | -21.0464 | 0.0875 | -21.1598 | 0.077 | -22.2702 | M4 | EM4 - HM4 | -9.6736 | -11.6464 | OK |  | -10.5511 | -12.8702 | 432.4675
33 | YES | GSM | 850 | Clam | 2006 | 33 | 69.69 | 36.8634 | 69.69 | 36.8634 | 199.3 | 45.9901 | M3 | 0.143 | -16.8933 | 0.0955 | -20.3999 | 0.231 | -12.7278 | M4 | EM3 - HM4 | -11.6366 | -14.9933 | OK |  | -2.5099 | -10.8278 | 862.7706
33 |  | GSM | 1900 |  |  | 30 | 23.41 | 27.388 | 19.83 | 25.9465 | 57.7 | 35.2235 | M3 | 0.08308 | -21.6101 | 0.0623 | -24.1102 | 0.165 | -15.6503 | M3 | EM3 - HM3 | -11.112 | -9.7101 | Not OK | 1.4019 | -3.2765 | -3.7503 | 349.697
33 |  | WCDMA | 850 |  |  | 24 | 75.71 | 37.5831 | 75.67 | 37.5785 | 68.1 | 36.6629 | M4 | 0.1465 | -16.6832 | 0.09778 | -20.195 | 0.088 | -21.1103 | M4 | EM4 - HM4 | -13.4169 | -17.2832 | OK |  | -14.3371 | -21.7103 | 773.8636
33 |  | WCDMA | 1900 |  |  | 24 | 38.19 | 31.639 | 32.31 | 30.1867 | 29.4 | 29.3669 | M4 | 0.1234 | -18.1737 | 0.0933 | -20.6024 | 0.084 | -21.5144 | M4 | EM4 - HM4 | -9.361 | -8.7737 | Not OK | 0.58731 | -11.6331 | -12.1144 | 350
34 | YES | GSM | 850 | Clam | 2007 | 33.3 | 55.64 | 34.9077 | 55.54 | 34.8921 | 158.3 | 43.9896 | M3 | 0.1153 | -18.7634 | 0.0846 | -21.4526 | 0.209 | -13.5971 | M4 | EM3 - HM4 | -13.5923 | -16.8634 | OK | FALSE | -4.5104 | -11.6971 | 757.4163
34 |  | GSM | 1900 |  |  | 30.25 | 27.48 | 28.7803 | 26.69 | 28.527 | 76.6 | 37.6846 | M3 | 0.09378 | -20.5578 | 0.0769 | -22.2815 | 0.203 | -13.8501 | M3 | EM3 - HM3 | -9.7197 | -8.6578 | Not OK | 1.0619 | -0.81542 | -1.9501 | 377.3399
34 |  | WCDMA | 850 |  |  | 24 | 50.55 | 34.0744 | 50.54 | 34.0727 | 46.5 | 33.3491 | M4 | 0.1014 | -19.8792 | 0.0758 | -22.4066 | 0.069 | -23.223 | M4 | EM4 - HM4 | -16.9256 | -20.4792 | OK | FALSE | -17.6509 | -23.823 | 673.913
34 |  | WCDMA | 1900 |  |  | 24 | 45.63 | 33.185 | 43.37 | 32.7438 | 39.9 | 32.0195 | M4 | 0.1494 | -16.513 | 0.1179 | -18.5697 | 0.105 | -19.5762 | M4 | EM4 - HM4 | -7.815 | -7.113 | Not OK | 0.702 | -8.9805 | -10.1762 | 380
35 | YES | GSM | 850 | Slider | 2007 | 32.5 | 85.86 | 38.6758 | 85.86 | 38.6758 | 244.7 | 47.7727 | M3 | 0.1612 | -15.8527 | 0.1121 | -19.0079 | 0.277 | -11.1504 | M4 | EM3 - HM4 | -9.8242 | -13.9527 | OK | FALSE | -0.72732 | -9.2504 | 883.3935
35 |  | GSM | 1900 |  |  | 30.25 | 27.48 | 28.7803 | 26.69 | 28.527 | 76.6 | 37.6846 | M3 | 0.09378 | -20.5578 | 0.0769 | -22.2815 | 0.203 | -13.8501 | M3 | EM3 - HM3 | -9.7197 | -8.6578 | Not OK | 1.0619 | -0.81542 | -1.9501 | 377.3399
35 |  | WCDMA | 850 |  |  | 24 | 72.78 | 37.2402 | 72.78 | 37.2402 | 66.95 | 36.515 | M4 | 0.1363 | -17.3101 | 0.1014 | -19.8792 | 0.0923 | -20.696 | M4 | EM4 - HM4 | -13.7598 | -17.9101 | OK | FALSE | -14.485 | -21.296 | 725.3521
35 |  | WCDMA | 1900 |  |  | 24 | 45.63 | 33.185 | 43.37 | 32.7438 | 39.9 | 32.0195 | M4 | 0.1494 | -16.513 | 0.1179 | -18.5697 | 0.105 | -19.5762 | M4 | EM4 - HM4 | -7.815 | -7.113 | Not OK | 0.702 | -8.9805 | -10.1762 | 380
36 | YES | GSM | 850 | Clam | 2006 | 33 | 89.7 | 39.0558 | 89.69 | 39.0549 | 257.4 | 48.2122 | M3 | 0.2326 | -12.6678 | 0.1595 | -15.9448 | 0.426 | -7.4118 | M4 | EM3 - HM4 | -9.4442 | -10.7678 | OK | FALSE | -0.28783 | -5.5118 | 604.2254
36 |  |  | 1900 |  |  | 30 | 33.04 | 30.3808 | 27.17 | 28.6818 | 81.5 | 38.2232 | M3 | 0.1013 | -19.8878 | 0.0785 | -22.1026 | 0.179 | -14.9429 | M3 | EM3 - HM3 | -8.1192 | -7.9878 | Not OK | 0.13139 | -0.27685 | -3.0429 | 455.3073
37 | YES | GSM | 850 | Clam | 2006 | 32.5 | 89.52 | 39.0384 | 89.51 | 39.0374 | 254.2 | 48.1035 | M3 | 0.1782 | -14.9818 | 0.1189 | -18.4964 | 0.315 | -10.0338 | M4 | EM3 - HM4 | -9.4616 | -13.0818 | OK | FALSE | -0.39649 | -8.1338 | 806.9841
37 |  |  | 1900 |  |  | 30 | 27.62 | 28.8245 | 23.81 | 27.5352 | 67 | 36.5215 | M3 | 0.08198 | -21.7258 | 0.0603 | -24.3937 | 0.129 | -17.7882 | M4 | EM3 - HM4 | -9.6755 | -9.8258 | OK | FALSE | -1.9785 | -5.8882 | 519.3798
38 | YES | GSM | 850 | Clam | 2006 | 32.5 | 86.9 | 38.7804 | 86.9 | 38.7804 | 246.8 | 47.8469 | M3 | 0.1433 | -16.8751 | 0.0955 | -20.3999 | 0.233 | -12.6529 | M4 | EM3 - HM4 | -9.7196 | -14.9751 | OK | FALSE | -0.6531 | -10.7529 | 1059.2275
38 |  |  | 1900 |  |  | 30.5 | 32.13 | 30.1382 | 26.79 | 28.5595 | 76.9 | 37.7185 | M3 | 0.085 | -21.4116 | 0.0743 | -22.5802 | 0.197 | -14.1107 | M3 | EM3 - HM3 | -8.3618 | -9.5116 | OK | FALSE | -0.78147 | -2.2107 | 390.3553
39 | YES | GSM | 850 | Clam | 2006 | 32.3 | 69.24 | 36.8071 | 69.23 | 36.8059 | 196.6 | 45.8717 | M3 | 0.1229 | -18.209 | 0.0902 | -20.8959 | 0.22 | -13.1515 | M4 | EM3 - HM4 | -11.6929 | -16.309 | OK | FALSE | -2.6283 | -11.2515 | 893.6364
39 |  |  | 1900 |  |  | 29.5 | 27.29 | 28.7201 | 23.41 | 27.388 | 67.2 | 36.5474 | M3 | 0.09441 | -20.4996 | 0.0638 | -23.9036 | 0.169 | -15.4423 | M3 | EM3 - HM3 | -9.7799 | -8.5996 | Not OK | 1.1803 | -1.9526 | -3.5423 | 397.6331
40 | YES | GSM | 850 | Clam | 2006 | 32.3 | 82.53 | 38.3322 | 82.09 | 38.2858 | 231.5 | 47.291 | M3 | 0.1867 | -14.5771 | 0.1191 | -18.4818 | 0.281 | -11.0259 | M4 | EM3 - HM4 | -10.1678 | -12.6771 | OK | FALSE | -1.209 | -9.1259 | 823.8434
40 |  |  | 1900 |  |  | 30 | 29.87 | 29.5047 | 28.05 | 28.9587 | 79.1 | 37.9635 | M3 | 0.1227 | -18.2231 | 0.0775 | -22.214 | 0.2 | -13.9794 | M3 | EM3 - HM3 | -8.9953 | -6.3231 | Not OK | 2.6722 | -0.53647 | -2.0794 | 395.5
41 | YES | GSM | 850 | Candy Bar | 2006 | 33 | 85.97 | 38.6869 | 85.53 | 38.6424 | 241.2 | 47.6475 | M3 | 0.2676 | -11.4503 | 0.2674 | -11.4568 | 0.631 | -3.9994 | M3 | EM3 - HM3 | -9.8131 | -9.5503 | Not OK | 0.26278 | -0.85245 | -2.0994 | 382.2504
41 |  |  | 1900 |  |  | 30 | 29.28 | 29.3314 | 28.05 | 28.9587 | 79.1 | 37.9635 | M3 | 0.1091 | -19.2435 | 0.0826 | -21.6604 | 0.213 | -13.4324 | M3 | EM3 - HM3 | -9.1686 | -7.3435 | Not OK | 1.8251 | -0.53647 | -1.5324 | 371.3615
42 | YES | GSM | 850 | Slider | 2006 | 32.3 | 86.62 | 38.7524 | 86.62 | 38.7524 | 247.7 | 47.8785 | M3 | 0.1463 | -16.6951 | 0.1021 | -19.8195 | 0.247 | -12.1461 | M4 | EM3 - HM4 | -9.7476 | -14.7951 | OK | FALSE | -0.62148 | -10.2461 | 1002.834
42 |  |  | 1900 |  |  | 30.5 | 27.35 | 28.7391 | 25.09 | 27.99 | 73 | 37.2665 | M3 | 0.0782 | -22.1359 | 0.0691 | -23.2104 | 0.183 | -14.751 | M3 | EM3 - HM3 | -9.7609 | -10.2359 | OK | FALSE | -1.2335 | -2.851 | 398.9071
43 | YES | GSM | 850 | Clam | 2006 | 32.8 | 54.29 | 34.6944 | 54.29 | 34.6944 | 153.1 | 43.6995 | M3 | 0.1372 | -17.2529 | 0.0877 | -21.14 | 0.207 | -13.6806 | M4 | EM3 - HM4 | -13.8056 | -15.3529 | OK | FALSE | -4.8005 | -11.7806 | 739.6135
43 |  |  | 1900 |  |  | 30 | 29.15 | 29.2928 | 28.58 | 29.1212 | 80.6 | 38.1267 | M3 | 0.1083 | -19.3074 | 0.0721 | -22.8413 | 0.186 | -14.6097 | M3 | EM3 - HM3 | -9.2072 | -7.4074 | Not OK | 1.7998 | -0.3733 | -2.7097 | 433.3333
44 | YES | GSM | 850 | Rotator | 2007 | 33 | 89.01 | 38.9888 | 89.01 | 38.9888 | 253.7 | 48.0864 | M3 | 0.1489 | -16.5421 | 0.1032 | -19.7264 | 0.255 | -11.8692 | M4 | EM3 - HM4 | -9.5112 | -14.6421 | OK | FALSE | -0.41359 | -9.9692 | 994.902
44 |  |  | 1900 |  |  | 30.5 | 29.29 | 29.3344 | 28.89 | 29.215 | 82.9 | 38.3711 | M3 | 0.08879 | -21.0327 | 0.0803 | -21.9057 | 0.212 | -13.4733 | M3 | EM3 - HM3 | -9.1656 | -9.1327 | Not OK | 0.032894 | -0.12891 | -1.5733 | 391.0377
45 | YES | GSM | 850 | Clam | 2007 | 33.5 | 24.18 | 27.6691 | 23.58 | 27.4509 | 67.2 | 36.5474 | M4 | 0.0533 | -25.4655 | 0.0508 | -25.8827 | 0.118 | -18.5624 | M4 | EM4 - HM4 | -20.8309 | -23.5655 | OK | FALSE | -11.9526 | -16.6624 | 569.4915
45 |  |  | 1900 |  |  | 30.5 | 19.24 | 25.6841 | 15.99 | 24.077 | 45.9 | 33.2363 | M4 | 0.05063 | -25.9118 | 0.0481 | -26.3571 | 0.129 | -17.7882 | M4 | EM4 - HM4 | -12.8159 | -14.0118 | OK | FALSE | -5.2637 | -5.8882 | 355.814
46 | YES | GSM | 850 | Clam | 2007 | 32.8 | 44.95 | 33.0546 | 44.95 | 33.0546 | 128.1 | 42.151 | M4 | 0.1057 | -19.5185 | 0.0746 | -22.5452 | 0.173 | -15.2391 | M4 | EM4 - HM4 | -15.4454 | -17.6185 | OK | FALSE | -6.349 | -13.3391 | 740.4624
46 |  |  | 1900 |  |  | 30 | 24.69 | 27.8504 | 23.76 | 27.5169 | 68.2 | 36.6757 | M3 | 0.0811 | -21.8196 | 0.0638 | -23.9036 | 0.171 | -15.3401 | M3 | EM3 - HM3 | -10.6496 | -9.9196 | Not OK | 0.73 | -1.8243 | -3.4401 | 398.8304
47 | YES | GSM | 850 | Clam | 2006 | 32.5 | 67.23 | 36.5513 | 67.23 | 36.5513 | 191.6 | 45.6479 | M3 | 0.1217 | -18.2942 | 0.0798 | -21.9599 | 0.197 | -14.1107 | M4 | EM3 - HM4 | -11.9487 | -16.3942 | OK | FALSE | -2.8521 | -12.2107 | 972.5888
47 |  |  | 1900 |  |  | 30 | 21.38 | 26.6002 | 19.41 | 25.7605 | 55.7 | 34.9171 | M3 | 0.07417 | -22.5954 | 0.0496 | -26.0904 | 0.131 | -17.6546 | M4 | EM3 - HM4 | -11.8998 | -10.6954 | Not OK | 1.2044 | -3.5829 | -5.7546 | 425.1908
48 | NO | iDEN | 800 |  | 2006 | 28.1 | 75.18 | 37.522 | 74.05 | 37.3905 | 371 | 51.3875 | M2 | 0.1451 | -16.7667 | 0.0946 | -20.4822 | 0.419 | -7.5557 | M4 | EM2 - HM4 | -13.478 | -17.3667 | OK | FALSE | 0.38748 | -8.1557 | 885.4415
48 |  |  | 900 |  |  | 28.1 | 59.04 | 35.4229 | 58.86 | 35.3964 | 284.3 | 49.0755 | M3 | 0.1413 | -16.9972 | 0.0933 | -20.6024 | 0.407 | -7.8081 | M4 | EM3 - HM4 | -15.5771 | -17.5972 | OK | FALSE | -1.9245 | -8.4081 | 698.5258
49 | NO | iDEN | 800 |  | 2006 | 28.1 | 76.56 | 37.68 | 75.25 | 37.5301 | 377 | 51.5268 | M2 | 0.1677 | -15.5093 | 0.1106 | -19.1249 | 0.49 | -6.1961 | M4 | EM2 - HM4 | -13.32 | -16.1093 | OK | FALSE | 0.52683 | -6.7961 | 769.3878
49 |  |  | 900 |  |  | 28.1 | 65.19 | 36.2836 | 63.87 | 36.1059 | 308.5 | 49.7851 | M3 | 0.1483 | -16.5772 | 0.0907 | -20.8479 | 0.438 | -7.1705 | M4 | EM3 - HM4 | -14.7164 | -17.1772 | OK | FALSE | -1.2149 | -7.7705 | 704.3379
50 | NO | iDEN | 800 | Candy Bar | 2006 | 28.1 | 69.97 | 36.8982 | 69.97 | 36.8982 | 343.5 | 50.7185 | M3 | 0.1152 | -18.771 | 0.1048 | -19.5928 | 0.456 | -6.8207 | M4 | EM3 - HM4 | -14.1018 | -19.371 | OK | FALSE | -0.28147 | -7.4207 | 753.2895
50 |  |  | 900 |  |  | 28.1 | 69.57 | 36.8484 | 69.57 | 36.8484 | 334 | 50.4749 | M3 | 0.109 | -19.2515 | 0.1018 | -19.845 | 0.445 | -7.0328 | M4 | EM3 - HM4 | -14.1516 | -19.8515 | OK | FALSE | -0.52507 | -7.6328 | 750.5618
51 | NO | iDEN | 800 | Candy Bar | 2006 | 28.06 | 83.52 | 38.4358 | 83.52 | 38.4358 | 410.1 | 52.2578 | M2 | 0.1195 | -18.4526 | 0.114 | -18.8619 | 0.496 | -6.0904 | M4 | EM2 - HM4 | -12.5642 | -19.0526 | OK | FALSE | 1.2578 | -6.6904 | 826.8145
51 |  |  | 900 |  |  | 28.06 | 94.83 | 39.5389 | 94.75 | 39.5316 | 454.8 | 53.1564 | M2 | 0.1345 | -17.4256 | 0.1341 | -17.4514 | 0.586 | -4.642 | M4 | EM2 - HM4 | -11.4611 | -18.0256 | OK | FALSE | 2.1564 | -5.242 | 776.1092
52 | NO | iDEN | 800 | Candy Bar | 2006 | 28.1 | 73.89 | 37.3717 | 73.89 | 37.3717 | 370.2 | 51.3687 | M2 | 0.1541 | -16.2439 | 0.1009 | -19.9222 | 0.447 | -6.9938 | M4 | EM2 - HM4 | -13.6283 | -16.8439 | OK | FALSE | 0.36873 | -7.5938 | 828.1879
52 |  |  | 900 |  |  | 28.1 | 63.69 | 36.0814 | 63.69 | 36.0814 | 307.6 | 49.7597 | M3 | 0.1399 | -17.0836 | 0.0943 | -20.5098 | 0.411 | -7.7232 | M4 | EM3 - HM4 | -14.9186 | -17.6836 | OK | FALSE | -1.2403 | -8.3232 | 748.4185
53 | NO | iDEN | 800 | Candy Bar | 2006 | 28.06 | 66.2 | 36.4172 | 65.9 | 36.3777 | 333.5 | 50.4619 | M3 | 0.1155 | -18.7484 | 0.1148 | -18.8012 | 0.488 | -6.2316 | M4 | EM3 - HM4 | -14.5828 | -19.3484 | OK | FALSE | -0.53808 | -6.8316 | 683.4016
53 |  |  | 900 |  |  | 28.06 | 100.9 | 40.0778 | 100.9 | 40.0778 | 491.5 | 53.8305 | M2 | 0.1897 | -14.4387 | 0.1482 | -16.583 | 0.63 | -4.0132 | M3 | EM2 - HM3 | -10.9222 | -15.0387 | OK | FALSE | 2.8305 | -4.6132 | 780.1587
54 | NO | GSM | 850 | QWERTY Candy Bar | 2006 | 32.5 | 83.48 | 38.4316 | 80.14 | 38.077 | 228.4 | 47.1739 | M3 | 0.1674 | -15.5249 | 0.115 | -18.786 | 0.284 | -10.9336 | M4 | EM3 - HM4 | -10.0684 | -13.6249 | OK | FALSE | -1.3261 | -9.0336 | 804.2254
54 |  |  | 1900 |  |  | 29.5 | 40.27 | 32.0996 | 32.02 | 30.1084 | 91.9 | 39.2663 | M2 | 0.1215 | -18.3085 | 0.0875 | -21.1598 | 0.231 | -12.7278 | M3 | EM2 - HM3 | -6.4004 | -6.4085 | OK | FALSE | 0.76631 | -0.82776 | 397.8355
54 | NO | WCDMA | 850 |  | 2006 | 24 | 85.75 | 38.6647 | 82.61 | 38.3407 | 76 | 37.6163 | M4 | 0.1517 | -16.3803 | 0.1187 | -18.511 | 0.108 | -19.3315 | M4 | EM4 - HM4 | -12.3353 | -16.9803 | OK | FALSE | -13.3837 | -19.9315 | 703.7037
54 |  |  | 1900 |  |  | 24 | 61.2 | 35.735 | 42.17 | 32.5001 | 38.8 | 31.7766 | M4 | 0.1607 | -15.8797 | 0.1213 | -18.3228 | 0.108 | -19.3315 | M4 | EM4 - HM4 | -5.265 | -6.4797 | OK | FALSE | -9.2234 | -9.9315 | 359.2593
55 | NO | GSM | 850 | Clam | 2006 | 33 | 99.32 | 39.9407 | 99.26 | 39.9355 | 281.9 | 49.0019 | M2 | 0.1532 | -16.2948 | 0.0981 | -20.1666 | 0.26 | -11.7005 | M4 | EM2 - HM4 | -8.5593 | -14.3948 | OK | FALSE | 0.5019 | -9.8005 | 1084.2308
55 |  |  | 1900 |  |  | 30 | 44.22 | 32.9124 | 44.22 | 32.9124 | 126 | 42.0074 | M2 | 0.1334 | -17.4969 | 0.1159 | -18.7183 | 0.248 | -12.111 | M2 | EM2 - HM2 | -5.5876 | -5.5969 | OK | FALSE | 3.5074 | -0.21097 | 508.0645
56 | NO | GSM | 850 | Candy Bar | 2006 | 32.6 | 129.5 | 42.2454 | 129.5 | 42.2454 | 370.4 | 51.3734 | M2 | 0.2553 | -11.859 | 0.1778 | -15.0014 | 0.416 | -7.6181 | M4 | EM2 - HM4 | -6.2546 | -9.959 | OK | FALSE | 2.8734 | -5.7181 | 890.3846
56 |  |  | 1900 |  |  | 30.4 | 52.83 | 34.4576 | 52.36 | 34.38 | 150.8 | 43.568 | M1 | 0.166 | -15.5978 | 0.1436 | -16.8569 | 0.359 | -8.8981 | M2 | EM1 - HM2 | -4.0424 | -3.6978 | Not OK | 0.34455 | 5.068 | 3.0019 | 420.0557
57 | NO | GSM | 850 | Clam | 2006 | 32.8 | 117.7 | 41.4155 | 117.7 | 41.4155 | 360.1 | 51.1285 | M2 | 0.2085 | -13.6179 | 0.1264 | -17.9651 | 0.34 | -9.3704 | M4 | EM2 - HM4 | -7.0845 | -11.7179 | OK | FALSE | 2.6285 | -7.4704 | 1059.1176
57 |  |  | 1900 |  |  | 30.2 | 40.64 | 32.1791 | 40.63 | 32.1769 | 117 | 41.3637 | M2 | 0.1305 | -17.6878 | 0.1065 | -19.453 | 0.247 | -12.1461 | M3 | EM2 - HM3 | -6.3209 | -5.7878 | Not OK | 0.53314 | 2.8637 | -0.24606 | 473.6842
58 | NO | GSM | 850 | Clam | Pre-2006 | 32.5 | 101.3 | 40.1122 | 101.3 | 40.1122 | 290.6 | 49.2659 | >M1 | 0.2008 | -13.9447 | 0.1333 | -17.5034 | 0.356 | -8.971 | M2 | E>M1 - HM2 | -8.3878 | -12.0447 | OK | FALSE | 0.76591 | -7.071 | 816.2921
58 |  |  | 1900 |  |  | 30 | 44.91 | 33.0469 | 44.7 | 33.0062 | 134.1 | 42.5486 | M2 | 0.1486 | -16.5596 | 0.1263 | -17.9719 | 0.288 | -10.8122 | M2 | EM2 - HM2 | -5.4531 | -4.6596 | Not OK | 0.79352 | 4.0486 | 1.0878 | 465.625
59 | NO | GSM | 850 | Candy Bar | Pre-2006 | 32.3 | 124 | 41.8684 | 124 | 41.8684 | 353.5 | 50.9678 | >M1 | 0.2851 | -10.9001 | 0.2 | -13.9794 | 0.54 | -5.3521 | M1 | E>M1 - HM1 | -6.6316 | -9.0001 | OK | FALSE | 2.4678 | -3.4521 | 654.6296
59 |  |  | 1900 |  |  | 30.2 | 55.73 | 34.9218 | 50.49 | 34.0641 | 145.4 | 43.2513 | M2 | 0.1509 | -16.4262 | 0.1339 | -17.4644 | 0.3 | -10.4576 | M2 | EM2 - HM2 | -3.5782 | -4.5262 | OK | FALSE | 4.7513 | 1.4424 | 484.6667
60 | NO | GSM | 850 | Slider | 2006 | 33 | 126.2 | 42.0212 | 126.2 | 42.0212 | 358.4 | 51.0874 | M2 | 0.2277 | -12.8527 | 0.1585 | -15.9994 | 0.42 | -7.535 | M4 | EM2 - HM4 | -6.4788 | -10.9527 | OK | FALSE | 2.5874 | -5.635 | 853.3333
60 |  |  | 1900 |  |  | 30 | 36.05 | 31.1381 | 26.74 | 28.5432 | 76.2 | 37.6391 | M3 | 0.1174 | -18.6066 | 0.0832 | -21.5975 | 0.178 | -14.9916 | M3 | EM3 - HM3 | -7.3619 | -6.7066 | Not OK | 0.65526 | -0.8609 | -3.0916 | 428.0899
61 | NO | GSM | 850 | Candy Bar | 2006 | 33 | 87.25 | 38.8153 | 87.25 | 38.8153 | 249.5 | 47.9414 | M3 | 0.1738 | -15.199 | 0.1181 | -18.555 | 0.289 | -10.782 | M4 | EM3 - HM4 | -9.6847 | -13.299 | OK | FALSE | -0.55859 | -8.882 | 863.3218
61 |  |  | 1900 |  |  | 30.3 | 48.72 | 33.7541 | 30.99 | 29.8244 | 90.2 | 39.1041 | M2 | 0.1296 | -17.7479 | 0.0966 | -20.3005 | 0.256 | -11.8352 | M2 | EM2 - HM2 | -4.7459 | -5.8479 | OK | FALSE | 0.60413 | 0.064799 | 352.3438
62 | NO | GSM | 850 | Clam | Pre-2006 | 33 | 63.06 | 35.9951 | 63.06 | 35.9951 | 184.1 | 45.3011 | M3 | 0.1145 | -18.8239 | 0.0776 | -22.2028 | 0.208 | -13.6387 | M4 | EM3 - HM4 | -12.5049 | -16.9239 | OK | FALSE | -3.1989 | -11.7387 | 885.0962
62 |  |  | 1900 |  |  | 30 | 38.38 | 31.6821 | 30.13 | 29.58 | 90.4 | 39.1234 | M2 | 0.08927 | -20.9859 | 0.0711 | -22.9626 | 0.162 | -15.8097 | M3 | EM2 - HM3 | -6.8179 | -9.0859 | OK | FALSE | 0.62337 | -3.9097 | 558.0247
63 | NO | GSM | 850 | Clam | 2006 | 33 | 105.5 | 40.465 | 105.5 | 40.465 | 301.6 | 49.5886 | M2 | 0.1972 | -14.1019 | 0.126 | -17.9926 | 0.305 | -10.314 | M4 | EM2 - HM4 | -8.035 | -12.2019 | OK | FALSE | 1.0886 | -8.414 | 988.8525
63 |  |  | 1900 |  |  | 30.5 | 38.03 | 31.6025 | 37.38 | 31.4528 | 108.8 | 40.7326 | M2 | 0.1125 | -18.9769 | 0.0928 | -20.649 | 0.246 | -12.1813 | M3 | EM2 - HM3 | -6.8975 | -7.0769 | OK | FALSE | 2.2326 | -0.2813 | 442.2764
64 | NO | GSM | 850 | Clam | 2006 | 33 | 112.9 | 41.0539 | 112.9 | 41.0539 | 329.7 | 50.3624 | M2 | 0.2283 | -12.8299 | 0.1552 | -16.1822 | 0.416 | -7.6181 | M4 | EM2 - HM4 | -7.4461 | -10.9299 | OK | FALSE | 1.8624 | -5.7181 | 792.5481
64 |  |  | 1900 |  |  | 30.5 | 37.1 | 31.3875 | 34.7 | 30.8066 | 104.7 | 40.3989 | M2 | 0.1236 | -18.1596 | 0.0947 | -20.473 | 0.216 | -13.3109 | M3 | EM2 - HM3 | -7.1125 | -6.2596 | Not OK | 0.85289 | 1.8989 | -1.4109 | 484.7222
65 | NO | GSM | 850 | QWERTY Side Slider | 2006 | 32.7 | 86.3 | 38.7202 | 69.61 | 36.8534 | 199.1 | 45.9814 | M3 | 0.2328 | -12.6603 | 0.2293 | -12.7919 | 0.555 | -5.1141 | M3 | EM3 - HM3 | -9.7798 | -10.7603 | OK | FALSE | -2.5186 | -3.2141 | 358.7387
65 |  |  | 1900 |  |  | 30 | 50.61 | 34.0847 | 44.95 | 33.0546 | 130.8 | 42.3322 | M2 | 0.1581 | -16.0214 | 0.1577 | -16.0434 | 0.418 | -7.5765 | M2 | EM2 - HM2 | -4.4153 | -4.1214 | Not OK | 0.29391 | 3.8322 | 4.3235 | 312.9187
66 | NO | GSM | 850 | Clam | Pre-2006 | 33 | 116.8 | 41.3489 | 116.8 | 41.3489 | 335.1 | 50.5035 | >M1 | 0.2576 | -11.7811 | 0.1693 | -15.4269 | 0.452 | -6.8972 | M1 | E>M1 - HM1 | -7.1511 | -9.8811 | OK | FALSE | 2.0035 | -4.9972 | 741.3717
66 |  |  | 1900 |  |  | 30 | 28.01 | 28.9463 | 22 | 26.8485 | 66 | 36.3909 | M3 | 0.07776 | -22.1849 | 0.057 | -24.8825 | 0.13 | -17.7211 | M4 | EM3 - HM4 | -9.5537 | -10.2849 | OK | FALSE | -2.1091 | -5.8211 | 507.6923
67 | NO | GSM | 850 | Slider | 2006 |  | 106.05 | 40.5102 | 106.05 | 40.5102 | 301.2 | 49.5771 | M2 | 0.1825 | -14.7747 | 0.1268 | -17.9376 | 0.336 | -9.4732 | M4 | EM2 - HM4 | -7.9898 | -12.8747 | OK | FALSE | 1.0771 | -7.5732 | 896.4286
67 |  |  | 1900 |  |  |  | 35.79 | 31.0752 | 27.72 | 28.8559 | 79 | 37.9525 | M3 | 0.1202 | -18.4019 | 0.0902 | -20.8959 | 0.193 | -14.2889 | M3 | EM3 - HM3 | -7.4248 | -6.5019 | Not OK | 0.92286 | -0.54746 | -2.3889 | 409.3264
68 | NO | GSM | 850 | Candy Bar | Pre-2006 | 32.8 | 67.65 | 36.6054 | 62.87 | 35.9689 | 179.2 | 45.0668 | M1 | 0.2757 | -11.1913 | 0.2707 | -11.3502 | 0.731 | -2.7217 | M1 | EM1 - HM1 | -11.8946 | -9.2913 | Not OK | 2.6034 | -3.4332 | -0.82165 | 245.1436
68 |  |  | 1900 |  |  | 30 | 31.75 | 30.0349 | 30.52 | 29.6917 | 87.9 | 38.8798 | M2 | 0.1058 | -19.5103 | 0.1058 | -19.5103 | 0.237 | -12.505 | M3 | EM2 - HM3 | -8.4651 | -7.6103 | Not OK | 0.85484 | 0.37978 | -0.60503 | 370.8861
69 | NO | GSM | 850 | Candy Bar | 2006 | 33.1 | 103.6 | 40.3072 | 103.6 | 40.3072 | 294.2 | 49.3729 | M2 | 0.165 | -15.6503 | 0.1105 | -19.1328 | 0.293 | -10.6626 | M4 | EM2 - HM4 | -8.1928 | -13.7503 | OK | FALSE | 0.87285 | -8.7626 | 1004.0956
69 |  |  | 1900 |  |  | 30 | 61.75 | 35.8127 | 40.88 | 32.2302 | 116.5 | 41.3265 | M2 | 0.15 | -16.4782 | 0.1495 | -16.5072 | 0.32 | -9.897 | M2 | EM2 - HM2 | -2.6873 | -4.5782 | OK | FALSE | 2.8265 | 2.003 | 364.0625
70 | NO | GSM | 850 | Candy Bar | 2006 | 32.5 | 102.9 | 40.2483 | 102.9 | 40.2483 | 293.5 | 49.3522 | M2 | 0.2672 | -11.4633 | 0.2129 | -13.4365 | 0.494 | -6.1255 | M3 | EM2 - HM3 | -8.2517 | -9.5633 | OK | FALSE | 0.85216 | -4.2255 | 594.1296
70 |  |  | 1900 |  |  | 30 | 24.9 | 27.924 | 24.04 | 27.6187 | 69 | 36.777 | M3 | 0.09421 | -20.5181 | 0.0649 | -23.7551 | 0.174 | -15.189 | M3 | EM3 - HM3 | -10.576 | -8.6181 | Not OK | 1.958 | -1.723 | -3.289 | 396.5517
71 | NO | GSM | 850 | Clam | 2006 | 32.5 | 77.7 | 37.8084 | 77.7 | 37.8084 | 222.2 | 46.9349 | M3 | 0.1624 | -15.7883 | 0.1089 | -19.2594 | 0.269 | -11.405 | M4 | EM3 - HM4 | -10.6916 | -13.8883 | OK | FALSE | -1.5651 | -9.505 | 826.0223
71 |  |  | 1900 |  |  | 29.7 | 24.55 | 27.801 | 23.89 | 27.5643 | 68.8 | 36.7518 | M3 | 0.08254 | -21.6667 | 0.0645 | -23.8088 | 0.169 | -15.4423 | M3 | EM3 - HM3 | -10.699 | -9.7667 | Not OK | 0.93226 | -1.7482 | -3.5423 | 407.1006
72 | NO | GSM | 850 | Candy Bar | 2006 | 33 | 109 | 40.7485 | 108.87 | 40.7382 | 307 | 49.7428 | M2 | 0.2374 | -12.4904 | 0.1563 | -16.1208 | 0.369 | -8.6595 | M4 | EM2 - HM4 | -7.7515 | -10.5904 | OK | FALSE | 1.2428 | -6.7595 | 831.9783
72 |  |  | 1900 |  |  | 30 | 43.79 | 32.8275 | 39.96 | 32.0325 | 112.7 | 41.0385 | M2 | 0.1416 | -16.9787 | 0.1078 | -19.3476 | 0.278 | -11.1191 | M2 | EM2 - HM2 | -5.6725 | -5.0787 | Not OK | 0.59377 | 2.5385 | 0.7809 | 405.3957
73 | NO | GSM | 850 | Candy Bar | 2006 | 33 | 105.1 | 40.4321 | 104.4 | 40.374 | 297.4 | 49.4668 | M2 | 0.2021 | -13.8887 | 0.1385 | -17.171 | 0.342 | -9.3195 | M4 | EM2 - HM4 | -8.0679 | -11.9887 | OK | FALSE | 0.96682 | -7.4195 | 869.5906
73 |  |  | 1900 |  |  | 30.5 | 49.14 | 33.8287 | 47.7 | 33.5704 | 136.9 | 42.7281 | M2 | 0.1734 | -15.219 | 0.1348 | -17.4062 | 0.356 | -8.971 | M2 | EM2 - HM2 | -4.6713 | -3.319 | Not OK | 1.3523 | 4.2281 | 2.929 | 384.5506
[6 empty] | Peak / pulsed averaged channel power (dBm) | [15 empty] | Delta (dB) E-Field relative to M3 (pass) | Delta (dB) H-Field relative to M3 (pass)
101 | Pass | iDEN 800 |  | Candy Bar PDA | 2006 | 28 |  |  |  |  | 266.3 | 48.5074 | 3 |  |  |  |  | 0.704 | -3.0485 | 3 | 3 | -2.4926 | -3.6485 | OK |  |  |  | 378.267
101 |  | iDEN 900 |  |  | 2006 | 28 |  |  |  |  | 239.4 | 47.5825 | 3 |  |  |  |  | 0.575 | -4.8066 | 4 |  | -3.4175 | -5.4066 | OK |  |  |  | 416.3478
102 | Pass | iDEN 800 |  | Candy Bar PDA | 2006 | 28 |  |  |  |  | 235.62 | 47.4442 | 3 |  |  |  |  | 0.693 | -3.1853 | 3 | 3 | -3.5558 | -3.7853 | OK |  |  |  | 340
103 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 238.1 | 47.5352 | 3 |  |  |  |  | 0.499 | -6.038 | 3 | 2 | -0.96481 | -4.138 | OK |  |  |  | 477.1543
103 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 96.6 | 39.6995 | 2 |  |  |  |  | 0.281 | -11.0259 | 2 |  | 1.1995 | 0.87413 | OK |  |  |  | 343.7722
104 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 197.2 | 45.8981 | 3 |  |  |  |  | 0.404 | -7.8724 | 3 | 3 | -2.6019 | -5.9724 | OK |  |  |  | 488.1188
104 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 112.9 | 41.0539 | 2 |  |  |  |  | 0.308 | -10.229 | 2 |  | 2.5539 | 1.671 | OK |  |  |  | 366.5584
105 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 259.6 | 48.2861 | 4 |  |  |  |  | 0.507 | -5.8998 | 3 | 2 | -0.21391 | -3.9998 | OK |  |  |  | 512.0316
105 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 114.9 | 41.2064 | 2 |  |  |  |  | 0.303 | -10.3711 | 2 |  | 2.7064 | 1.5289 | OK |  |  |  | 379.2079
106 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 279.1 | 48.9152 | 2 |  |  |  |  | 0.499 | -6.038 | 3 | 2 | 0.4152 | -4.138 | OK |  |  |  | 559.3186
106 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 147.4 | 43.3699 | 2 |  |  |  |  | 0.37 | -8.636 | 2 |  | 4.8699 | 3.264 | OK |  |  |  | 398.3784
107 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 210.4 | 46.4609 | 3 |  |  |  |  | 0.41 | -7.7443 | 4 | 3 | -2.0391 | -5.8443 | OK |  |  |  | 513.1707
107 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 82.68 | 38.348 | 3 |  |  |  |  | 0.23 | -12.7654 | 3 |  | -0.15199 | -0.86544 | OK |  |  |  | 359.4783
108 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 224.8 | 47.0359 | 3 |  |  |  |  | 0.41 | -7.7443 | 4 | 3 | -1.4641 | -5.8443 | OK |  |  |  | 548.2927
108 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 82.68 | 38.348 | 3 |  |  |  |  | 0.23 | -12.7654 | 3 |  | -0.15199 | -0.86544 | OK |  |  |  | 359.4783
109 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 224.9 | 47.0398 | 3 |  |  |  |  | 0.41 | -7.7443 | 4 | 2 | -1.4602 | -5.8443 | OK |  |  |  | 548.5366
109 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 132.1 | 42.4181 | 2 |  |  |  |  | 0.327 | -9.709 | 2 |  | 3.9181 | 2.191 | OK |  |  |  | 403.9755
110 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 160.6 | 44.1149 | 3 |  |  |  |  | 0.273 | -11.2767 | 4 | 2 | -4.3851 | -9.3767 | OK |  |  |  | 588.2784
110 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 105.6 | 40.4733 | 2 |  |  |  |  | 0.284 | -10.9336 | 2 |  | 1.9733 | 0.96637 | OK |  |  |  | 371.831
111 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 229.9 | 47.2308 | 3 |  |  |  |  | 0.413 | -7.681 | 4 | 2 | -1.2692 | -5.781 | OK |  |  |  | 556.6586
111 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 128.6 | 42.1848 | 2 |  |  |  |  | 0.32 | -9.897 | 2 |  | 3.6848 | 2.003 | OK |  |  |  | 401.875
112 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 298.1 | 49.4872 | 2 |  |  |  |  | 0.53 | -5.5145 | 3 | 1 | 0.98724 | -3.6145 | OK |  |  |  | 562.4528
112 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 158.2 | 43.9841 | 1 |  |  |  |  | 0.43 | -7.3306 | 2 |  | 5.4841 | 4.5694 | OK |  |  |  | 367.907
113 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 227.6 | 47.1434 | 3 |  |  |  |  | 0.45 | -6.9357 | 4 | 3 | -1.3566 | -5.0357 | OK |  |  |  | 505.7778
113 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 67.6 | 36.5989 | 3 |  |  |  |  | 0.17 | -15.391 | 3 |  | -1.9011 | -3.491 | OK |  |  |  | 397.6471
114 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 282.2 | 49.0111 | 2 |  |  |  |  | 0.529 | -5.5309 | 3 | 2 | 0.51114 | -3.6309 | OK |  |  |  | 533.4594
114 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 103.2 | 40.2736 | 2 |  |  |  |  | 0.29 | -10.752 | 2 |  | 1.7736 | 1.148 | OK |  |  |  | 355.8621
115 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 257.5 | 48.2155 | 3 |  |  |  |  | 0.34 | -9.3704 | 4 | 2 | -0.28446 | -7.4704 | OK |  |  |  | 757.3529
115 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 147.6 | 43.3817 | 2 |  |  |  |  | 0.39 | -8.1787 | 2 |  | 4.8817 | 3.7213 | OK |  |  |  | 378.4615
116 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 230.7 | 47.261 | 3 |  |  |  |  | 0.354 | -9.0199 | 4 | 1 | -1.239 | -7.1199 | OK |  |  |  | 651.6949
116 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 154.5 | 43.7786 | 1 |  |  |  |  | 0.46 | -6.7448 | 1 |  | 5.2786 | 5.1552 | OK |  |  |  | 335.8696
117 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 244 | 47.7478 | 3 |  |  |  |  | 0.363 | -8.8019 | 4 | 1 | -0.7522 | -6.9019 | OK |  |  |  | 672.1763
117 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 153 | 43.6938 | 1 |  |  |  |  | 0.448 | -6.9744 | 2 |  | 5.1938 | 4.9256 | OK |  |  |  | 341.5179
118 | Pass | GSM 1900 |  | Candy Bar PDA | 2006 | 30 |  |  |  |  | 79.1 | 37.9635 | 3 |  |  |  |  | 0.22 | -13.1515 | 3 | 3 | -0.53647 | -1.2515 | OK |  |  |  | 359.5455
119 | Fail | GSM 850 |  |  | 2006 | 33 |  |  |  |  | 106.5 | 40.547 | 4 |  |  |  |  | 0.286 | -10.8727 | 4 | 2 | -7.953 | -8.9727 | OK |  |  |  | 372.3776
119 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 97.2 | 39.7533 | 2 |  |  |  |  | 0.287 | -10.8424 | 2 |  | 1.2533 | 1.0576 | OK |  |  |  | 338.676
120 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | off |  |  |  |  | - |  | - |  |  |  |  | - |  | - | 3
120 |  | GSM 1900 |  |  | 2006 | 29 |  |  |  |  | 81.6 | 38.2338 | 3 |  |  |  |  | 0.242 | -12.3237 | 3 |  | -0.2662 | -0.42369 | OK |  |  |  | 337.1901
121 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 29 |  |  |  |  | 77.1 | 37.7411 | 4 |  |  |  |  | 0.264 | -11.5679 | 4 | 2 | -10.7589 | -9.6679 | Not OK | 1.091 |  |  | 292.0455
121 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 81.3 | 38.2018 | 3 |  |  |  |  | 0.247 | -12.1461 | 3 |  | -0.29819 | -0.24606 | Not OK | 0.052128 |  |  | 329.1498
122 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 29 |  |  |  |  | 88.2 | 38.9094 | 4 |  |  |  |  | 0.323 | -9.8159 | 4 | 2 | -9.5906 | -7.9159 | Not OK | 1.6747 |  |  | 273.065
122 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 86.9 | 38.7804 | 2 |  |  |  |  | 0.287 | -10.8424 | 2 |  | 0.2804 | 1.0576 | Not OK | 0.77724 |  |  | 302.7875
123 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 29 |  |  |  |  | 86.5 | 38.7403 | 4 |  |  |  |  | 0.306 | -10.2856 | 4 | 2 | -9.7597 | -8.3856 | Not OK | 1.3741 |  |  | 282.6797
123 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 91.52 | 39.2303 | 2 |  |  |  |  | 0.26 | -11.7005 | 2 |  | 0.73032 | 0.19947 | OK |  |  |  | 352
124 | Fail | GSM 1900 |  | Candy Bar PDA | 2006 | 30 |  |  |  |  | 108.8 | 40.7326 | 2 |  |  |  |  | 0.343 | -9.2941 | 2 | 2 | 2.2326 | 2.6059 | Not OK | 0.3733 |  |  | 317.2012
125 | Fail | GSM 1900 |  | Candy Bar PDA | 2006 | 30 |  |  |  |  | 96.4 | 39.6815 | 2 |  |  |  |  | 0.319 | -9.9242 | 2 | 2 | 1.1815 | 1.9758 | Not OK | 0.79427 |  |  | 302.1944
126 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 170.1 | 44.6141 | 3 |  |  |  |  | 0.303 | -10.3711 | 4 | 2 | -3.8859 | -8.4711 | OK |  |  |  | 561.3861
126 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 145.1 | 43.2333 | 2 |  |  |  |  | 0.39 | -8.1787 | 2 |  | 4.7333 | 3.7213 | OK |  |  |  | 372.0513
127 | Pass | 802.11b/g |  | Candy Bar PDA | 2006 | 18 |  |  |  |  | 14.4 | 23.1672 | 4 |  |  |  |  | 0.037 | -28.636 | 4 | 4 | -17.8328 | -19.236 | OK |  |  |  | 389.1892
128 | Pass | 802.11b./g |  |  | 2006 | 18 |  |  |  |  | 29.1 | 29.2779 | 4 |  |  |  |  | 0.074 | -22.6154 | 4 | 4 | -11.7221 | -13.2154 | OK |  |  |  | 393.2432
129 | Pass | 802.11b/g |  | Candy Bar PDA | 2006 | 18 |  |  |  |  | 25.6 | 28.1648 | 4 |  |  |  |  | 0.05 | -26.0206 | 4 | 4 | -12.8352 | -16.6206 | OK |  |  |  | 512
130 | Pass | 802.11a |  | Candy Bar PDA | 2006 | 17 |  |  |  |  | 13.1 | 22.3454 | 4 | [7 empty] | 4 | -18.6546
131 | Pass | 802.11b/g |  | Candy Bar PDA | 2006 | 18 |  |  |  |  | 29.4 | 29.3669 | 4 |  |  |  |  | 0.09 | -20.9151 |  | 4 | -11.6331 | -11.5151 | Not OK | 0.1179 |  |  | 326.6667
132 | Pass | 802.11a |  | Candy Bar PDA | 2006 | 17 |  |  |  |  | 13.1 | 22.3454 | 4 | [7 empty] | 4 | -18.6546
133 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 92.1 | 39.2852 | 4 |  |  |  |  | 0.165 | -15.6503 | 4 | 4 | -11.7148 | -16.2503 | OK |  |  |  | 558.1818
133 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 36.2 | 31.1742 | 4 |  |  |  |  | 0.1 | -20 | 4 |  | -9.8258 | -10.6 | OK |  |  |  | 362
134 | Pass | WCDMA FDD V 850 |  | Candy Bar PDA | 2006 | 23.5 |  |  |  |  | 63 | 35.9868 | 4 |  |  |  |  | 0.13 | -17.7211 | 4 | 4 | -15.0132 | -18.3211 | OK |  |  |  | 484.6154
134 |  | WCDMA FDD II 1900 |  |  | 2006 | 24 |  |  |  |  | 47.9 | 33.6067 | 4 |  |  |  |  | 0.117 | -18.6363 | 4 |  | -7.3933 | -9.2363 | OK |  |  |  | 409.4017
135 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 94.5 | 39.5086 | 4 |  |  |  |  | 0.169 | -15.4423 | 4 | 4 | -11.4914 | -16.0423 | OK |  |  |  | 559.1716
135 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 62.8 | 35.9592 | 4 |  |  |  |  | 0.161 | -15.8635 | 4 |  | -5.0408 | -6.4635 | OK |  |  |  | 390.0621
136 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 145 | 43.2274 | 4 |  |  |  |  | 0.26 | -11.7005 | 4 | 3 | -7.7726 | -12.3005 | OK |  |  |  | 557.6923
136 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 68.2 | 36.6757 | 3 |  |  |  |  | 0.15 | -16.4782 | 4 |  | -4.3243 | -7.0782 | OK |  |  |  | 454.6667
137 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 106.5 | 40.547 | 4 |  |  |  |  | 0.23 | -12.7654 | 4 | 4 | -10.453 | -13.3654 | OK |  |  |  | 463.0435
137 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 48.7 | 33.7506 | 4 |  |  |  |  | 0.12 | -18.4164 | 4 |  | -7.2494 | -9.0164 | OK |  |  |  | 405.8333
138 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 91.1 | 39.1904 | 4 |  |  |  |  | 0.134 | -17.4579 | 4 | 3 | -11.8096 | -18.0579 | OK |  |  |  | 679.8507
138 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 72.2 | 37.1707 | 3 |  |  |  |  | 0.191 | -14.3793 | 3 |  | -3.8293 | -4.9793 | OK |  |  |  | 378.0105
139 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 52.3 | 34.37 | 4 |  |  |  |  | 0.192 | -14.334 | 4 | 3 | -16.63 | -14.934 | Not OK | 1.696 |  |  | 272.3958
139 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 34.9 | 30.8565 | 4 |  |  |  |  | 0.092 | -20.7242 | 4 |  | -10.1435 | -11.3242 | OK |  |  |  | 379.3478
140 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 60.8 | 35.6781 | 4 |  |  |  |  | 0.121 | -18.3443 | 4 | 4 | -15.3219 | -18.9443 | OK |  |  |  | 502.4793
140 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 38.8 | 31.7766 | 4 |  |  |  |  | 0.133 | -17.523 | 4 |  | -9.2234 | -8.123 | Not OK | 1.1004 |  |  | 291.7293
141 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 63.1 | 36.0006 | 4 |  |  |  |  | 0.13 | -17.7211 | 4 | 4 | -14.9994 | -18.3211 | OK |  |  |  | 485.3846
141 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 36.2 | 31.1742 | 4 |  |  |  |  | 0.128 | -17.8558 | 4 |  | -9.8258 | -8.4558 | Not OK | 1.37 |  |  | 282.8125
142 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 66.2 | 36.4172 | 4 |  |  |  |  | 0.127 | -17.9239 | 4 |  | -14.5828 | -18.5239 | OK |  |  |  | 521.2598
142 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 33.9 | 30.604 | 4 |  |  |  |  | 0.119 | -18.4891 | 4 |  | -10.396 | -9.0891 | Not OK | 1.3069 |  |  | 284.8739
201 | Fail | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 237.8 | 47.5242 | 3 |  |  |  |  | 0.491 | -6.1784 | 3 | 2 | -0.97576 | -4.2784 | OK |  |  |  | 484.3177
201 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 128.4 | 42.1713 | 2 |  |  |  |  | 0.302 | -10.3999 | 2 |  | 3.6713 | 1.5001 | OK |  |  |  | 425.1656
202 | Pass | GSM 850 |  | Flip | 2007 | 33 |  |  |  |  | 188.09 | 45.4873 | 3 |  |  |  |  | 0.396 | -8.0461 | 3 | 3 | -3.0127 | -6.1461 | OK |  |  |  | 474.9747
202 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 75.98 | 37.614 | 3 |  |  |  |  | 0.207 | -13.6806 | 3 |  | -0.88601 | -1.7806 | OK |  |  |  | 367.0531
203 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 211.2 | 46.4939 | 3 |  |  |  |  | 0.5 | -6.0206 | 4 | 3 | -2.0061 | -4.1206 | OK |  |  |  | 422.4
203 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 67.1 | 36.5345 | 3 |  |  |  |  | 0.216 | -13.3109 | 3 |  | -1.9655 | -1.4109 | Not OK | 0.55462 |  |  | 310.6481
203 |  | CDMA 800 |  |  | 2007 | 24.5 |  |  |  |  | 74.3 | 37.4198 | 4 |  |  |  |  | 0.183 | -14.751 | 4 |  | -13.5802 | -15.351 | OK |  |  |  | 406.0109
203 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 25.4 | 28.0967 | 4 |  |  |  |  | 0.089 | -21.0122 | 4 |  | -12.9033 | -11.6122 | Not OK | 1.2911 |  |  | 285.3933
204 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 212 | 46.5267 | 3 |  |  |  |  | 0.462 | -6.7072 | 3 | 3 | -1.9733 | -4.8072 | OK |  |  |  | 458.8745
204 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 72.7 | 37.2307 | 3 |  |  |  |  | 0.225 | -12.9563 | 3 |  | -1.2693 | -1.0563 | Not OK | 0.21296 |  |  | 323.1111
205 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 182.6 | 45.23 | 3 |  |  |  |  | 0.391 | -8.1565 | 4 | 3 | -3.27 | -6.2565 | OK |  |  |  | 467.0077
205 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 75.6 | 37.5704 | 3 |  |  |  |  | 0.228 | -12.8413 | 3 |  | -0.92956 | -0.9413 | OK |  |  |  | 331.5789
205 |  | CDMA 800 |  |  | 2007 | 24.5 |  |  |  |  | 58.1 | 35.2835 | 4 |  |  |  |  | 0.107 | -19.4123 | 4 |  | -15.7165 | -20.0123 | OK |  |  |  | 542.9907
205 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 36.7 | 31.2933 | 4 |  |  |  |  | 0.104 | -19.6593 | 4 |  | -9.7067 | -10.2593 | OK |  |  |  | 352.8846
206 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 179.2 | 45.0668 | 3 |  |  |  |  | 0.37 | -8.636 | 4 | 3 | -3.4332 | -6.736 | OK |  |  |  | 484.3243
206 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 70.4 | 36.9515 | 3 |  |  |  |  | 0.175 | -15.1392 | 3 |  | -1.5485 | -3.2392 | OK |  |  |  | 402.2857
207 | Pass | iDEN 800 |  | Candy Bar PDA | 2007 | 27.5 |  |  |  |  | 225.6 | 47.0668 | 3 |  |  |  |  | 0.517 | -5.7302 | 4 | 3 | -3.9332 | -6.3302 | OK |  |  |  | 436.3636
207 |  | iDEN 900 |  |  | 2007 | 27.5 |  |  |  |  | 253.6 | 48.083 | 3 |  |  |  |  | 0.546 | -5.2561 | 4 |  | -2.917 | -5.8561 | OK |  |  |  | 464.4689
208 | Pass | CDMA 800 |  | Flip | 2007 | 24.5 |  |  |  |  | 88.2 | 38.9094 | 4 |  |  |  |  | 0.175 | -15.1392 | 4 | 4 | -12.0906 | -15.7392 | OK |  |  |  | 504
208 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 31.6 | 29.9937 | 4 |  |  |  |  | 0.085 | -21.4116 | 4 |  | -11.0063 | -12.0116 | OK |  |  |  | 371.7647
209 | Pass | CDMA 800 |  | Candy Bar PDA | 2007 | 24.5 |  |  |  |  | 78.7 | 37.9195 | 4 |  |  |  |  | 0.178 | -14.9916 | 4 | 4 | -13.0805 | -15.5916 | OK |  |  |  | 442.1348
209 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 49.8 | 33.9446 | 4 |  |  |  |  | 0.142 | -16.9542 | 4 |  | -7.0554 | -7.5542 | OK |  |  |  | 350.7042
300 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 248 | 47.889 | [5 empty] | 0.36 | -8.8739 |  | 2 | -0.61097 | -6.9739 | OK |  |  |  | 688.8889
300 |  | GSM1900 |  |  | 2006 | [5 empty] | 104.5 | 40.3823 | [5 empty] | 0.34 | -9.3704 |  |  | 1.8823 | 2.5296 | Not OK | 0.64725 |  |  | 307.3529
301 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 219.5 | 46.8287 | [5 empty] | 0.281 | -11.0259 |  | 3 | -1.6713 | -9.1259 | OK |  |  |  | 781.1388
301 |  | GSM1900 |  |  | 2006 | [5 empty] | 79.8 | 38.0401 | [5 empty] | 0.155 | -16.1934 |  |  | -0.45994 | -4.2934 | OK |  |  |  | 514.8387
302 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 333 | 50.4489 | [5 empty] | 0.394 | -8.0901 |  | 2 | 1.9489 | -6.1901 | OK |  |  |  | 845.1777
302 |  | GSM1900 |  |  | 2006 | [5 empty] | 133.5 | 42.5096 | [5 empty] | 0.376 | -8.4962 |  |  | 4.0096 | 3.4038 | OK |  |  |  | 355.0532
303 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 273.3 | 48.7328 | [5 empty] | 0.452 | -6.8972 |  | 2 | 0.23279 | -4.9972 | OK |  |  |  | 604.646
303 |  | GSM1900 |  |  | 2006 | [5 empty] | 106.8 | 40.5714 | [5 empty] | 0.401 | -7.9371 |  |  | 2.0714 | 3.9629 | Not OK | 1.8915 |  |  | 266.3342
304 |  | GSM 850 |  | Monoblock | 2006
304 |  | GSM1900 |  |  | 2006
305 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 256.7 | 48.1885 | [5 empty] | 0.378 | -8.4502 |  | 2 | -0.31148 | -6.5502 | OK |  |  |  | 679.1005
305 |  | GSM1900 |  |  | 2006 | [5 empty] | 118.3 | 41.4597 | [5 empty] | 0.34 | -9.3704 |  |  | 2.9597 | 2.5296 | OK |  |  |  | 347.9412
306 |  | GSM 850 |  | Slide | 2006
306 |  | GSM1900 |  |  | 2006
307 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 302.8 | 49.6231 | [5 empty] | 0.401 | -7.9371 |  | 2 | 1.1231 | -6.0371 | OK |  |  |  | 755.1122
307 |  | GSM1900 |  |  | 2006 | [5 empty] | 119.3 | 41.5328 | [5 empty] | 0.289 | -10.782 |  |  | 3.0328 | 1.118 | OK |  |  |  | 412.8028
308 |  | GSM 850 |  | Slide | 2006
308 |  | GSM1900 |  |  | 2006
309 |  | GSM 850 |  | Slide | 2006
309 |  | GSM1900 |  |  | 2006
310 |  | GSM 850 |  | Slide | 2006
310 |  | GSM1900 |  |  | 2006
311 |  | GSM 850 |  | Slide | 2006
311 |  | GSM1900 |  |  | 2006 | [5 empty] | 101.4 | 40.1208 | [5 empty] | 0.164 | -15.7031 |  |  | 1.6208 | -3.8031 | OK |  |  |  | 618.2927
312 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 141.1 | 42.9905 | [5 empty] | 0.225 | -12.9563 |  | 2 | -5.5095 | -11.0563 | OK |  |  |  | 627.1111
312 |  | GSM1900 |  |  | 2007 | [5 empty] | 117.5 | 41.4008 | [5 empty] | 0.288 | -10.8122 |  |  | 2.9008 | 1.0878 | OK |  |  |  | 407.9861
313 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 219.2 | 46.8168 | [5 empty] | 0.304 | -10.3425 |  | 2 | -1.6832 | -8.4425 | OK |  |  |  | 721.0526
313 |  | GSM1900 |  |  | 2006 | [5 empty] | 107.6 | 40.6362 | [5 empty] | 0.218 | -13.2309 |  |  | 2.1362 | -1.3309 | OK |  |  |  | 493.578
314 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 155.7 | 43.8458 | [5 empty] | 0.446 | -7.0133 |  | 2 | -4.6542 | -5.1133 | OK |  |  |  | 349.1031
314 |  | GSM1900 |  |  | 2007 | [5 empty] | 105.4 | 40.4568 | [5 empty] | 0.317 | -9.9788 |  |  | 1.9568 | 1.9212 | OK |  |  |  | 332.4921
315 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 179.5 | 45.0813 | [5 empty] | 0.257 | -11.8013 |  | 2 | -3.4187 | -9.9013 | OK |  |  |  | 698.4436
315 |  | GSM1900 |  |  | 2006 | [5 empty] | 109.6 | 40.7962 | [5 empty] | 0.335 | -9.4991 |  |  | 2.2962 | 2.4009 | Not OK | 0.10469 |  |  | 327.1642
316 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 95.5 | 39.6001 | [5 empty] | 0.268 | -11.4373 |  | 2 | -8.8999 | -9.5373 | OK |  |  |  | 356.3433
316 |  | GSM1900 |  |  | 2006 | [5 empty] | 139.2 | 42.8728 | [5 empty] | 0.211 | -13.5144 |  |  | 4.3728 | -1.6144 | OK |  |  |  | 659.7156
317 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 213.2 | 46.5757 | [5 empty] | 0.471 | -6.5396 |  | 2 | -1.9243 | -4.6396 | OK |  |  |  | 452.6539
317 |  | GSM1900 |  |  | 2006 | [5 empty] | 138.6 | 42.8353 | [5 empty] | 0.45 | -6.9357 |  |  | 4.3353 | 4.9643 | Not OK | 0.62899 |  |  | 308
318 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 176.4 | 44.93 | [5 empty] | 0.351 | -9.0939 |  | 3 | -3.57 | -7.1939 | OK |  |  |  | 502.5641
318 |  | GSM1900 |  |  | 2007 | [5 empty] | 49 | 33.8039 | [5 empty] | 0.172 | -15.2894 |  |  | -4.6961 | -3.3894 | Not OK | 1.3066 |  |  | 284.8837
319 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 265.3 | 48.4747 | [5 empty] | 0.285 | -10.9031 |  | 1 | -0.025255 | -9.0031 | OK |  |  |  | 930.8772
319 |  | GSM1900 |  |  | 2006 | [5 empty] | 171 | 44.6599 | [5 empty] | 0.364 | -8.778 |  |  | 6.1599 | 3.122 | OK | FALSE |  |  | 469.7802
320 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 208.2 | 46.3696 | [5 empty] | 0.241 | -12.3597 |  | 2 | -2.1304 | -10.4597 | OK |  |  |  | 863.9004
320 |  | GSM1900 |  |  | 2006 | [5 empty] | 108.8 | 40.7326 | [5 empty] | 0.26 | -11.7005 |  |  | 2.2326 | 0.19947 | OK |  |  |  | 418.4615
321 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 196.1 | 45.8496 | [5 empty] | 0.301 | -10.4287 |  | 2 | -2.6504 | -8.5287 | OK |  |  |  | 651.495
321 |  | GSM1900 |  |  | 2006 | [5 empty] | 91.1 | 39.1904 | [5 empty] | 0.242 | -12.3237 |  |  | 0.69037 | -0.42369 | OK |  |  |  | 376.4463
322 |  | GSM 850 |  | monoblock | 2006 | [5 empty] | 136.9 | 42.7281 | [5 empty] | 0.43 | -7.3306 |  | 3 | -5.7719 | -5.4306 | Not OK | 0.3413 |  |  | 318.3721
322 |  | GSM1900 |  |  | 2006 | [5 empty] | 73.9 | 37.3729 | [5 empty] | 0.213 | -13.4324 |  |  | -1.1271 | -1.5324 | OK |  |  |  | 346.9484
323 |  | GSM 850 |  | monoblock | 2006 | [5 empty] | 153.7 | 43.7335 | [5 empty] | 0.518 | -5.7134 |  | 2 | -4.7665 | -3.8134 | Not OK | 0.95312 |  |  | 296.7181
323 |  | GSM1900 |  |  | 2006 | [5 empty] | 86.8 | 38.7704 | [5 empty] | 0.278 | -11.1191 |  |  | 0.27039 | 0.7809 | Not OK | 0.5105 |  |  | 312.2302
324 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 205.5 | 46.2562 | [5 empty] | 0.644 | -3.8223 |  | 3 | -2.2438 | -1.9223 | Not OK | 0.32148 |  |  | 319.0994
324 |  | GSM1900 |  |  | 2006 | [5 empty] | 72 | 37.1466 | [5 empty] | 0.179 | -14.9429 |  |  | -1.3534 | -3.0429 | OK |  |  |  | 402.2346
325 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 263.4 | 48.4123 | [5 empty] | 0.297 | -10.5449 |  | 2 | -0.087685 | -8.6449 | OK |  |  |  | 886.8687
325 |  | GSM1900 |  |  | 2006 | [5 empty] | 124.5 | 41.9034 | [5 empty] | 0.308 | -10.229 |  |  | 3.4034 | 1.671 | OK |  |  |  | 404.2208
326 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 209.2 | 46.4112 | [5 empty] | 0.273 | -11.2767 |  | 2 | -2.0888 | -9.3767 | OK |  |  |  | 766.3004
326 |  | GSM1900 |  |  | 2006 | [5 empty] | 108.2 | 40.6845 | [5 empty] | 0.317 | -9.9788 |  |  | 2.1845 | 1.9212 | OK |  |  |  | 341.3249
327 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 231.3 | 47.2835 | [5 empty] | 0.314 | -10.0614 |  | 2 | -1.2165 | -8.1614 | OK |  |  |  | 736.6242
327 |  | GSM1900 |  |  | 2006 | [5 empty] | 86.7 | 38.7604 | [5 empty] | 0.24 | -12.3958 |  |  | 0.26038 | -0.49578 | OK |  |  |  | 361.25
328 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 197.8 | 45.9245 | [5 empty] | 0.281 | -11.0259 |  | 2 | -2.5755 | -9.1259 | OK |  |  |  | 703.9146
328 |  | GSM1900 |  |  | 2006 | [5 empty] | 95.6 | 39.6092 | [5 empty] | 0.24 | -12.3958 |  |  | 1.1092 | -0.49578 | OK |  |  |  | 398.3333
329 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 209.5 | 46.4237 | [5 empty] | 0.417 | -7.5973 |  | 3 | -2.0763 | -5.6973 | OK |  |  |  | 502.3981
329 |  | GSM1900 |  |  | 2006 | [5 empty] | 77.1 | 37.7411 | [5 empty] | 0.221 | -13.1122 |  |  | -0.75891 | -1.2122 | OK |  |  |  | 348.8688
330 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 243.5 | 47.73 | [5 empty] | 0.331 | -9.6034 |  | 3 | -0.77002 | -7.7034 | OK |  |  |  | 735.6495
330 |  | GSM1900 |  |  | 2006 | [5 empty] | 79.7 | 38.0292 | [5 empty] | 0.185 | -14.6566 |  |  | -0.47083 | -2.7566 | OK |  |  |  | 430.8108
331 |  | GSM 850 |  | Fold | 2006
331 |  | GSM1900 |  |  | 2006 | [5 empty] | 83.2 | 38.4025 | [5 empty] | 0.168 | -15.4938 |  | 3 | -0.097533 | -3.5938 | OK |  |  |  | 495.2381
332 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 252.5 | 48.0452 | [5 empty] | 0.378 | -8.4502 |  | 3 | -0.45477 | -6.5502 | OK |  |  |  | 667.9894
332 |  | GSM1900 |  |  | 2006 | [5 empty] | 83.2 | 38.4025 | [5 empty] | 0.168 | -15.4938 |  |  | -0.097533 | -3.5938 | OK |  |  |  | 495.2381
333 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 200.5 | 46.0423 | [5 empty] | 0.28 | -11.0568 |  | 3 | -2.4577 | -9.1568 | OK |  |  |  | 716.0714
333 |  | GSM1900 |  |  | 2006 | [5 empty] | 73.4 | 37.3139 | [5 empty] | 0.206 | -13.7227 |  |  | -1.1861 | -1.8227 | OK |  |  |  | 356.3107
334 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 210.9 | 46.4815 | [5 empty] | 0.296 | -10.5742 |  | 3 | -2.0185 | -8.6742 | OK |  |  |  | 712.5
334 |  | GSM1900 |  |  | 2006 | [5 empty] | 78.4 | 37.8863 | [5 empty] | 0.215 | -13.3512 |  |  | -0.61368 | -1.4512 | OK |  |  |  | 364.6512
335 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 178.7 | 45.0425 | [5 empty] | 0.54 | -5.3521 |  | 3 | -3.4575 | -3.4521 | Not OK | 0.0053841 |  |  | 330.9259
335 |  | GSM1900 |  |  | 2006 | [5 empty] | 92.3 | 39.304 | [5 empty] | 0.39 | -8.1787 |  |  | 0.80403 | 3.7213 | Not OK | 2.9173 |  |  | 236.6667
336 |  | GSM 850 |  | Communicator | 2006 | [5 empty] | 211.4 | 46.5021 | [5 empty] | 0.29 | -10.752 |  | 2 | -1.9979 | -8.852 | OK |  |  |  | 728.9655
336 |  | GSM1900 |  |  | 2006 | [5 empty] | 96.5 | 39.6905 | [5 empty] | 0.27 | -11.3727 |  |  | 1.1905 | 0.52728 | OK | FALSE |  |  | 357.4074
337 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 77.7 | 37.8084 | [5 empty] | 0.13 | -17.7211 |  | 2 | -10.6916 | -15.8211 | OK |  |  |  | 597.6923
337 |  | GSM1900 |  |  | 2006 | [5 empty] | 65.2 | 36.285 | [5 empty] | 0.29 | -10.752 |  |  | -2.215 | 1.148 | Not OK | 3.363 |  |  | 224.8276
338 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 143.5 | 43.137 | [5 empty] | 0.41 | -7.7443 |  | 2 | -5.363 | -5.8443 | OK |  |  |  | 350
338 |  | GSM1900 |  |  | 2006 | [5 empty] | 97.9 | 39.8157 | [5 empty] | 0.29 | -10.752 |  |  | 1.3157 | 1.148 | OK |  |  |  | 337.5862
339 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 207.3 | 46.332 | [5 empty] | 0.53 | -5.5145 |  | 1 | -2.168 | -3.6145 | OK |  |  |  | 391.1321
339 |  | GSM1900 |  |  | 2006 | [5 empty] | 112.3 | 41.0076 | [5 empty] | 0.47 | -6.558 |  |  | 2.5076 | 5.342 | Not OK | 2.8344 |  |  | 238.9362
340 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 246.4 | 47.8328 | [5 empty] | 0.51 | -5.8486 |  | 2 | -0.66719 | -3.9486 | OK |  |  |  | 483.1373
340 |  | GSM1900 |  |  | 2006 | [5 empty] | 118.8 | 41.4963 | [5 empty] | 0.41 | -7.7443 |  |  | 2.9963 | 4.1557 | Not OK | 1.1593 |  |  | 289.7561
341 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 206.8 | 46.311 | [5 empty] | 0.62 | -4.1522 |  | 2 | -2.189 | -2.2522 | OK |  |  |  | 333.5484
341 |  | GSM1900 |  |  | 2006 | [5 empty] | 93.5 | 39.4162 | [5 empty] | 0.44 | -7.1309 |  |  | 0.91623 | 4.7691 | Not OK | 3.8528 |  |  | 212.5
342 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 279 | 48.9121 | [5 empty] | 0.326 | -9.7356 |  | 1 | 0.41208 | -7.8356 | OK |  |  |  | 855.8282
342 |  | GSM1900 |  |  | 2006 | [5 empty] | 150 | 43.5218 | [5 empty] | 0.649 | -3.7551 |  |  | 5.0218 | 8.1449 | Not OK | 3.1231 |  |  | 231.1248
343 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 199.8 | 46.0119 | [8 empty] | 2
343 |  | GSM1900 |  |  | 2006 | [5 empty] | 85.8 | 38.6697
344 |  | GSM 850 |  | Monoblock | 2006
344 |  | GSM1900 |  |  | 2006
345 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 187.9 | 45.4785 | [5 empty] | 0.555 | -5.1141 |  | 1 | 6.9785 | 6.7859 | OK |  |  |  | 338.5586
345 |  | GSM1900 |  |  | 2006 | [5 empty] | 130.8 | 42.3322 | [5 empty] | 1.15 | 1.214 |  |  | -6.1678 | 3.114 | Not OK | 9.2818 |  |  | 113.7391
346 |  | GSM 850 |  | Slide | 2006 | [5 empty] | 337 | 50.5526 | [5 empty] | 0.423 | -7.4732 |  | 0 | 12.0526 | 4.4268 | OK |  |  |  | 796.6903
346 |  | GSM1900 |  |  | 2006
347 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 195.2 | 45.8096 | [5 empty] | 0.61 | -4.2934 |  | 1 | -2.6904 | -2.3934 | Not OK | 0.297 |  |  | 320
347 |  | GSM1900 |  |  | 2006 | [5 empty] | 102.5 | 40.2145 | [5 empty] | 0.542 | -5.32 |  |  | 1.7145 | 6.58 | Not OK | 4.8655 |  |  | 189.1144
348 |  | GSM 850 |  | Slide | 2007 | [5 empty] | 235.7 | 47.4472 | [5 empty] | 0.314 | -10.0614 |  | 2 | -1.0528 | -8.1614 | OK |  |  |  | 750.6369
348 |  | GSM1900 |  |  | 2007 | [5 empty] | 130.1 | 42.2855 | [5 empty] | 0.385 | -8.2908 |  |  | 3.7855 | 3.6092 | OK | FALSE |  |  | 337.9221
349 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 115.4 | 41.2441 | [5 empty] | 0.214 | -13.3917 |  | 3 | -7.2559 | -11.4917 | OK |  |  |  | 539.2523
349 |  | GSM1900 |  |  | 2006 | [5 empty] | 56.3 | 35.0102 | [5 empty] | 0.109 | -19.2515 |  |  | -3.4898 | -7.3515 | OK |  |  |  | 516.5138
350 |  | GSM 850 |  | Fold | 2006 | [16 empty] | -48.5 | 1.9 | Not OK |  |  |  |  | 50.4
350 |  | GSM1900 |  |  | 2006 | [16 empty] | -38.5 | 11.9 | Not OK |  |  |  |  | 50.4
351 |  | GSM 850 |  | Fold | 2006 | [16 empty] | -48.5 | 1.9 | Not OK |  |  |  |  | 50.4
351 |  | GSM1900 |  |  | 2006 | [16 empty] | -38.5 | 11.9 | Not OK |  |  |  |  | 50.4
352 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 280.2 | 48.9494 | [5 empty] | 0.399 | -7.9805 |  | 2 | 0.44936 | -6.0805 | OK |  |  |  | 702.2556 | FALSE
352 |  | GSM1900 |  |  | 2006 | [5 empty] | 135.1 | 42.6131 | [5 empty] | 0.361 | -8.8499 |  |  | 4.1131 | 3.0501 | OK |  |  |  | 374.2382 | FALSE
353 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 229.6 | 47.2194 | [5 empty] | 0.73 | -2.7335 |  | 2 | -1.2806 | -0.83354 | Not OK | 0.44702 |  |  | 314.5205
353 |  | GSM1900 |  |  | 2006 | [5 empty] | 95 | 39.5545 | [5 empty] | 0.34 | -9.3704 |  |  | 1.0545 | 2.5296 | Not OK | 1.4751 |  |  | 279.4118
354 |  | GSM 850 |  | Slide | 2006
354 |  | GSM1900 |  |  | 2006 | [5 empty] | 75.3 | 37.5359 | [5 empty] | 0.176 | -15.0897 |  | 2 | -0.9641 | -3.1897 | OK |  |  |  | 427.8409
355 |  | GSM 850 |  | Slide | 2006
355 |  | GSM1900 |  |  | 2006
356 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 271 | 48.6594 | [5 empty] | 0.5 | -6.0206 |  | 2 | 0.15939 | -4.1206 | OK |  |  |  | 542
356 |  | GSM1900 |  |  | 2006 | [5 empty] | 74.5 | 37.4431 | [5 empty] | 0.2 | -13.9794 |  |  | -1.0569 | -2.0794 | OK |  |  |  | 372.5
357 |  | GSM 850 |  | Swivel | 2006 | [5 empty] | 185.6 | 45.3716 | [5 empty] | 0.296 | -10.5742 |  | 2 | -3.1284 | -8.6742 | OK |  |  |  | 627.027
357 |  | GSM1900 |  |  | 2006 | [5 empty] | 95 | 39.5545 | [5 empty] | 0.16 | -15.9176 |  |  | 1.0545 | -4.0176 | OK |  |  |  | 593.75
358 |  | GSM 850 |  | Slide | 2007 | [5 empty] | 242.7 | 47.7014 | [5 empty] | 0.287 | -10.8424 |  | 2 | -0.7986 | -8.9424 | OK |  |  |  | 845.6446
358 |  | GSM1900 |  |  | 2007 | [5 empty] | 150.3 | 43.5392 | [5 empty] | 0.442 | -7.0916 |  |  | 5.0392 | 4.8084 | OK |  |  |  | 340.0452
359 |  | GSM 850 |  | Monoblock | 2006 | [18 empty] | Not OK |  |  |  |  | 0
359 |  | GSM1900 |  |  | 2006 | [5 empty] | 125.9 | 42.0005 | [5 empty] | 0.343 | -9.2941 |  | 1 | 3.5005 | 2.6059 | OK |  |  |  | 367.0554
360 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 236.2 | 47.4656 | [5 empty] | 0.289 | -10.782 |  | 2 | -1.0344 | -8.882 | OK |  |  |  | 817.301
360 |  | GSM1900 |  |  | 2006 | [5 empty] | 142.2 | 43.058 | [5 empty] | 0.366 | -8.7304 |  |  | 2.3436 | 3.1696 | Not OK | 0.82599 |  |  | 388.5246
361 |  | GSM 850 |  | Monoblock | 2007 | [18 empty] | Not OK
361 |  | GSM1900 |  |  | 2007 | [5 empty] | 110.2 | 40.8436 | [5 empty] | 0.32 | -9.897 |  | 2 | -0.24493 | 2.003 | Not OK | 2.2479 |  |  | 344.375
362 |  | GSM 850 |  | Fold | 2007 | [5 empty] | 220.2 | 46.8563 | [5 empty] | 0.308 | -10.229 |  | 3 | -1.6437 | -8.329 | OK |  |  |  | 714.9351
362 |  | GSM1900 |  |  | 2007 | [5 empty] | 81.8 | 38.2551 | [5 empty] | 0.272 | -11.3086 |  |  | -0.24493 | 0.59138 | Not OK | 0.83631 |  |  | 300.7353
363 |  | GSM 850 |  | Slide | 2007 | [5 empty] | 229.8 | 47.227 | [5 empty] | 0.339 | -9.396 |  | 2 | -1.273 | -7.496 | OK |  |  |  | 677.8761
363 |  | GSM1900 |  |  | 2007 | [5 empty] | 134.8 | 42.5938 | [5 empty] | 0.428 | -7.3711 |  |  | 4.0938 | 4.5289 | Not OK | 0.43508 |  |  | 314.9533
364 |  | GSM 850 |  | Block | 2007 | [5 empty] | 192.3 | 45.6796 | [5 empty] | 0.328 | -9.6825 |  | 2 | -2.8204 | -7.7825 | OK |  |  |  | 586.2805
364 |  | GSM1900 |  |  | 2007 | [5 empty] | 93.2 | 39.3883 | [5 empty] | 0.305 | -10.314 |  |  | 0.88832 | 1.586 | Not OK | 0.69768 |  |  | 305.5738
365 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 186.2 | 45.3996 | [5 empty] | 0.801 | -1.9273 |  | 2 | -3.1004 | -0.02735 | Not OK | 3.0731 |  |  | 232.4594
365 |  | GSM1900 |  |  | 2006 | [5 empty] | 133.4 | 42.5031 | [5 empty] | 0.548 | -5.2244 |  |  | 4.0031 | 6.6756 | Not OK | 2.6725 |  |  | 243.4307
366 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 186.4 | 45.4089 | [5 empty] | 0.227 | -12.8795 |  | 2 | -3.0911 | -10.9795 | OK |  |  |  | 821.1454
366 |  | GSM1900 |  |  | 2006 | [5 empty] | 55.5 | 34.8859 | [5 empty] | 0.203 | -13.8501 |  |  | -3.6141 | -1.9501 | Not OK | 1.6641 |  |  | 273.399
367 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 279.2 | 48.9183 | [5 empty] | 0.337 | -9.4474 |  | 3 | 0.41831 | -7.5474 | OK |  |  |  | 828.4866
367 |  | GSM1900 |  |  | 2006 | [5 empty] | 110.7 | 40.883 | [5 empty] | 0.393 | -8.1121 |  |  | 2.383 | 3.7879 | Not OK | 1.4049 |  |  | 281.6794
368 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 151.4 | 43.6025 | [5 empty] | 0.23 | -12.7654 |  | 2 | -4.8975 | -10.8654 | OK |  |  |  | 658.2609
368 |  | GSM1900 |  |  | 2006 | [5 empty] | 133 | 42.477 | [5 empty] | 0.459 | -6.7637 |  |  | 3.977 | 5.1363 | Not OK | 1.1592 |  |  | 289.7603
369 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 278.9 | 48.909 | [5 empty] | 0.342 | -9.3195 |  | 2 | 0.40897 | -7.4195 | OK |  |  |  | 815.4971
369 |  | GSM1900 |  |  | 2006 | [5 empty] | 144.1 | 43.1733 | [5 empty] | 0.505 | -5.9342 |  |  | 4.6733 | 5.9658 | Not OK | 1.2925 |  |  | 285.3465
370 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 188.9 | 45.5246 | [5 empty] | 0.915 | -0.77158 |  | 1 | -2.9754 | 1.1284 | Not OK | 4.1038 |  |  | 206.4481
370 |  | GSM1900 |  |  | 2006 | [5 empty] | 155.2 | 43.8178 | [5 empty] | 0.662 | -3.5828 |  |  | 5.3178 | 8.3172 | Not OK | 2.9993 |  |  | 234.4411
371 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 187 | 45.4368 | [5 empty] | 0.794 | -2.0036 |  | 1 | -3.0632 | -0.10359 | Not OK | 2.9596 |  |  | 235.5164
371 |  | GSM1900 |  |  | 2006 | [5 empty] | 149.6 | 43.4986 | [5 empty] | 0.617 | -4.1943 |  |  | 4.9986 | 7.7057 | Not OK | 2.7071 |  |  | 242.4635
372 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 199.6 | 46.0032 | [5 empty] | 0.822 | -1.7026 |  | 1 | -2.4968 | 0.19744 | Not OK | 2.6942 |  |  | 242.8224
372 |  | GSM1900 |  |  | 2006 | [5 empty] | 151.1 | 43.5853 | [5 empty] | 0.581 | -4.7165 |  |  | 5.0853 | 7.1835 | Not OK | 2.0982 |  |  | 260.0688
373 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 210.2 | 46.4527 | [5 empty] | 0.896 | -0.95384 |  | 1 | -2.0473 | 0.94616 | Not OK | 2.9935 |  |  | 234.5982
373 |  | GSM1900 |  |  | 2006 | [5 empty] | 136.1 | 42.6772 | [5 empty] | 0.601 | -4.4225 |  |  | 4.1772 | 7.4775 | Not OK | 3.3003 |  |  | 226.4559
374 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 212 | 46.5267 | [5 empty] | 1.04 | 0.34067 |  | 1 | -1.9733 | 2.2407 | Not OK | 4.2139 |  |  | 203.8462
374 |  | GSM1900 |  |  | 2006 | [5 empty] | 164.6 | 44.3286 | [5 empty] | 0.614 | -4.2366 |  |  | 5.8286 | 7.6634 | Not OK | 1.8348 |  |  | 268.0782
375 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 217.3 | 46.7412 | [5 empty] | 1.02 | 0.172 |  |  | -1.7588 | 2.072 | Not OK | 3.8308 |  |  | 213.0392
375 |  | GSM1900 |  |  | 2006 | [5 empty] | 133.1 | 42.4836 | [5 empty] | 0.479 | -6.3933 |  |  | 3.9836 | 5.5067 | Not OK | 1.5231 |  |  | 277.8706
376 |  | GSM 850 |  | slide | 2007 | [5 empty] | 288.7 | 49.2089 | [5 empty] | 0.357 | -8.9466 |  | 2 | 0.70894 | -7.0466 | OK |  |  |  | 808.6835
376 |  | GSM1900 |  |  | 2007 | [5 empty] | 112.1 | 40.9921 | [5 empty] | 0.29 | -10.752 |  |  | 2.4921 | 1.148 | OK |  |  |  | 386.5517
377 |  | GSM 850 |  | slide | 2007 | [5 empty] | 192.8 | 45.7021 | [5 empty] | 0.239 | -12.432 |  | 2 | -2.7979 | -10.532 | OK |  |  |  | 806.6946
377 |  | GSM1900 |  |  | 2007 | [5 empty] | 86.7 | 38.7604 | [5 empty] | 0.23 | -12.7654 |  |  | 0.26038 | -0.86544 | OK |  |  |  | 376.9565
378 |  | GSM 850 |  | slide | 2007 | [5 empty] | 267.7 | 48.553 | [5 empty] | 0.342 | -9.3195 |  | 2 | 0.052967 | -7.4195 | OK |  |  |  | 782.7485
378 |  | GSM1900 |  |  | 2007 | [5 empty] | 102.4 | 40.206 | [5 empty] | 0.223 | -13.0339 |  |  | 1.706 | -1.1339 | OK |  |  |  | 459.1928
379 |  | GSM 850 |  | slide | 2007 | [5 empty] | 154.6 | 43.7842 | [5 empty] | 0.232 | -12.6902 |  | 2 | -4.7158 | -10.7902 | OK |  |  |  | 666.3793
379 |  | GSM1900 |  |  | 2007 | [5 empty] | 88.6 | 38.9487 | [5 empty] | 0.223 | -13.0339 |  |  | 0.44867 | -1.1339 | OK |  |  |  | 397.3094
380 |  | GSM 850 |  | monoblock | 2007 | [18 empty] | Not OK |  |  |  |  | 0
380 |  | GSM1900 |  |  | 2007 | [5 empty] | 114.8 | 41.1988 | [5 empty] | 0.326 | -9.7356 |  | 2 | 2.6988 | 2.1644 | OK |  |  |  | 352.1472
381 |  | GSM 850 |  | Dual slider | 2007 | [5 empty] | 175.1 | 44.8657 | [5 empty] | 0.31 | -10.1728 |  | 2 | -3.6343 | -8.2728 | OK |  |  |  | 564.8387
381 |  | GSM1900 |  |  | 2007 | [5 empty] | 85 | 38.5884 | [5 empty] | 0.239 | -12.432 |  |  | 0.088379 | -0.53204 | OK |  |  |  | 355.6485
382 |  | GSM 850 |  | Dual slider | 2007 | [5 empty] | 206.8 | 46.311 | [5 empty] | 0.235 | -12.5786 |  | 3 | -2.189 | -10.6786 | OK |  |  |  | 880
382 |  | GSM1900 |  |  | 2007 | [5 empty] | 83.5 | 38.4337 | [5 empty] | 0.152 | -16.3631 |  |  | -0.06627 | -4.4631 | OK |  |  |  | 549.3421
383 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 262.4 | 48.3793 | [5 empty] | 248.6 | 47.91 |  | 3 | -0.12072 | 49.81 | Not OK |  |  |  | 1.0555 | 49.9307
383 |  | GSM1900 |  |  | 2006 | [5 empty] | 74.2 | 37.4081 | [5 empty] | 0.147 | -16.6537 |  |  | -1.0919 | -4.7537 | OK |  |  |  | 504.7619
384 |  | GSM 850 |  | Fold | 2006 | [5 empty] | 256.9 | 48.1953 | [5 empty] | 0.346 | -9.2185 |  | 2 | -0.30472 | -7.3185 | OK |  |  |  | 742.4855
384 |  | GSM1900 |  |  | 2006 | [5 empty] | 94 | 39.4626 | [5 empty] | 0.263 | -11.6009 |  |  | 0.96256 | 0.29911 | OK |  |  |  | 357.4144
385 |  | GSM 850 |  | Fold | 2007 | [5 empty] | 172.5 | 44.7358 | [5 empty] | 0.318 | -9.9515 |  | 2 | -3.7642 | -8.0515 | OK |  |  |  | 542.4528
385 |  | GSM1900 |  |  | 2007 | [5 empty] | 76.6 | 37.6846 | [5 empty] | 0.293 | -10.6626 |  |  | -0.81542 | 1.2374 | Not OK | 2.0528 |  |  | 261.4334
386 |  | GSM 850 |  | Monoblock | 2006 | [5 empty] | 215.7 | 46.677 | [5 empty] | 0.341 | -9.3449 |  | 2 | -1.823 | -7.4449 | OK |  |  |  | 632.5513
386 |  | GSM1900 |  |  | 2006 | [5 empty] | 91.8 | 39.2569 | [5 empty] | 0.251 | -12.0065 |  |  | 0.75685 | -0.10653 | OK |  |  |  | 365.7371
387 |  | GSM 850 |  | Fold | 2007 | [5 empty] | 245.5 | 47.801 | [5 empty] | 0.326 | -9.7356 |  | 3 | -0.69897 | -7.8356 | OK |  |  |  | 753.0675
387 |  | GSM1900 |  |  | 2007 | [5 empty] | 77.6 | 37.7972 | [5 empty] | 0.206 | -13.7227 |  |  | -0.70277 | -1.8227 | OK |  |  |  | 376.699
388 |  | GSM 850 |  | Fold | 2007 | [5 empty] | 913.1 | 59.2104 | [5 empty] | 0.353 | -9.0445 |  | 2 | 10.7104 | -7.1445 | OK |  |  |  | 2586.6856
388 |  | GSM1900 |  |  | 2007 | [5 empty] | 65.6 | 36.3381 | [5 empty] | 0.262 | -11.634 |  |  | -2.1619 | 0.26603 | Not OK | 2.4279 |  |  | 250.3817
389 |  | GSM 850 |  | Dual slider | 2007 | [5 empty] | 263.4 | 48.4123 | [5 empty] | 0.328 | -9.6825 |  | 2 | -0.087685 | -7.7825 | OK |  |  |  | 803.0488
389 |  | GSM1900 |  |  | 2007 | [5 empty] | 102.3 | 40.1975 | [5 empty] | 0.193 | -14.2889 |  |  | 1.6975 | -2.3889 | OK |  |  |  | 530.0518
390 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 152.7 | 43.6768 | [5 empty] | 0.247 | -12.1461 |  | 2 | -4.8232 | -10.2461 | OK |  |  |  | 618.2186
390 |  | GSM1900 |  |  | 2007 | [5 empty] | 117.7 | 41.4155 | [5 empty] | 0.297 | -10.5449 |  |  | 2.9155 | 1.3551 | OK |  |  |  | 396.2963
391 |  | GSM 850 |  | Slide | 2007 | [5 empty] | 208.4 | 46.378 | [5 empty] | 0.299 | -10.4866 |  | 2 | -2.122 | -8.5866 | OK |  |  |  | 696.99
391 |  | GSM1900 |  |  | 2007 | [5 empty] | 131.8 | 42.3983 | [5 empty] | 0.342 | -9.3195 |  |  | 3.8983 | 2.5805 | OK |  |  |  | 385.3801
392 |  | GSM 850 |  | Slide | 2007 | [5 empty] | 147.7 | 43.3876 | [5 empty] | 0.243 | -12.2879 |  | 2 | -5.1124 | -10.3879 | OK |  |  |  | 607.8189
392 |  | GSM1900 |  |  | 2007 | [5 empty] | 85 | 38.5884 | [5 empty] | 0.231 | -12.7278 |  |  | 0.088379 | -0.82776 | OK |  |  |  | 367.9654
393 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 129.8 | 42.2655 | [5 empty] | 0.395 | -8.0681 |  | 2 | -6.2345 | -6.1681 | Not OK | 0.066448 |  |  | 328.6076
393 |  | GSM1900 |  |  | 2007 | [5 empty] | 86.3 | 38.7202 | [5 empty] | 0.246 | -12.1813 |  |  | 0.22022 | -0.2813 | OK |  |  |  | 350.813
394 |  | GSM 850 |  | side slider | 2007 | [5 empty] | 108.6 | 40.7166 | [5 empty] | 0.253 | -11.9376 |  | 3 | -7.7834 | -10.0376 | OK |  |  |  | 429.249
394 |  | GSM1900 |  |  | 2007 | [5 empty] | 52.7 | 34.4362 | [5 empty] | 0.141 | -17.0156 |  |  | -4.0638 | -5.1156 | OK |  |  |  | 373.7589
395 |  | GSM 850 |  | side slider | 2007 | [5 empty] | 227.4 | 47.1358 | [5 empty] | 0.326 | -9.7356 |  | 3 | -1.3642 | -7.8356 | OK |  |  |  | 697.546
395 |  | GSM1900 |  |  | 2007 | [5 empty] | 124.3 | 41.8894 | [5 empty] | 0.414 | -7.66 |  |  | 3.3894 | 4.24 | Not OK | 0.85058 |  |  | 300.2415
396 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 125.8 | 41.9936 | [5 empty] | 0.176 | -15.0897 |  | 2 | -6.5064 | -13.1897 | OK |  |  |  | 714.7727
396 |  | GSM1900 |  |  | 2007 | [5 empty] | 85.4 | 38.6292 | [5 empty] | 0.195 | -14.1993 |  |  | 0.12916 | -2.2993 | OK |  |  |  | 437.9487
397 |  | GSM 850 |  | Side slider | 2007 | [5 empty] | 218.9 | 46.8049 | [5 empty] | 0.292 | -10.6923 |  | 2 | -1.6951 | -8.7923 | OK |  |  |  | 749.6575
397 |  | GSM1900 |  |  | 2007 | [5 empty] | 65.6 | 36.3381 | [5 empty] | 0.176 | -15.0897 |  |  | -2.1619 | -3.1897 | OK |  |  |  | 372.7273
398 |  | GSM 850 |  | Fold | 2007 | [5 empty] | 201.3 | 46.0769 | [5 empty] | 0.358 | -8.9223 |  | 2 | -2.4231 | -7.0223 | OK |  |  |  | 562.2905
398 |  | GSM1900 |  |  | 2007 | [5 empty] | 68.5 | 36.7138 | [5 empty] | 0.243 | -12.2879 |  |  | -1.7862 | -0.38787 | Not OK | 1.3983 |  |  | 281.893
399 |  | GSM 850 |  | Fold | 2007 | [5 empty] | 234.5 | 47.4029 | [5 empty] | 0.398 | -8.0023 |  | 2 | -1.0971 | -6.1023 | OK |  |  |  | 589.196
399 |  | GSM1900 |  |  | 2007 | [5 empty] | 106.2 | 40.5225 | [5 empty] | 0.238 | -12.4685 |  |  | 2.0225 | -0.56846 | OK |  |  |  | 446.2185
400 |  | GSM 850 |  | Side Slide | 2007 | [5 empty] | 210.6 | 46.4692 | [5 empty] | 0.307 | -10.2572 |  | 3 | -2.0308 | -8.3572 | OK |  |  |  | 685.9935
400 |  | GSM1900 |  |  | 2007 | [5 empty] | 67.8 | 36.6246 | [5 empty] | 0.223 | -13.0339 |  |  | -1.8754 | -1.1339 | Not OK | 0.7415 |  |  | 304.0359
401 |  | GSM 850 |  | Fold | 2007 | [5 empty] | 213 | 46.5676 | [5 empty] | 0.295 | -10.6036 |  | 3 | -1.9324 | -8.7036 | OK |  |  |  | 722.0339
401 |  | GSM1900 |  |  | 2007 | [5 empty] | 68.7 | 36.7391 | [5 empty] | 0.201 | -13.9361 |  |  | -1.7609 | -2.0361 | OK |  |  |  | 341.791
402 |  | GSM 850 |  | Fold | 2007
402 |  | GSM1900 |  |  | 2007
403 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 191.1 | 45.6252 | [5 empty] | 0.462 | -6.7072 |  | 2 | -2.8748 | -4.8072 | OK |  |  |  | 413.6364
403 |  | GSM1900 |  |  | 2007 | [5 empty] | 98.5 | 39.8687 | [5 empty] | 0.231 | -12.7278 |  |  | 1.3687 | -0.82776 | OK |  |  |  | 426.4069
404 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 174.6 | 44.8409 | [5 empty] | 0.044 | -27.1309 |  | 2 | -3.6591 | -25.2309 | OK |  |  |  | 3968.1818
404 |  | GSM1900 |  |  | 2007 | [5 empty] | 134.5 | 42.5744 | [5 empty] | 0.038 | -28.4043 |  |  | 4.0744 | -16.5043 | OK |  |  |  | 3539.4737
405 |  | GSM 850 |  | Monoblock | 2007 | [5 empty] | 116.7 | 41.3414 | [5 empty] | 0.327 | -9.709 |  | 3 | -7.1586 | -7.809 | OK |  |  |  | 356.8807
405 |  | GSM1900 |  |  | 2007 | [5 empty] | 71 | 37.0252 | [5 empty] | 0.204 | -13.8074 |  |  | -1.4748 | -1.9074 | OK |  |  |  | 348.0392

## Data low band
Product Number | HAC Compliant? | Air Interface | Bands | Form Factor | Standard Version | Max Power dBm | Max Avg. E Field (V/m) | Max Avg. E Field dB(V/m) | After Exclusion E field (V/m) | After Exclusion E field dB(V/m) | PMF Scaled E field (V/m) | PMF Scaled E field dB(V/m) | Rating E field | Max Avg. H Field (A/m) | Max Avg. H Field dB(A/m) | After Exclusion H field (A/m) | After Exclusion H field dB(A/m) | PMF Scaled H field (A/m) | PMF Scaled H field dB(A/m) | Rating H field | E H Ratings | Delta (dB) E-Field relative to M3 (pass) | Delta (dB) H-Field relative to M3 (pass) | OK E-Field Dominate | Amount H-Field is Dominate in dB | Corrected Delta (dB) E-Field relative to M3 (pass) | Corrected Delta (dB) H-Field relative to M3 (pass) | E/H ratio (Ohms) | peak dB(V/m) - peak dBm | peak dB(A/m) - peak dBm
1 | YES | iDEN | 800 | Candy Bar | 2007 | 28.06 | 69.97 | 36.8982 | 69.96 | 36.897 | 343.5 | 50.7185 | M3 | 0.1152 | -18.771 | 0.1048 | -19.5928 | 0.456 | -6.8207 | M4 | EM3 - HM4 | -14.1018 | -19.371 | OK |  | -0.28147 | -7.4207 | 753.2895 | 22.6585 | -34.8807
1 |  |  | 900 |  |  | 28.06 | 69.57 | 36.8484 | 69.57 | 36.8484 | 334 | 50.4749 | M3 | 0.109 | -19.2515 | 0.1018 | -19.845 | 0.445 | -7.0328 | M4 | EM3 - HM4 | -14.1516 | -19.8515 | OK |  | -0.52507 | -7.6328 | 750.5618 | 22.4149 | -35.0928
2 | YES | iDEN | 800 | Clam | 2007 | 28.6 | 30.44 | 29.6689 | 30.1 | 29.5713 | 147.8 | 43.3935 | M4 | 0.07575 | -22.4123 | 0.06276 | -24.0463 | 0.273 | -11.2767 | M4 | EM4 - HM4 | -21.3311 | -23.0123 | OK |  | -7.6065 | -11.8767 | 541.3919 | 14.7935 | -39.8767
2 |  |  | 900 |  |  | 28.6 | 18.82 | 25.4924 | 17.375 | 24.7985 | 83.4 | 38.4233 | M4 | 0.06907 | -23.2142 | 0.05515 | -25.1691 | 0.241 | -12.3597 | M4 | EM4 - HM4 | -25.5076 | -23.8142 | Not OK | 1.6934 | -12.5767 | -12.9597 | 346.0581 | 9.8233 | -40.9597
3 | YES | iDEN | 800 | Clam | 2006 | 28.06 | 45.03 | 33.07 | 45.03 | 33.07 | 213.9 | 46.6042 | M3 | 0.08649 | -21.2607 | 0.06137 | -24.2409 | 0.286 | -10.8727 | M4 | EM3 - HM4 | -17.93 | -21.8607 | OK |  | -4.3958 | -11.4727 | 747.9021 | 18.5442 | -38.9327
3 |  |  | 900 |  |  | 28.06 | 30.42 | 29.6632 | 30.41 | 29.6603 | 146.9 | 43.3404 | M4 | 0.07638 | -22.3404 | 0.05668 | -24.9314 | 0.263 | -11.6009 | M4 | EM4 - HM4 | -21.3368 | -22.9404 | OK |  | -7.6596 | -12.2009 | 558.5551 | 15.2804 | -39.6609
4 | YES | CDMA | 850 | Clam | 2006 | 25 | 73.15 | 37.2843 | 73.15 | 37.2843 | 71 | 37.0252 | M3 | 0.1532 | -16.2948 | 0.1032 | -19.7264 | 0.096 | -20.3546 | M4 | EM3 - HM4 | -13.7157 | -16.8948 | OK |  | -13.9748 | -20.9546 | 739.5833 | 12.0252 | -45.3546
5 | YES | CDMA | 850 | Clam | 2006 | 25 | 76.86 | 37.714 | 76.87 | 37.7151 | 76.1 | 37.6277 | M4 | 0.1479 | -16.6006 | 0.1051 | -19.5679 | 0.103 | -19.7433 | M4 | EM4 - HM4 | -13.286 | -17.2006 | OK |  | -13.3723 | -20.3433 | 738.835 | 12.6277 | -44.7433
6 | YES | CDMA | 850 | Clam | 2006 | 25 | 113.5 | 41.0999 | 113.5 | 41.0999 | 126 | 42.0074 | M4 | 0.2116 | -13.4897 | 0.1467 | -16.6714 | 0.157 | -16.082 | M4 | EM4 - HM4 | -9.9001 | -14.0897 | OK |  | -8.9926 | -16.682 | 802.5478 | 17.0074 | -41.082
7 | YES | CDMA | 850 | Candy Bar | 2006 | 25 | 124.5 | 41.9034 | 124.5 | 41.9034 | 123.3 | 41.8193 | M4 | 0.2126 | -13.4487 | 0.15 | -16.4782 | 0.147 | -16.6537 | M4 | EM4 - HM4 | -9.0966 | -14.0487 | OK |  | -9.1807 | -17.2537 | 838.7755 | 16.8193 | -41.6537
8 | YES | CDMA | 850 | Clam | 2006 | 25 | 65.42 | 36.3142 | 65.38 | 36.3089 | 69.3 | 36.8147 | M4 | 0.1569 | -16.0875 | 0.1152 | -18.771 | 0.121 | -18.3443 | M4 | EM4 - HM4 | -14.6858 | -16.6875 | OK |  | -14.1853 | -18.9443 | 572.7273 | 11.8147 | -43.3443
10 | YES | CDMA | 850 | Clam | 2006 | 25 | 95 | 39.5545 | 95 | 39.5545 | 105.4 | 40.4568 | M4 | 0.1991 | -14.0186 | 0.143 | -16.8933 | 0.153 | -16.3062 | M4 | EM4 - HM4 | -11.4455 | -14.6186 | OK |  | -10.5432 | -16.9062 | 688.8889 | 15.4568 | -41.3062
11 | YES | CDMA | 850 | Clam | 2006 | 25 | 105.3 | 40.4486 | 105.3 | 40.4486 | 111.6 | 40.9533 | M4 | 0.3516 | -9.079 | 0.2667 | -11.4795 | 0.28 | -11.0568 | M4 | EM4 - HM4 | -10.5514 | -9.679 | Not OK | 0.87241 | -10.0467 | -11.6568 | 398.5714 | 15.9533 | -36.0568
12 | YES | CDMA | 850 | Clam | 2006 | 25 | 87.18 | 38.8083 | 87.17 | 38.8073 | 92.4 | 39.3134 | M4 | 0.1586 | -15.9939 | 0.1057 | -19.5185 | 0.111 | -19.0935 | M4 | EM4 - HM4 | -12.1917 | -16.5939 | OK |  | -11.6866 | -19.6935 | 832.4324 | 14.3134 | -44.0935
14 | YES | CDMA | 850 | QWERTY Candy Bar | 2006 | 26 | 102.2 | 40.189 | 101.8 | 40.155 | 98.7 | 39.8863 | M4 | 0.1978 | -14.0755 | 0.1935 | -14.2664 | 0.18 | -14.8945 | M4 | EM4 - HM4 | -10.811 | -14.6755 | OK |  | -11.1137 | -15.4945 | 548.3333 | 13.8863 | -40.8945
15 | YES | CDMA | 850 | Clam | 2006 | 25 | 100.9 | 40.0778 | 100.9 | 40.0778 | 103.9 | 40.3323 | M4 | 0.19 | -14.4249 | 0.1288 | -17.8017 | 0.134 | -17.4579 | M4 | EM4 - HM4 | -10.9222 | -15.0249 | OK |  | -10.6677 | -18.0579 | 775.3731 | 15.3323 | -42.4579
16 | YES | CDMA | 850 | Slider | 2006 | 25 | 153.6 | 43.7278 | 145.1 | 43.2333 | 158.2 | 43.9841 | M4 | 0.2335 | -12.6343 | 0.1683 | -15.4783 | 0.175 | -15.1392 | M4 | EM4 - HM4 | -7.2722 | -13.2343 | OK |  | -7.0159 | -15.7392 | 904 | 18.9841 | -40.1392
17 | YES | CDMA | 850 | Slider | 2006 | 25 | 170.5 | 44.6345 | 170.5 | 44.6345 | 179 | 45.0571 | M4 | 0.2577 | -11.7777 | 0.1816 | -14.8177 | 0.187 | -14.5632 | M4 | EM4 - HM4 | -6.3655 | -12.3777 | OK |  | -5.9429 | -15.1632 | 957.2193 | 20.0571 | -39.5632
18 | YES | CDMA | 850 | Clam | 2006 | 25 | 94.47 | 39.5059 | 94.47 | 39.5059 | 99.2 | 39.9302 | M4 | 0.1857 | -14.6238 | 0.1137 | -18.8848 | 0.121 | -18.3443 | M4 | EM4 - HM4 | -11.4941 | -15.2238 | OK |  | -11.0698 | -18.9443 | 819.8347 | 14.9302 | -43.3443
19 | YES | CDMA | 850 | QWERTY Candy Bar | 2006 | 26 | 114.4 | 41.1685 | 114.4 | 41.1685 | 117.8 | 41.4229 | M4 | 0.1907 | -14.393 | 0.1366 | -17.291 | 0.138 | -17.2024 | M4 | EM4 - HM4 | -9.8315 | -14.993 | OK |  | -9.5771 | -17.8024 | 853.6232 | 15.4229 | -43.2024
20 | YES | CDMA | 850 | Slider | 2006 | 25 | 154.6 | 43.7842 | 154.6 | 43.7842 | 162.3 | 44.2064 | M4 | 0.2427 | -12.2986 | 0.168 | -15.4938 | 0.173 | -15.2391 | M4 | EM4 - HM4 | -7.2158 | -12.8986 | OK |  | -6.7936 | -15.8391 | 938.1503 | 19.2064 | -40.2391
21 | YES | CDMA | 850 | Clam | 2006 | 25 | 106.9 | 40.5796 | 106.9 | 40.5796 | 112.2 | 40.9999 | M4 | 0.2077 | -13.6513 | 0.1447 | -16.7906 | 0.149 | -16.5363 | M4 | EM4 - HM4 | -10.4204 | -14.2513 | OK |  | -10.0001 | -17.1363 | 753.0201 | 15.9999 | -41.5363
22 | YES | CDMA | 850 | Clam | 2006 | 25 | 77.02 | 37.7321 | 77.02 | 37.7321 | 80.9 | 38.159 | M4 | 0.1442 | -16.8207 | 0.101 | -19.9136 | 0.104 | -19.6593 | M4 | EM4 - HM4 | -13.2679 | -17.4207 | OK |  | -12.841 | -20.2593 | 777.8846 | 13.159 | -44.6593
23 | YES | CDMA | 850 | Clam | 2006 | 25 | 97.66 | 39.7943 | 94.61 | 39.5187 | 96.5 | 39.6905 | M4 | 0.171 | -15.3401 | 0.1155 | -18.7484 | 0.119 | -18.4891 | M4 | EM4 - HM4 | -11.2057 | -15.9401 | OK |  | -11.3095 | -19.0891 | 810.9244 | 14.6905 | -43.4891
24 | YES | CDMA | 850 | Clam | 2006 | 25 | 101.2 | 40.1036 | 101.2 | 40.1036 | 103.3 | 40.282 | M4 | 0.1816 | -14.8177 | 0.1214 | -18.3156 | 0.125 | -18.0618 | M4 | EM4 - HM4 | -10.8964 | -15.4177 | OK |  | -10.718 | -18.6618 | 826.4 | 15.282 | -43.0618
25 | YES | CDMA | 850 | Clam | 2006 | 25 | 109.7 | 40.8041 | 107.8 | 40.6524 | 110 | 40.8279 | M4 | 0.2228 | -13.0417 | 0.1631 | -15.7509 | 0.168 | -15.4938 | M4 | EM4 - HM4 | -10.1959 | -13.6417 | OK |  | -10.1721 | -16.0938 | 654.7619 | 15.8279 | -40.4938
26 | YES | CDMA | 850 | Clam | 2006 | 25 | 93.1 | 39.379 | 93.1 | 39.379 | 95.9 | 39.6364 | M4 | 0.1883 | -14.503 | 0.1204 | -18.3875 | 0.124 | -18.1316 | M4 | EM4 - HM4 | -11.621 | -15.103 | OK |  | -11.3636 | -18.7316 | 773.3871 | 14.6364 | -43.1316
27 | YES | CDMA | 850 | Clam | 2006 | 25 | 100.8 | 40.0692 | 100.6 | 40.052 | 103.6 | 40.3072 | M4 | 0.2249 | -12.9602 | 0.1602 | -15.9067 | 0.165 | -15.6503 | M4 | EM4 - HM4 | -10.9308 | -13.5602 | OK |  | -10.6928 | -16.2503 | 627.8788 | 15.3072 | -40.6503
28 | YES | CDMA | 850 | Clam | 2007 | 24.5 | 88.2 | 38.9094 | 88.16 | 38.9054 | 90.8 | 39.1617 | M4 | 0.1766 | -15.0602 | 0.1243 | -18.1106 | 0.128 | -17.8558 | M4 | EM4 - HM4 | -12.0906 | -15.6602 | OK |  | -11.8383 | -18.4558 | 709.375 | 14.6617 | -42.3558
29 | YES | CDMA | 850 | Slider | 2007 | 25 | 129.3 | 42.232 | 128.8 | 42.1983 | 132.7 | 42.4574 | M4 | 0.2289 | -12.8071 | 0.1718 | -15.2995 | 0.177 | -15.0405 | M4 | EM4 - HM4 | -8.768 | -13.4071 | OK |  | -8.5426 | -15.6405 | 749.7175 | 17.4574 | -40.0405
30 | YES | CDMA | 850 | Slider | 2007 | 25 | 107.2 | 40.6039 | 107.2 | 40.6039 | 110.4 | 40.8594 | M4 | 0.1821 | -14.7938 | 0.1369 | -17.2719 | 0.141 | -17.0156 | M4 | EM4 - HM4 | -10.3961 | -15.3938 | OK |  | -10.1406 | -17.6156 | 782.9787 | 15.8594 | -42.0156
31 | YES | CDMA | 850 | Candy Bar | 2007 | 25 | 128.6 | 42.1848 | 128.58 | 42.1835 | 136.3 | 42.6899 | M4 | 0.2472 | -12.139 | 0.1828 | -14.7605 | 0.181 | -14.8464 | M4 | EM4 - HM4 | -8.8152 | -12.739 | OK |  | -8.3101 | -15.4464 | 753.0387 | 17.6899 | -39.8464
32 | YES | GSM | 850 | Clam | 2006 | 32.5 | 75.66 | 37.5773 | 75.66 | 37.5773 | 214.9 | 46.6447 | M3 | 0.1672 | -15.5353 | 0.1181 | -18.555 | 0.313 | -10.0891 | M4 | EM3 - HM4 | -10.9227 | -13.6353 | OK |  | -1.8553 | -8.1891 | 686.5815 | 14.1447 | -42.5891
32 |  | WCDMA | 850 |  |  | 24 | 87.21 | 38.8113 | 87.17 | 38.8073 | 80.2 | 38.0835 | M4 | 0.1766 | -15.0602 | 0.1258 | -18.0064 | 0.117 | -18.6363 | M4 | EM4 - HM4 | -12.1887 | -15.6602 | OK |  | -12.9165 | -19.2363 | 685.4701 | 14.0835 | -42.6363
33 | YES | GSM | 850 | Clam | 2006 | 33 | 69.69 | 36.8634 | 69.69 | 36.8634 | 199.3 | 45.9901 | M3 | 0.143 | -16.8933 | 0.0955 | -20.3999 | 0.231 | -12.7278 | M4 | EM3 - HM4 | -11.6366 | -14.9933 | OK |  | -2.5099 | -10.8278 | 862.7706 | 12.9901 | -45.7278
33 |  | WCDMA | 850 |  |  | 24 | 75.71 | 37.5831 | 75.67 | 37.5785 | 68.1 | 36.6629 | M4 | 0.1465 | -16.6832 | 0.09778 | -20.195 | 0.088 | -21.1103 | M4 | EM4 - HM4 | -13.4169 | -17.2832 | OK |  | -14.3371 | -21.7103 | 773.8636 | 12.6629 | -45.1103
34 | YES | GSM | 850 | Clam | 2007 | 33.3 | 55.64 | 34.9077 | 55.54 | 34.8921 | 158.3 | 43.9896 | M3 | 0.1153 | -18.7634 | 0.0846 | -21.4526 | 0.209 | -13.5971 | M4 | EM3 - HM4 | -13.5923 | -16.8634 | OK | FALSE | -4.5104 | -11.6971 | 757.4163 | 10.6896 | -46.8971
34 |  | WCDMA | 850 |  |  | 24 | 50.55 | 34.0744 | 50.54 | 34.0727 | 46.5 | 33.3491 | M4 | 0.1014 | -19.8792 | 0.0758 | -22.4066 | 0.069 | -23.223 | M4 | EM4 - HM4 | -16.9256 | -20.4792 | OK | FALSE | -17.6509 | -23.823 | 673.913 | 9.3491 | -47.223
35 | YES | GSM | 850 | Slider | 2007 | 32.5 | 85.86 | 38.6758 | 85.86 | 38.6758 | 244.7 | 47.7727 | M3 | 0.1612 | -15.8527 | 0.1121 | -19.0079 | 0.277 | -11.1504 | M4 | EM3 - HM4 | -9.8242 | -13.9527 | OK | FALSE | -0.72732 | -9.2504 | 883.3935 | 15.2727 | -43.6504
35 |  | WCDMA | 850 |  |  | 24 | 72.78 | 37.2402 | 72.78 | 37.2402 | 66.95 | 36.515 | M4 | 0.1363 | -17.3101 | 0.1014 | -19.8792 | 0.0923 | -20.696 | M4 | EM4 - HM4 | -13.7598 | -17.9101 | OK | FALSE | -14.485 | -21.296 | 725.3521 | 12.515 | -44.696
36 | YES | GSM | 850 | Clam | 2006 | 33 | 89.7 | 39.0558 | 89.69 | 39.0549 | 257.4 | 48.2122 | M3 | 0.2326 | -12.6678 | 0.1595 | -15.9448 | 0.426 | -7.4118 | M4 | EM3 - HM4 | -9.4442 | -10.7678 | OK | FALSE | -0.28783 | -5.5118 | 604.2254 | 15.2122 | -40.4118
37 | YES | GSM | 850 | Clam | 2006 | 32.5 | 89.52 | 39.0384 | 89.51 | 39.0374 | 254.2 | 48.1035 | M3 | 0.1782 | -14.9818 | 0.1189 | -18.4964 | 0.315 | -10.0338 | M4 | EM3 - HM4 | -9.4616 | -13.0818 | OK | FALSE | -0.39649 | -8.1338 | 806.9841 | 15.6035 | -42.5338
38 | YES | GSM | 850 | Clam | 2006 | 32.5 | 86.9 | 38.7804 | 86.9 | 38.7804 | 246.8 | 47.8469 | M3 | 0.1433 | -16.8751 | 0.0955 | -20.3999 | 0.233 | -12.6529 | M4 | EM3 - HM4 | -9.7196 | -14.9751 | OK | FALSE | -0.6531 | -10.7529 | 1059.2275 | 15.3469 | -45.1529
39 | YES | GSM | 850 | Clam | 2006 | 32.3 | 69.24 | 36.8071 | 69.23 | 36.8059 | 196.6 | 45.8717 | M3 | 0.1229 | -18.209 | 0.0902 | -20.8959 | 0.22 | -13.1515 | M4 | EM3 - HM4 | -11.6929 | -16.309 | OK | FALSE | -2.6283 | -11.2515 | 893.6364 | 13.5717 | -45.4515
40 | YES | GSM | 850 | Clam | 2006 | 32.3 | 82.53 | 38.3322 | 82.09 | 38.2858 | 231.5 | 47.291 | M3 | 0.1867 | -14.5771 | 0.1191 | -18.4818 | 0.281 | -11.0259 | M4 | EM3 - HM4 | -10.1678 | -12.6771 | OK | FALSE | -1.209 | -9.1259 | 823.8434 | 14.991 | -43.3259
41 | YES | GSM | 850 | Candy Bar | 2006 | 33 | 85.97 | 38.6869 | 85.53 | 38.6424 | 241.2 | 47.6475 | M3 | 0.2676 | -11.4503 | 0.2674 | -11.4568 | 0.631 | -3.9994 | M3 | EM3 - HM3 | -9.8131 | -9.5503 | Not OK | 0.26278 | -0.85245 | -2.0994 | 382.2504 | 14.6475 | -36.9994
42 | YES | GSM | 850 | Slider | 2006 | 32.3 | 86.62 | 38.7524 | 86.62 | 38.7524 | 247.7 | 47.8785 | M3 | 0.1463 | -16.6951 | 0.1021 | -19.8195 | 0.247 | -12.1461 | M4 | EM3 - HM4 | -9.7476 | -14.7951 | OK | FALSE | -0.62148 | -10.2461 | 1002.834 | 15.5785 | -44.4461
43 | YES | GSM | 850 | Clam | 2006 | 32.8 | 54.29 | 34.6944 | 54.29 | 34.6944 | 153.1 | 43.6995 | M3 | 0.1372 | -17.2529 | 0.0877 | -21.14 | 0.207 | -13.6806 | M4 | EM3 - HM4 | -13.8056 | -15.3529 | OK | FALSE | -4.8005 | -11.7806 | 739.6135 | 10.8995 | -46.4806
44 | YES | GSM | 850 | Rotator | 2007 | 33 | 89.01 | 38.9888 | 89.01 | 38.9888 | 253.7 | 48.0864 | M3 | 0.1489 | -16.5421 | 0.1032 | -19.7264 | 0.255 | -11.8692 | M4 | EM3 - HM4 | -9.5112 | -14.6421 | OK | FALSE | -0.41359 | -9.9692 | 994.902 | 15.0864 | -44.8692
45 | YES | GSM | 850 | Clam | 2007 | 33.5 | 24.18 | 27.6691 | 23.58 | 27.4509 | 67.2 | 36.5474 | M4 | 0.0533 | -25.4655 | 0.0508 | -25.8827 | 0.118 | -18.5624 | M4 | EM4 - HM4 | -20.8309 | -23.5655 | OK | FALSE | -11.9526 | -16.6624 | 569.4915 | 3.0474 | -52.0624
46 | YES | GSM | 850 | Clam | 2007 | 32.8 | 44.95 | 33.0546 | 44.95 | 33.0546 | 128.1 | 42.151 | M4 | 0.1057 | -19.5185 | 0.0746 | -22.5452 | 0.173 | -15.2391 | M4 | EM4 - HM4 | -15.4454 | -17.6185 | OK | FALSE | -6.349 | -13.3391 | 740.4624 | 9.351 | -48.0391
47 | YES | GSM | 850 | Clam | 2006 | 32.5 | 67.23 | 36.5513 | 67.23 | 36.5513 | 191.6 | 45.6479 | M3 | 0.1217 | -18.2942 | 0.0798 | -21.9599 | 0.197 | -14.1107 | M4 | EM3 - HM4 | -11.9487 | -16.3942 | OK | FALSE | -2.8521 | -12.2107 | 972.5888 | 13.1479 | -46.6107
48 | NO | iDEN | 800 |  | 2006 | 28.1 | 75.18 | 37.522 | 74.05 | 37.3905 | 371 | 51.3875 | M2 | 0.1451 | -16.7667 | 0.0946 | -20.4822 | 0.419 | -7.5557 | M4 | EM2 - HM4 | -13.478 | -17.3667 | OK | FALSE | 0.38748 | -8.1557 | 885.4415 | 23.2875 | -35.6557
48 |  |  | 900 |  |  | 28.1 | 59.04 | 35.4229 | 58.86 | 35.3964 | 284.3 | 49.0755 | M3 | 0.1413 | -16.9972 | 0.0933 | -20.6024 | 0.407 | -7.8081 | M4 | EM3 - HM4 | -15.5771 | -17.5972 | OK | FALSE | -1.9245 | -8.4081 | 698.5258 | 20.9755 | -35.9081
49 | NO | iDEN | 800 |  | 2006 | 28.1 | 76.56 | 37.68 | 75.25 | 37.5301 | 377 | 51.5268 | M2 | 0.1677 | -15.5093 | 0.1106 | -19.1249 | 0.49 | -6.1961 | M4 | EM2 - HM4 | -13.32 | -16.1093 | OK | FALSE | 0.52683 | -6.7961 | 769.3878 | 23.4268 | -34.2961
49 |  |  | 900 |  |  | 28.1 | 65.19 | 36.2836 | 63.87 | 36.1059 | 308.5 | 49.7851 | M3 | 0.1483 | -16.5772 | 0.0907 | -20.8479 | 0.438 | -7.1705 | M4 | EM3 - HM4 | -14.7164 | -17.1772 | OK | FALSE | -1.2149 | -7.7705 | 704.3379 | 21.6851 | -35.2705
50 | NO | iDEN | 800 | Candy Bar | 2006 | 28.1 | 69.97 | 36.8982 | 69.97 | 36.8982 | 343.5 | 50.7185 | M3 | 0.1152 | -18.771 | 0.1048 | -19.5928 | 0.456 | -6.8207 | M4 | EM3 - HM4 | -14.1018 | -19.371 | OK | FALSE | -0.28147 | -7.4207 | 753.2895 | 22.6185 | -34.9207
50 |  |  | 900 |  |  | 28.1 | 69.57 | 36.8484 | 69.57 | 36.8484 | 334 | 50.4749 | M3 | 0.109 | -19.2515 | 0.1018 | -19.845 | 0.445 | -7.0328 | M4 | EM3 - HM4 | -14.1516 | -19.8515 | OK | FALSE | -0.52507 | -7.6328 | 750.5618 | 22.3749 | -35.1328
51 | NO | iDEN | 800 | Candy Bar | 2006 | 28.06 | 83.52 | 38.4358 | 83.52 | 38.4358 | 410.1 | 52.2578 | M2 | 0.1195 | -18.4526 | 0.114 | -18.8619 | 0.496 | -6.0904 | M4 | EM2 - HM4 | -12.5642 | -19.0526 | OK | FALSE | 1.2578 | -6.6904 | 826.8145 | 24.1978 | -34.1504
51 |  |  | 900 |  |  | 28.06 | 94.83 | 39.5389 | 94.75 | 39.5316 | 454.8 | 53.1564 | M2 | 0.1345 | -17.4256 | 0.1341 | -17.4514 | 0.586 | -4.642 | M4 | EM2 - HM4 | -11.4611 | -18.0256 | OK | FALSE | 2.1564 | -5.242 | 776.1092 | 25.0964 | -32.702
52 | NO | iDEN | 800 | Candy Bar | 2006 | 28.1 | 73.89 | 37.3717 | 73.89 | 37.3717 | 370.2 | 51.3687 | M2 | 0.1541 | -16.2439 | 0.1009 | -19.9222 | 0.447 | -6.9938 | M4 | EM2 - HM4 | -13.6283 | -16.8439 | OK | FALSE | 0.36873 | -7.5938 | 828.1879 | 23.2687 | -35.0938
52 |  |  | 900 |  |  | 28.1 | 63.69 | 36.0814 | 63.69 | 36.0814 | 307.6 | 49.7597 | M3 | 0.1399 | -17.0836 | 0.0943 | -20.5098 | 0.411 | -7.7232 | M4 | EM3 - HM4 | -14.9186 | -17.6836 | OK | FALSE | -1.2403 | -8.3232 | 748.4185 | 21.6597 | -35.8232
53 | NO | iDEN | 800 | Candy Bar | 2006 | 28.06 | 66.2 | 36.4172 | 65.9 | 36.3777 | 333.5 | 50.4619 | M3 | 0.1155 | -18.7484 | 0.1148 | -18.8012 | 0.488 | -6.2316 | M4 | EM3 - HM4 | -14.5828 | -19.3484 | OK | FALSE | -0.53808 | -6.8316 | 683.4016 | 22.4019 | -34.2916
53 |  |  | 900 |  |  | 28.06 | 100.9 | 40.0778 | 100.9 | 40.0778 | 491.5 | 53.8305 | M2 | 0.1897 | -14.4387 | 0.1482 | -16.583 | 0.63 | -4.0132 | M3 | EM2 - HM3 | -10.9222 | -15.0387 | OK | FALSE | 2.8305 | -4.6132 | 780.1587 | 25.7705 | -32.0732
54 | NO | GSM | 850 | QWERTY Candy Bar | 2006 | 32.5 | 83.48 | 38.4316 | 80.14 | 38.077 | 228.4 | 47.1739 | M3 | 0.1674 | -15.5249 | 0.115 | -18.786 | 0.284 | -10.9336 | M4 | EM3 - HM4 | -10.0684 | -13.6249 | OK | FALSE | -1.3261 | -9.0336 | 804.2254 | 14.6739 | -43.4336
54 | NO | WCDMA | 850 |  | 2006 | 24 | 85.75 | 38.6647 | 82.61 | 38.3407 | 76 | 37.6163 | M4 | 0.1517 | -16.3803 | 0.1187 | -18.511 | 0.108 | -19.3315 | M4 | EM4 - HM4 | -12.3353 | -16.9803 | OK | FALSE | -13.3837 | -19.9315 | 703.7037 | 13.6163 | -43.3315
55 | NO | GSM | 850 | Clam | 2006 | 33 | 99.32 | 39.9407 | 99.26 | 39.9355 | 281.9 | 49.0019 | M2 | 0.1532 | -16.2948 | 0.0981 | -20.1666 | 0.26 | -11.7005 | M4 | EM2 - HM4 | -8.5593 | -14.3948 | OK | FALSE | 0.5019 | -9.8005 | 1084.2308 | 16.0019 | -44.7005
56 | NO | GSM | 850 | Candy Bar | 2006 | 32.6 | 129.5 | 42.2454 | 129.5 | 42.2454 | 370.4 | 51.3734 | M2 | 0.2553 | -11.859 | 0.1778 | -15.0014 | 0.416 | -7.6181 | M4 | EM2 - HM4 | -6.2546 | -9.959 | OK | FALSE | 2.8734 | -5.7181 | 890.3846 | 18.7734 | -40.2181
57 | NO | GSM | 850 | Clam | 2006 | 32.8 | 117.7 | 41.4155 | 117.7 | 41.4155 | 360.1 | 51.1285 | M2 | 0.2085 | -13.6179 | 0.1264 | -17.9651 | 0.34 | -9.3704 | M4 | EM2 - HM4 | -7.0845 | -11.7179 | OK | FALSE | 2.6285 | -7.4704 | 1059.1176 | 18.3285 | -42.1704
58 | NO | GSM | 850 | Clam | Pre-2006 | 32.5 | 101.3 | 40.1122 | 101.3 | 40.1122 | 290.6 | 49.2659 | >M1 | 0.2008 | -13.9447 | 0.1333 | -17.5034 | 0.356 | -8.971 | M2 | E>M1 - HM2 | -8.3878 | -12.0447 | OK | FALSE | 0.76591 | -7.071 | 816.2921 | 16.7659 | -41.471
59 | NO | GSM | 850 | Candy Bar | Pre-2006 | 32.3 | 124 | 41.8684 | 124 | 41.8684 | 353.5 | 50.9678 | >M1 | 0.2851 | -10.9001 | 0.2 | -13.9794 | 0.54 | -5.3521 | M1 | E>M1 - HM1 | -6.6316 | -9.0001 | OK | FALSE | 2.4678 | -3.4521 | 654.6296 | 18.6678 | -37.6521
60 | NO | GSM | 850 | Slider | 2006 | 33 | 126.2 | 42.0212 | 126.2 | 42.0212 | 358.4 | 51.0874 | M2 | 0.2277 | -12.8527 | 0.1585 | -15.9994 | 0.42 | -7.535 | M4 | EM2 - HM4 | -6.4788 | -10.9527 | OK | FALSE | 2.5874 | -5.635 | 853.3333 | 18.0874 | -40.535
61 | NO | GSM | 850 | Candy Bar | 2006 | 33 | 87.25 | 38.8153 | 87.25 | 38.8153 | 249.5 | 47.9414 | M3 | 0.1738 | -15.199 | 0.1181 | -18.555 | 0.289 | -10.782 | M4 | EM3 - HM4 | -9.6847 | -13.299 | OK | FALSE | -0.55859 | -8.882 | 863.3218 | 14.9414 | -43.782
62 | NO | GSM | 850 | Clam | Pre-2006 | 33 | 63.06 | 35.9951 | 63.06 | 35.9951 | 184.1 | 45.3011 | M3 | 0.1145 | -18.8239 | 0.0776 | -22.2028 | 0.208 | -13.6387 | M4 | EM3 - HM4 | -12.5049 | -16.9239 | OK | FALSE | -3.1989 | -11.7387 | 885.0962 | 12.3011 | -46.6387
63 | NO | GSM | 850 | Clam | 2006 | 33 | 105.5 | 40.465 | 105.5 | 40.465 | 301.6 | 49.5886 | M2 | 0.1972 | -14.1019 | 0.126 | -17.9926 | 0.305 | -10.314 | M4 | EM2 - HM4 | -8.035 | -12.2019 | OK | FALSE | 1.0886 | -8.414 | 988.8525 | 16.5886 | -43.314
64 | NO | GSM | 850 | Clam | 2006 | 33 | 112.9 | 41.0539 | 112.9 | 41.0539 | 329.7 | 50.3624 | M2 | 0.2283 | -12.8299 | 0.1552 | -16.1822 | 0.416 | -7.6181 | M4 | EM2 - HM4 | -7.4461 | -10.9299 | OK | FALSE | 1.8624 | -5.7181 | 792.5481 | 17.3624 | -40.6181
65 | NO | GSM | 850 | QWERTY Side Slider | 2006 | 32.7 | 86.3 | 38.7202 | 69.61 | 36.8534 | 199.1 | 45.9814 | M3 | 0.2328 | -12.6603 | 0.2293 | -12.7919 | 0.555 | -5.1141 | M3 | EM3 - HM3 | -9.7798 | -10.7603 | OK | FALSE | -2.5186 | -3.2141 | 358.7387 | 13.2814 | -37.8141
66 | NO | GSM | 850 | Clam | Pre-2006 | 33 | 116.8 | 41.3489 | 116.8 | 41.3489 | 335.1 | 50.5035 | >M1 | 0.2576 | -11.7811 | 0.1693 | -15.4269 | 0.452 | -6.8972 | M1 | E>M1 - HM1 | -7.1511 | -9.8811 | OK | FALSE | 2.0035 | -4.9972 | 741.3717 | 17.5035 | -39.8972
67 | NO | GSM | 850 | Slider | 2006 | 33 | 106.05 | 40.5102 | 106.05 | 40.5102 | 301.2 | 49.5771 | M2 | 0.1825 | -14.7747 | 0.1268 | -17.9376 | 0.336 | -9.4732 | M4 | EM2 - HM4 | -7.9898 | -12.8747 | OK | FALSE | 1.0771 | -7.5732 | 896.4286 | 16.5771 | -42.4732
68 | NO | GSM | 850 | Candy Bar | Pre-2006 | 32.8 | 67.65 | 36.6054 | 62.87 | 35.9689 | 179.2 | 45.0668 | M1 | 0.2757 | -11.1913 | 0.2707 | -11.3502 | 0.731 | -2.7217 | M1 | EM1 - HM1 | -11.8946 | -9.2913 | Not OK | 2.6034 | -3.4332 | -0.82165 | 245.1436 | 12.2668 | -35.5217
69 | NO | GSM | 850 | Candy Bar | 2006 | 33.1 | 103.6 | 40.3072 | 103.6 | 40.3072 | 294.2 | 49.3729 | M2 | 0.165 | -15.6503 | 0.1105 | -19.1328 | 0.293 | -10.6626 | M4 | EM2 - HM4 | -8.1928 | -13.7503 | OK | FALSE | 0.87285 | -8.7626 | 1004.0956 | 16.2729 | -43.7626
70 | NO | GSM | 850 | Candy Bar | 2006 | 32.5 | 102.9 | 40.2483 | 102.9 | 40.2483 | 293.5 | 49.3522 | M2 | 0.2672 | -11.4633 | 0.2129 | -13.4365 | 0.494 | -6.1255 | M3 | EM2 - HM3 | -8.2517 | -9.5633 | OK | FALSE | 0.85216 | -4.2255 | 594.1296 | 16.8522 | -38.6255
71 | NO | GSM | 850 | Clam | 2006 | 32.5 | 77.7 | 37.8084 | 77.7 | 37.8084 | 222.2 | 46.9349 | M3 | 0.1624 | -15.7883 | 0.1089 | -19.2594 | 0.269 | -11.405 | M4 | EM3 - HM4 | -10.6916 | -13.8883 | OK | FALSE | -1.5651 | -9.505 | 826.0223 | 14.4349 | -43.905
72 | NO | GSM | 850 | Candy Bar | 2006 | 33 | 109 | 40.7485 | 108.87 | 40.7382 | 307 | 49.7428 | M2 | 0.2374 | -12.4904 | 0.1563 | -16.1208 | 0.369 | -8.6595 | M4 | EM2 - HM4 | -7.7515 | -10.5904 | OK | FALSE | 1.2428 | -6.7595 | 831.9783 | 16.7428 | -41.6595
73 | NO | GSM | 850 | Candy Bar | 2006 | 33 | 105.1 | 40.4321 | 104.4 | 40.374 | 297.4 | 49.4668 | M2 | 0.2021 | -13.8887 | 0.1385 | -17.171 | 0.342 | -9.3195 | M4 | EM2 - HM4 | -8.0679 | -11.9887 | OK | FALSE | 0.96682 | -7.4195 | 869.5906 | 16.4668 | -42.3195
[6 empty] | Peak / pulsed averaged channel power (dBm) | [15 empty] | Delta (dB) E-Field relative to M3 (pass) | Delta (dB) H-Field relative to M3 (pass)
101 | Pass | iDEN 800 |  | Candy Bar PDA | 2006 | 28 |  |  |  |  | 266.3 | 48.5074 | 3 |  |  |  |  | 0.704 | -3.0485 | 3 | 3 | -2.4926 | -3.6485 | OK |  |  |  | 378.267 | 20.5074 | -31.0485
101 |  | iDEN 900 |  |  | 2006 | 28 |  |  |  |  | 239.4 | 47.5825 | 3 |  |  |  |  | 0.575 | -4.8066 | 4 |  | -3.4175 | -5.4066 | OK |  |  |  | 416.3478 | 19.5825 | -32.8066
102 | Pass | iDEN 800 |  | Candy Bar PDA | 2006 | 28 |  |  |  |  | 235.62 | 47.4442 | 3 |  |  |  |  | 0.693 | -3.1853 | 3 | 3 | -3.5558 | -3.7853 | OK |  |  |  | 340 | 19.4442 | -31.1853
103 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 238.1 | 47.5352 | 3 |  |  |  |  | 0.499 | -6.038 | 3 | 2 | -0.96481 | -4.138 | OK |  |  |  | 477.1543 | 14.5352 | -39.038
104 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 197.2 | 45.8981 | 3 |  |  |  |  | 0.404 | -7.8724 | 3 | 3 | -2.6019 | -5.9724 | OK |  |  |  | 488.1188 | 12.8981 | -40.8724
105 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 259.6 | 48.2861 | 4 |  |  |  |  | 0.507 | -5.8998 | 3 | 2 | -0.21391 | -3.9998 | OK |  |  |  | 512.0316 | 15.2861 | -38.8998
106 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 279.1 | 48.9152 | 2 |  |  |  |  | 0.499 | -6.038 | 3 | 2 | 0.4152 | -4.138 | OK |  |  |  | 559.3186 | 15.9152 | -39.038
107 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 210.4 | 46.4609 | 3 |  |  |  |  | 0.41 | -7.7443 | 4 | 3 | -2.0391 | -5.8443 | OK |  |  |  | 513.1707 | 13.4609 | -40.7443
108 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 224.8 | 47.0359 | 3 |  |  |  |  | 0.41 | -7.7443 | 4 | 3 | -1.4641 | -5.8443 | OK |  |  |  | 548.2927 | 14.0359 | -40.7443
109 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 224.9 | 47.0398 | 3 |  |  |  |  | 0.41 | -7.7443 | 4 | 2 | -1.4602 | -5.8443 | OK |  |  |  | 548.5366 | 14.0398 | -40.7443
110 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 160.6 | 44.1149 | 3 |  |  |  |  | 0.273 | -11.2767 | 4 | 2 | -4.3851 | -9.3767 | OK |  |  |  | 588.2784 | 11.1149 | -44.2767
111 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 229.9 | 47.2308 | 3 |  |  |  |  | 0.413 | -7.681 | 4 | 2 | -1.2692 | -5.781 | OK |  |  |  | 556.6586 | 14.2308 | -40.681
112 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 298.1 | 49.4872 | 2 |  |  |  |  | 0.53 | -5.5145 | 3 | 1 | 0.98724 | -3.6145 | OK |  |  |  | 562.4528 | 16.4872 | -38.5145
113 | Pass | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 227.6 | 47.1434 | 3 |  |  |  |  | 0.45 | -6.9357 | 4 | 3 | -1.3566 | -5.0357 | OK |  |  |  | 505.7778 | 14.1434 | -39.9357
114 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 282.2 | 49.0111 | 2 |  |  |  |  | 0.529 | -5.5309 | 3 | 2 | 0.51114 | -3.6309 | OK |  |  |  | 533.4594 | 16.0111 | -38.5309
115 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 257.5 | 48.2155 | 3 |  |  |  |  | 0.34 | -9.3704 | 4 | 2 | -0.28446 | -7.4704 | OK |  |  |  | 757.3529 | 15.2155 | -42.3704
116 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 230.7 | 47.261 | 3 |  |  |  |  | 0.354 | -9.0199 | 4 | 1 | -1.239 | -7.1199 | OK |  |  |  | 651.6949 | 14.261 | -42.0199
117 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 244 | 47.7478 | 3 |  |  |  |  | 0.363 | -8.8019 | 4 | 1 | -0.7522 | -6.9019 | OK |  |  |  | 672.1763 | 14.7478 | -41.8019
119 | Fail | GSM 850 |  |  | 2006 | 33 |  |  |  |  | 106.5 | 40.547 | 4 |  |  |  |  | 0.286 | -10.8727 | 4 | 2 | -7.953 | -8.9727 | OK |  |  |  | 372.3776 | 7.547 | -43.8727
121 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 29 |  |  |  |  | 77.1 | 37.7411 | 4 |  |  |  |  | 0.264 | -11.5679 | 4 | 2 | -10.7589 | -9.6679 | Not OK | 1.091 |  |  | 292.0455 | 8.7411 | -40.5679
122 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 29 |  |  |  |  | 88.2 | 38.9094 | 4 |  |  |  |  | 0.323 | -9.8159 | 4 | 2 | -9.5906 | -7.9159 | Not OK | 1.6747 |  |  | 273.065 | 9.9094 | -38.8159
123 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 29 |  |  |  |  | 86.5 | 38.7403 | 4 |  |  |  |  | 0.306 | -10.2856 | 4 | 2 | -9.7597 | -8.3856 | Not OK | 1.3741 |  |  | 282.6797 | 9.7403 | -39.2856
126 | Fail | GSM 850 |  | Candy Bar PDA | 2006 | 33 |  |  |  |  | 170.1 | 44.6141 | 3 |  |  |  |  | 0.303 | -10.3711 | 4 | 2 | -3.8859 | -8.4711 | OK |  |  |  | 561.3861 | 11.6141 | -43.3711
133 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 92.1 | 39.2852 | 4 |  |  |  |  | 0.165 | -15.6503 | 4 | 4 | -11.7148 | -16.2503 | OK |  |  |  | 558.1818 | 14.7852 | -40.1503
134 | Pass | WCDMA FDD V 850 |  | Candy Bar PDA | 2006 | 23.5 |  |  |  |  | 63 | 35.9868 | 4 |  |  |  |  | 0.13 | -17.7211 | 4 | 4 | -15.0132 | -18.3211 | OK |  |  |  | 484.6154 | 12.4868 | -41.2211
135 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 94.5 | 39.5086 | 4 |  |  |  |  | 0.169 | -15.4423 | 4 | 4 | -11.4914 | -16.0423 | OK |  |  |  | 559.1716 | 15.0086 | -39.9423
136 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 145 | 43.2274 | 4 |  |  |  |  | 0.26 | -11.7005 | 4 | 3 | -7.7726 | -12.3005 | OK |  |  |  | 557.6923 | 18.7274 | -36.2005
137 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 106.5 | 40.547 | 4 |  |  |  |  | 0.23 | -12.7654 | 4 | 4 | -10.453 | -13.3654 | OK |  |  |  | 463.0435 | 16.047 | -37.2654
138 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 91.1 | 39.1904 | 4 |  |  |  |  | 0.134 | -17.4579 | 4 | 3 | -11.8096 | -18.0579 | OK |  |  |  | 679.8507 | 14.6904 | -41.9579
139 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 52.3 | 34.37 | 4 |  |  |  |  | 0.192 | -14.334 | 4 | 3 | -16.63 | -14.934 | Not OK | 1.696 |  |  | 272.3958 | 9.87 | -38.834
140 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 60.8 | 35.6781 | 4 |  |  |  |  | 0.121 | -18.3443 | 4 | 4 | -15.3219 | -18.9443 | OK |  |  |  | 502.4793 | 11.1781 | -42.8443
141 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 63.1 | 36.0006 | 4 |  |  |  |  | 0.13 | -17.7211 | 4 | 4 | -14.9994 | -18.3211 | OK |  |  |  | 485.3846 | 11.5006 | -42.2211
142 | Pass | CDMA 800 |  | Candy Bar PDA | 2006 | 24.5 |  |  |  |  | 66.2 | 36.4172 | 4 |  |  |  |  | 0.127 | -17.9239 | 4 |  | -14.5828 | -18.5239 | OK |  |  |  | 521.2598 | 11.9172 | -42.4239
201 | Fail | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 237.8 | 47.5242 | 3 |  |  |  |  | 0.491 | -6.1784 | 3 | 2 | -0.97576 | -4.2784 | OK |  |  |  | 484.3177 | 14.5242 | -39.1784
202 | Pass | GSM 850 |  | Flip | 2007 | 33 |  |  |  |  | 188.09 | 45.4873 | 3 |  |  |  |  | 0.396 | -8.0461 | 3 | 3 | -3.0127 | -6.1461 | OK |  |  |  | 474.9747 | 12.4873 | -41.0461
203 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 211.2 | 46.4939 | 3 |  |  |  |  | 0.5 | -6.0206 | 4 | 3 | -2.0061 | -4.1206 | OK |  |  |  | 422.4 | 13.4939 | -39.0206
203 |  | CDMA 800 |  |  | 2007 | 24.5 |  |  |  |  | 74.3 | 37.4198 | 4 |  |  |  |  | 0.183 | -14.751 | 4 |  | -13.5802 | -15.351 | OK |  |  |  | 406.0109 | 12.9198 | -39.251
204 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 212 | 46.5267 | 3 |  |  |  |  | 0.462 | -6.7072 | 3 | 3 | -1.9733 | -4.8072 | OK |  |  |  | 458.8745 | 13.5267 | -39.7072
205 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 182.6 | 45.23 | 3 |  |  |  |  | 0.391 | -8.1565 | 4 | 3 | -3.27 | -6.2565 | OK |  |  |  | 467.0077 | 12.23 | -41.1565
205 |  | CDMA 800 |  |  | 2007 | 24.5 |  |  |  |  | 58.1 | 35.2835 | 4 |  |  |  |  | 0.107 | -19.4123 | 4 |  | -15.7165 | -20.0123 | OK |  |  |  | 542.9907 | 10.7835 | -43.9123
206 | Pass | GSM 850 |  | Candy Bar PDA | 2007 | 33 |  |  |  |  | 179.2 | 45.0668 | 3 |  |  |  |  | 0.37 | -8.636 | 4 | 3 | -3.4332 | -6.736 | OK |  |  |  | 484.3243 | 12.0668 | -41.636
207 | Pass | iDEN 800 |  | Candy Bar PDA | 2007 | 27.5 |  |  |  |  | 225.6 | 47.0668 | 3 |  |  |  |  | 0.517 | -5.7302 | 4 | 3 | -3.9332 | -6.3302 | OK |  |  |  | 436.3636 | 19.5668 | -33.2302
207 |  | iDEN 900 |  |  | 2007 | 27.5 |  |  |  |  | 253.6 | 48.083 | 3 |  |  |  |  | 0.546 | -5.2561 | 4 |  | -2.917 | -5.8561 | OK |  |  |  | 464.4689 | 20.583 | -32.7561
208 | Pass | CDMA 800 |  | Flip | 2007 | 24.5 |  |  |  |  | 88.2 | 38.9094 | 4 |  |  |  |  | 0.175 | -15.1392 | 4 | 4 | -12.0906 | -15.7392 | OK |  |  |  | 504 | 14.4094 | -39.6392
209 | Pass | CDMA 800 |  | Candy Bar PDA | 2007 | 24.5 |  |  |  |  | 78.7 | 37.9195 | 4 |  |  |  |  | 0.178 | -14.9916 | 4 | 4 | -13.0805 | -15.5916 | OK |  |  |  | 442.1348 | 13.4195 | -39.4916
300 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 248 | 47.889 | [5 empty] | 0.36 | -8.8739 |  | 2 | -0.61097 | -6.9739 | OK |  |  |  | 688.8889 | 14.889 | -41.8739
301 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 219.5 | 46.8287 | [5 empty] | 0.281 | -11.0259 |  | 3 | -1.6713 | -9.1259 | OK |  |  |  | 781.1388 | 13.8287 | -44.0259
302 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 333 | 50.4489 | [5 empty] | 0.394 | -8.0901 |  | 2 | 1.9489 | -6.1901 | OK |  |  |  | 845.1777 | 17.4489 | -41.0901
303 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 273.3 | 48.7328 | [5 empty] | 0.452 | -6.8972 |  | 2 | 0.23279 | -4.9972 | OK |  |  |  | 604.646 | 15.7328 | -39.8972
305 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 256.7 | 48.1885 | [5 empty] | 0.378 | -8.4502 |  | 2 | -0.31148 | -6.5502 | OK |  |  |  | 679.1005 | 15.1885 | -41.4502
307 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 302.8 | 49.6231 | [5 empty] | 0.401 | -7.9371 |  | 2 | 1.1231 | -6.0371 | OK |  |  |  | 755.1122 | 16.6231 | -40.9371
312 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 141.1 | 42.9905 | [5 empty] | 0.225 | -12.9563 |  | 2 | -5.5095 | -11.0563 | OK |  |  |  | 627.1111 | 9.9905 | -45.9563
313 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 219.2 | 46.8168 | [5 empty] | 0.304 | -10.3425 |  | 2 | -1.6832 | -8.4425 | OK |  |  |  | 721.0526 | 13.8168 | -43.3425
314 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 155.7 | 43.8458 | [5 empty] | 0.446 | -7.0133 |  | 2 | -4.6542 | -5.1133 | OK |  |  |  | 349.1031 | 10.8458 | -40.0133
315 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 179.5 | 45.0813 | [5 empty] | 0.257 | -11.8013 |  | 2 | -3.4187 | -9.9013 | OK |  |  |  | 698.4436 | 12.0813 | -44.8013
316 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 95.5 | 39.6001 | [5 empty] | 0.268 | -11.4373 |  | 2 | -8.8999 | -9.5373 | OK |  |  |  | 356.3433 | 6.6001 | -44.4373
317 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 213.2 | 46.5757 | [5 empty] | 0.471 | -6.5396 |  | 2 | -1.9243 | -4.6396 | OK |  |  |  | 452.6539 | 13.5757 | -39.5396
318 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 176.4 | 44.93 | [5 empty] | 0.351 | -9.0939 |  | 3 | -3.57 | -7.1939 | OK |  |  |  | 502.5641 | 11.93 | -42.0939
319 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 265.3 | 48.4747 | [5 empty] | 0.285 | -10.9031 |  | 1 | -0.025255 | -9.0031 | OK |  |  |  | 930.8772 | 15.4747 | -43.9031
320 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 208.2 | 46.3696 | [5 empty] | 0.241 | -12.3597 |  | 2 | -2.1304 | -10.4597 | OK |  |  |  | 863.9004 | 13.3696 | -45.3597
321 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 196.1 | 45.8496 | [5 empty] | 0.301 | -10.4287 |  | 2 | -2.6504 | -8.5287 | OK |  |  |  | 651.495 | 12.8496 | -43.4287
322 |  | GSM 850 |  | monoblock | 2006 | 33 |  |  |  |  | 136.9 | 42.7281 | [5 empty] | 0.43 | -7.3306 |  | 3 | -5.7719 | -5.4306 | Not OK | 0.3413 |  |  | 318.3721 | 9.7281 | -40.3306
323 |  | GSM 850 |  | monoblock | 2006 | 33 |  |  |  |  | 153.7 | 43.7335 | [5 empty] | 0.518 | -5.7134 |  | 2 | -4.7665 | -3.8134 | Not OK | 0.95312 |  |  | 296.7181 | 10.7335 | -38.7134
324 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 205.5 | 46.2562 | [5 empty] | 0.644 | -3.8223 |  | 3 | -2.2438 | -1.9223 | Not OK | 0.32148 |  |  | 319.0994 | 13.2562 | -36.8223
325 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 263.4 | 48.4123 | [5 empty] | 0.297 | -10.5449 |  | 2 | -0.087685 | -8.6449 | OK |  |  |  | 886.8687 | 15.4123 | -43.5449
326 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 209.2 | 46.4112 | [5 empty] | 0.273 | -11.2767 |  | 2 | -2.0888 | -9.3767 | OK |  |  |  | 766.3004 | 13.4112 | -44.2767
327 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 231.3 | 47.2835 | [5 empty] | 0.314 | -10.0614 |  | 2 | -1.2165 | -8.1614 | OK |  |  |  | 736.6242 | 14.2835 | -43.0614
328 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 197.8 | 45.9245 | [5 empty] | 0.281 | -11.0259 |  | 2 | -2.5755 | -9.1259 | OK |  |  |  | 703.9146 | 12.9245 | -44.0259
329 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 209.5 | 46.4237 | [5 empty] | 0.417 | -7.5973 |  | 3 | -2.0763 | -5.6973 | OK |  |  |  | 502.3981 | 13.4237 | -40.5973
330 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 243.5 | 47.73 | [5 empty] | 0.331 | -9.6034 |  | 3 | -0.77002 | -7.7034 | OK |  |  |  | 735.6495 | 14.73 | -42.6034
332 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 252.5 | 48.0452 | [5 empty] | 0.378 | -8.4502 |  | 3 | -0.45477 | -6.5502 | OK |  |  |  | 667.9894 | 15.0452 | -41.4502
333 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 200.5 | 46.0423 | [5 empty] | 0.28 | -11.0568 |  | 3 | -2.4577 | -9.1568 | OK |  |  |  | 716.0714 | 13.0423 | -44.0568
334 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 210.9 | 46.4815 | [5 empty] | 0.296 | -10.5742 |  | 3 | -2.0185 | -8.6742 | OK |  |  |  | 712.5 | 13.4815 | -43.5742
335 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 178.7 | 45.0425 | [5 empty] | 0.54 | -5.3521 |  | 3 | -3.4575 | -3.4521 | Not OK | 0.0053841 |  |  | 330.9259 | 12.0425 | -38.3521
336 |  | GSM 850 |  | Communicator | 2006 | 33 |  |  |  |  | 211.4 | 46.5021 | [5 empty] | 0.29 | -10.752 |  | 2 | -1.9979 | -8.852 | OK |  |  |  | 728.9655 | 13.5021 | -43.752
337 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 77.7 | 37.8084 | [5 empty] | 0.13 | -17.7211 |  | 2 | -10.6916 | -15.8211 | OK |  |  |  | 597.6923 | 4.8084 | -50.7211
338 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 143.5 | 43.137 | [5 empty] | 0.41 | -7.7443 |  | 2 | -5.363 | -5.8443 | OK |  |  |  | 350 | 10.137 | -40.7443
339 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 207.3 | 46.332 | [5 empty] | 0.53 | -5.5145 |  | 1 | -2.168 | -3.6145 | OK |  |  |  | 391.1321 | 13.332 | -38.5145
340 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 246.4 | 47.8328 | [5 empty] | 0.51 | -5.8486 |  | 2 | -0.66719 | -3.9486 | OK |  |  |  | 483.1373 | 14.8328 | -38.8486
341 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 206.8 | 46.311 | [5 empty] | 0.62 | -4.1522 |  | 2 | -2.189 | -2.2522 | OK |  |  |  | 333.5484 | 13.311 | -37.1522
342 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 279 | 48.9121 | [5 empty] | 0.326 | -9.7356 |  | 1 | 0.41208 | -7.8356 | OK |  |  |  | 855.8282 | 15.9121 | -42.7356
345 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 187.9 | 45.4785 | [5 empty] | 0.555 | -5.1141 |  | 1 | 6.9785 | 6.7859 | OK |  |  |  | 338.5586 | 12.4785 | -38.1141
346 |  | GSM 850 |  | Slide | 2006 | 33 |  |  |  |  | 337 | 50.5526 | [5 empty] | 0.423 | -7.4732 |  | 0 | 12.0526 | 4.4268 | OK |  |  |  | 796.6903 | 17.5526 | -40.4732
347 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 195.2 | 45.8096 | [5 empty] | 0.61 | -4.2934 |  | 1 | -2.6904 | -2.3934 | Not OK | 0.297 |  |  | 320 | 12.8096 | -37.2934
348 |  | GSM 850 |  | Slide | 2007 | 33 |  |  |  |  | 235.7 | 47.4472 | [5 empty] | 0.314 | -10.0614 |  | 2 | -1.0528 | -8.1614 | OK |  |  |  | 750.6369 | 14.4472 | -43.0614
349 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 115.4 | 41.2441 | [5 empty] | 0.214 | -13.3917 |  | 3 | -7.2559 | -11.4917 | OK |  |  |  | 539.2523 | 8.2441 | -46.3917
352 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 280.2 | 48.9494 | [5 empty] | 0.399 | -7.9805 |  | 2 | 0.44936 | -6.0805 | OK |  |  |  | 702.2556 | 15.9494 | -40.9805
353 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 229.6 | 47.2194 | [5 empty] | 0.73 | -2.7335 |  | 2 | -1.2806 | -0.83354 | Not OK | 0.44702 |  |  | 314.5205 | 14.2194 | -35.7335
356 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 271 | 48.6594 | [5 empty] | 0.5 | -6.0206 |  | 2 | 0.15939 | -4.1206 | OK |  |  |  | 542 | 15.6594 | -39.0206
357 |  | GSM 850 |  | Swivel | 2006 | 33 |  |  |  |  | 185.6 | 45.3716 | [5 empty] | 0.296 | -10.5742 |  | 2 | -3.1284 | -8.6742 | OK |  |  |  | 627.027 | 12.3716 | -43.5742
358 |  | GSM 850 |  | Slide | 2007 | 33 |  |  |  |  | 242.7 | 47.7014 | [5 empty] | 0.287 | -10.8424 |  | 2 | -0.7986 | -8.9424 | OK |  |  |  | 845.6446 | 14.7014 | -43.8424
360 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 236.2 | 47.4656 | [5 empty] | 0.289 | -10.782 |  | 2 | -1.0344 | -8.882 | OK |  |  |  | 817.301 | 14.4656 | -43.782
362 |  | GSM 850 |  | Fold | 2007 | 33 |  |  |  |  | 220.2 | 46.8563 | [5 empty] | 0.308 | -10.229 |  | 3 | -1.6437 | -8.329 | OK |  |  |  | 714.9351 | 13.8563 | -43.229
363 |  | GSM 850 |  | Slide | 2007 | 33 |  |  |  |  | 229.8 | 47.227 | [5 empty] | 0.339 | -9.396 |  | 2 | -1.273 | -7.496 | OK |  |  |  | 677.8761 | 14.227 | -42.396
364 |  | GSM 850 |  | Block | 2007 | 33 |  |  |  |  | 192.3 | 45.6796 | [5 empty] | 0.328 | -9.6825 |  | 2 | -2.8204 | -7.7825 | OK |  |  |  | 586.2805 | 12.6796 | -42.6825
365 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 186.2 | 45.3996 | [5 empty] | 0.801 | -1.9273 |  | 2 | -3.1004 | -0.02735 | Not OK | 3.0731 |  |  | 232.4594 | 12.3996 | -34.9273
366 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 186.4 | 45.4089 | [5 empty] | 0.227 | -12.8795 |  | 2 | -3.0911 | -10.9795 | OK |  |  |  | 821.1454 | 12.4089 | -45.8795
367 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 279.2 | 48.9183 | [5 empty] | 0.337 | -9.4474 |  | 3 | 0.41831 | -7.5474 | OK |  |  |  | 828.4866 | 15.9183 | -42.4474
368 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 151.4 | 43.6025 | [5 empty] | 0.23 | -12.7654 |  | 2 | -4.8975 | -10.8654 | OK |  |  |  | 658.2609 | 10.6025 | -45.7654
369 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 278.9 | 48.909 | [5 empty] | 0.342 | -9.3195 |  | 2 | 0.40897 | -7.4195 | OK |  |  |  | 815.4971 | 15.909 | -42.3195
370 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 188.9 | 45.5246 | [5 empty] | 0.915 | -0.77158 |  | 1 | -2.9754 | 1.1284 | Not OK | 4.1038 |  |  | 206.4481 | 12.5246 | -33.7716
371 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 187 | 45.4368 | [5 empty] | 0.794 | -2.0036 |  | 1 | -3.0632 | -0.10359 | Not OK | 2.9596 |  |  | 235.5164 | 12.4368 | -35.0036
372 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 199.6 | 46.0032 | [5 empty] | 0.822 | -1.7026 |  | 1 | -2.4968 | 0.19744 | Not OK | 2.6942 |  |  | 242.8224 | 13.0032 | -34.7026
373 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 210.2 | 46.4527 | [5 empty] | 0.896 | -0.95384 |  | 1 | -2.0473 | 0.94616 | Not OK | 2.9935 |  |  | 234.5982 | 13.4527 | -33.9538
374 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 212 | 46.5267 | [5 empty] | 1.04 | 0.34067 |  | 1 | -1.9733 | 2.2407 | Not OK | 4.2139 |  |  | 203.8462 | 13.5267 | -32.6593
375 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 217.3 | 46.7412 | [5 empty] | 1.02 | 0.172 |  |  | -1.7588 | 2.072 | Not OK | 3.8308 |  |  | 213.0392 | 13.7412 | -32.828
376 |  | GSM 850 |  | slide | 2007 | 33 |  |  |  |  | 288.7 | 49.2089 | [5 empty] | 0.357 | -8.9466 |  | 2 | 0.70894 | -7.0466 | OK |  |  |  | 808.6835 | 16.2089 | -41.9466
377 |  | GSM 850 |  | slide | 2007 | 33 |  |  |  |  | 192.8 | 45.7021 | [5 empty] | 0.239 | -12.432 |  | 2 | -2.7979 | -10.532 | OK |  |  |  | 806.6946 | 12.7021 | -45.432
378 |  | GSM 850 |  | slide | 2007 | 33 |  |  |  |  | 267.7 | 48.553 | [5 empty] | 0.342 | -9.3195 |  | 2 | 0.052967 | -7.4195 | OK |  |  |  | 782.7485 | 15.553 | -42.3195
379 |  | GSM 850 |  | slide | 2007 | 33 |  |  |  |  | 154.6 | 43.7842 | [5 empty] | 0.232 | -12.6902 |  | 2 | -4.7158 | -10.7902 | OK |  |  |  | 666.3793 | 10.7842 | -45.6902
381 |  | GSM 850 |  | Dual slider | 2007 | 33 |  |  |  |  | 175.1 | 44.8657 | [5 empty] | 0.31 | -10.1728 |  | 2 | -3.6343 | -8.2728 | OK |  |  |  | 564.8387 | 11.8657 | -43.1728
382 |  | GSM 850 |  | Dual slider | 2007 | 33 |  |  |  |  | 206.8 | 46.311 | [5 empty] | 0.235 | -12.5786 |  | 3 | -2.189 | -10.6786 | OK |  |  |  | 880 | 13.311 | -45.5786
384 |  | GSM 850 |  | Fold | 2006 | 33 |  |  |  |  | 256.9 | 48.1953 | [5 empty] | 0.346 | -9.2185 |  | 2 | -0.30472 | -7.3185 | OK |  |  |  | 742.4855 | 15.1953 | -42.2185
385 |  | GSM 850 |  | Fold | 2007 | 33 |  |  |  |  | 172.5 | 44.7358 | [5 empty] | 0.318 | -9.9515 |  | 2 | -3.7642 | -8.0515 | OK |  |  |  | 542.4528 | 11.7358 | -42.9515
386 |  | GSM 850 |  | Monoblock | 2006 | 33 |  |  |  |  | 215.7 | 46.677 | [5 empty] | 0.341 | -9.3449 |  | 2 | -1.823 | -7.4449 | OK |  |  |  | 632.5513 | 13.677 | -42.3449
387 |  | GSM 850 |  | Fold | 2007 | 33 |  |  |  |  | 245.5 | 47.801 | [5 empty] | 0.326 | -9.7356 |  | 3 | -0.69897 | -7.8356 | OK |  |  |  | 753.0675 | 14.801 | -42.7356
388 |  | GSM 850 |  | Fold | 2007 | 33 |  |  |  |  | 913.1 | 59.2104 | [5 empty] | 0.353 | -9.0445 |  | 2 | 10.7104 | -7.1445 | OK |  |  |  | 2586.6856 | 26.2104 | -42.0445
389 |  | GSM 850 |  | Dual slider | 2007 | 33 |  |  |  |  | 263.4 | 48.4123 | [5 empty] | 0.328 | -9.6825 |  | 2 | -0.087685 | -7.7825 | OK |  |  |  | 803.0488 | 15.4123 | -42.6825
390 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 152.7 | 43.6768 | [5 empty] | 0.247 | -12.1461 |  | 2 | -4.8232 | -10.2461 | OK |  |  |  | 618.2186 | 10.6768 | -45.1461
391 |  | GSM 850 |  | Slide | 2007 | 33 |  |  |  |  | 208.4 | 46.378 | [5 empty] | 0.299 | -10.4866 |  | 2 | -2.122 | -8.5866 | OK |  |  |  | 696.99 | 13.378 | -43.4866
392 |  | GSM 850 |  | Slide | 2007 | 33 |  |  |  |  | 147.7 | 43.3876 | [5 empty] | 0.243 | -12.2879 |  | 2 | -5.1124 | -10.3879 | OK |  |  |  | 607.8189 | 10.3876 | -45.2879
393 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 129.8 | 42.2655 | [5 empty] | 0.395 | -8.0681 |  | 2 | -6.2345 | -6.1681 | Not OK | 0.066448 |  |  | 328.6076 | 9.2655 | -41.0681
394 |  | GSM 850 |  | side slider | 2007 | 33 |  |  |  |  | 108.6 | 40.7166 | [5 empty] | 0.253 | -11.9376 |  | 3 | -7.7834 | -10.0376 | OK |  |  |  | 429.249 | 7.7166 | -44.9376
395 |  | GSM 850 |  | side slider | 2007 | 33 |  |  |  |  | 227.4 | 47.1358 | [5 empty] | 0.326 | -9.7356 |  | 3 | -1.3642 | -7.8356 | OK |  |  |  | 697.546 | 14.1358 | -42.7356
396 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 125.8 | 41.9936 | [5 empty] | 0.176 | -15.0897 |  | 2 | -6.5064 | -13.1897 | OK |  |  |  | 714.7727 | 8.9936 | -48.0897
397 |  | GSM 850 |  | Side slider | 2007 | 33 |  |  |  |  | 218.9 | 46.8049 | [5 empty] | 0.292 | -10.6923 |  | 2 | -1.6951 | -8.7923 | OK |  |  |  | 749.6575 | 13.8049 | -43.6923
398 |  | GSM 850 |  | Fold | 2007 | 33 |  |  |  |  | 201.3 | 46.0769 | [5 empty] | 0.358 | -8.9223 |  | 2 | -2.4231 | -7.0223 | OK |  |  |  | 562.2905 | 13.0769 | -41.9223
399 |  | GSM 850 |  | Fold | 2007 | 33 |  |  |  |  | 234.5 | 47.4029 | [5 empty] | 0.398 | -8.0023 |  | 2 | -1.0971 | -6.1023 | OK |  |  |  | 589.196 | 14.4029 | -41.0023
400 |  | GSM 850 |  | Side Slide | 2007 | 33 |  |  |  |  | 210.6 | 46.4692 | [5 empty] | 0.307 | -10.2572 |  | 3 | -2.0308 | -8.3572 | OK |  |  |  | 685.9935 | 13.4692 | -43.2572
401 |  | GSM 850 |  | Fold | 2007 | 33 |  |  |  |  | 213 | 46.5676 | [5 empty] | 0.295 | -10.6036 |  | 3 | -1.9324 | -8.7036 | OK |  |  |  | 722.0339 | 13.5676 | -43.6036
403 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 191.1 | 45.6252 | [5 empty] | 0.462 | -6.7072 |  | 2 | -2.8748 | -4.8072 | OK |  |  |  | 413.6364 | 12.6252 | -39.7072
404 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 174.6 | 44.8409 | [5 empty] | 0.044 | -27.1309 |  | 2 | -3.6591 | -25.2309 | OK |  |  |  | 3968.1818 | 11.8409 | -60.1309
405 |  | GSM 850 |  | Monoblock | 2007 | 33 |  |  |  |  | 116.7 | 41.3414 | [5 empty] | 0.327 | -9.709 |  | 3 | -7.1586 | -7.809 | OK |  |  |  | 356.8807 | 8.3414 | -42.709
assumed numbers

## Data high band
Product Number | HAC Compliant? | Air Interface | Bands | Form Factor | Standard Version | Max Power dBm | Max Avg. E Field (V/m) | Max Avg. E Field dB(V/m) | After Exclusion E field (V/m) | After Exclusion E field dB(V/m) | PMF Scaled E field (V/m) | PMF Scaled E field dB(V/m) | Rating E field | Max Avg. H Field (A/m) | Max Avg. H Field dB(A/m) | After Exclusion H field (A/m) | After Exclusion H field dB(A/m) | PMF Scaled H field (A/m) | PMF Scaled H field dB(A/m) | Rating H field | E H Ratings | Delta (dB) E-Field relative to M3 (pass) | Delta (dB) H-Field relative to M3 (pass) | OK E-Field Dominate | Amount H-Field is Dominate in dB | Corrected Delta (dB) E-Field relative to M3 (pass) | Corrected Delta (dB) H-Field relative to M3 (pass) | E/H ratio (Ohms) | peak dB(V/m) - peak dBm | peak dB(A/m) - peak dBm
4 |  |  | 1900 |  |  | 25 | 61.07 | 35.7166 | 60.68 | 35.6609 | 62.5 | 35.9176 | M4 | 0.1514 | -16.3975 | 0.1407 | -17.0341 | 0.128 | -17.8558 | M4 | EM4 - HM4 | -5.2834 | -6.9975 | OK |  | -5.0824 | -8.4558 | 488.2812 | 10.9176 | -42.8558
5 |  |  | 1900 |  |  | 25 | 43.03 | 32.6754 | 43 | 32.6694 | 43 | 32.6694 | M4 | 0.1028 | -19.7601 | 0.1021 | -19.8195 | 0.098 | -20.1755 | M4 | EM4 - HM4 | -8.3246 | -10.3601 | OK |  | -8.3306 | -10.7755 | 438.7755 | 7.6694 | -45.1755
6 |  |  | 1900 |  |  | 25 | 47.53 | 33.5394 | 43.1 | 32.6895 | 48.7 | 33.7506 | M4 | 0.1326 | -17.5491 | 0.1228 | -18.216 | 0.124 | -18.1316 | M4 | EM4 - HM4 | -7.4606 | -8.1491 | OK |  | -7.2494 | -8.7316 | 392.7419 | 8.7506 | -43.1316
7 |  |  | 1900 |  |  | 25 | 57.35 | 35.1707 | 51.2 | 34.1854 | 51.2 | 34.1854 | M4 | 0.1586 | -15.9939 | 0.1385 | -17.171 | 0.133 | -17.523 | M4 | EM4 - HM4 | -5.8293 | -6.5939 | OK |  | -6.8146 | -8.123 | 384.9624 | 9.1854 | -42.523
8 |  |  | 1900 |  |  | 25 | 46.75 | 33.3956 | 45.67 | 33.1926 | 47.5 | 33.5339 | M4 | 0.1148 | -18.8012 | 0.1089 | -19.2594 | 0.11 | -19.1721 | M4 | EM4 - HM4 | -7.6044 | -9.4012 | OK |  | -7.4661 | -9.7721 | 431.8182 | 8.5339 | -44.1721
9 |  |  | 1900 |  |  | 24.7 | 51.05 | 34.1599 | 50.62 | 34.0864 | 57.2 | 35.1479 | M4 | 0.1332 | -17.5099 | 0.1208 | -18.3587 | 0.122 | -18.2728 | M4 | EM4 - HM4 |  |  |  |  | 0.98801 | -0.76289 | 468.8525 | 10.4479 | -42.9728
10 |  |  | 1900 |  |  | 25 | 51.7 | 34.2698 | 51.7 | 34.2698 | 58.4 | 35.3283 | M4 | 0.1941 | -14.2395 | 0.1366 | -17.291 | 0.138 | -17.2024 | M4 | EM4 - HM4 | -6.7302 | -4.8395 | Not OK | 1.8907 | -5.6717 | -7.8024 | 423.1884 | 10.3283 | -42.2024
11 |  |  | 1900 |  |  | 25 | 44.96 | 33.0565 | 39.6 | 31.9539 | 41.2 | 32.2979 | M4 | 0.1273 | -17.9034 | 0.1109 | -19.1014 | 0.112 | -19.0156 | M4 | EM4 - HM4 | -7.9435 | -8.5034 | OK |  | -8.7021 | -9.6156 | 367.8571 | 7.2979 | -44.0156
12 |  |  | 1900 |  |  | 25 | 54.58 | 34.7407 | 53.27 | 34.5297 | 55.4 | 34.8702 | M4 | 0.1579 | -16.0324 | 0.1495 | -16.5072 | 0.151 | -16.4205 | M4 | EM4 - HM4 | -6.2593 | -6.6324 | OK |  | -6.1298 | -7.0205 | 366.8874 | 9.8702 | -41.4205
13 |  |  | 1900 |  |  | 24.7 | 48.06 | 33.6357 | 47.98 | 33.6212 | 52.3 | 34.37 | M4 | 0.1099 | -19.18 | 0.1066 | -19.4449 | 0.097 | -20.2646 | M4 | EM4 - HM4 |  |  |  |  | 0.73436 | -1.0845 | 539.1753 | 9.67 | -44.9646
14 |  |  | 1900 |  |  | 25 | 71.75 | 37.1164 | 65.63 | 36.342 | 67.6 | 36.5989 | M3 | 0.2121 | -13.4692 | 0.2121 | -13.4692 | 0.193 | -14.2889 | M3 | EM3 - HM3 | -3.8836 | -4.0692 | OK |  | -4.4011 | -4.8889 | 350.2591 | 11.5989 | -39.2889
15 |  |  | 1900 |  |  | 25 | 63.44 | 36.0473 | 61.28 | 35.7464 | 66.8 | 36.4955 | M3 | 0.1696 | -15.4115 | 0.1407 | -17.0341 | 0.128 | -17.8558 | M4 | EM3 - HM4 | -4.9527 | -6.0115 | OK |  | -4.5045 | -8.4558 | 521.875 | 11.4955 | -42.8558
16 |  |  | 1900 |  |  | 25 | 63.9 | 36.11 | 55.87 | 34.9436 | 60.9 | 35.6923 | M4 | 0.1565 | -16.1097 | 0.1099 | -19.18 | 0.1 | -20 | M4 | EM4 - HM4 | -4.89 | -6.7097 | OK |  | -5.3077 | -10.6 | 609 | 10.6923 | -45
17 |  |  | 1900 |  |  | 25 | 71.89 | 37.1334 | 62.08 | 35.859 | 65.8 | 36.3645 | M3 | 0.1731 | -15.2341 | 0.1243 | -18.1106 | 0.128 | -17.8558 | M4 | EM3 - HM4 | -3.8666 | -5.8341 | OK |  | -4.6355 | -8.4558 | 514.0625 | 11.3645 | -42.8558
18 |  |  | 1900 |  |  | 25 | 51.45 | 34.2277 | 48.02 | 33.6284 | 48.5 | 33.7148 | M4 | 0.1509 | -16.4262 | 0.1067 | -19.4367 | 0.095 | -20.4455 | M4 | EM4 - HM4 | -6.7723 | -7.0262 | OK |  | -7.2852 | -11.0455 | 510.5263 | 8.7148 | -45.4455
19 |  |  | 1900 |  |  | 25 | 92.34 | 39.3078 | 76.85 | 37.7129 | 83 | 38.3816 | M3 | 0.2821 | -10.9919 | 0.2817 | -11.0043 | 0.293 | -10.6626 | M3 | EM3 - HM3 | -1.6922 | -1.5919 | Not OK | 0.10026 | -2.6184 | -1.2626 | 283.2765 | 13.3816 | -35.6626
20 |  |  | 1900 |  |  | 25 | 71.46 | 37.0813 | 63.3 | 36.0281 | 67.1 | 36.5345 | M3 | 0.1855 | -14.6331 | 0.1398 | -17.0899 | 0.144 | -16.8328 | M4 | EM3 - HM4 | -3.9187 | -5.2331 | OK |  | -4.4655 | -7.4328 | 465.9722 | 11.5345 | -41.8328
21 |  |  | 1900 |  |  | 25 | 48.23 | 33.6663 | 48.21 | 33.6627 | 51.1 | 34.1684 | M4 | 0.158 | -16.0269 | 0.133 | -17.523 | 0.137 | -17.2656 | M4 | EM4 - HM4 | -7.3337 | -6.6269 | Not OK | 0.7068 | -6.8316 | -7.8656 | 372.9927 | 9.1684 | -42.2656
22 |  |  | 1900 |  |  | 25 | 46.43 | 33.336 | 46.13 | 33.2797 | 48.9 | 33.7862 | M4 | 0.1386 | -17.1647 | 0.133 | -17.523 | 0.137 | -17.2656 | M4 | EM4 - HM4 | -7.664 | -7.7647 | OK |  | -7.2138 | -7.8656 | 356.9343 | 8.7862 | -42.2656
23 |  |  | 1900 |  |  | 25 | 53.48 | 34.5638 | 52.04 | 34.3267 | 56.2 | 34.9947 | M4 | 0.1635 | -15.7296 | 0.1447 | -16.7906 | 0.149 | -16.5363 | M4 | EM4 - HM4 | -6.4362 | -6.3296 | Not OK | 0.10653 | -6.0053 | -7.1363 | 377.1812 | 9.9947 | -41.5363
24 |  |  | 1900 |  |  | 25 | 46.02 | 33.2589 | 46.02 | 33.2589 | 49.7 | 33.9271 | M4 | 0.1363 | -17.3101 | 0.1049 | -19.5845 | 0.108 | -19.3315 | M4 | EM4 - HM4 | -7.7411 | -7.9101 | OK |  | -7.0729 | -9.9315 | 460.1852 | 8.9271 | -44.3315
25 |  |  | 1900 |  |  | 25 | 40.98 | 32.2514 | 38.61 | 31.734 | 41.7 | 32.4027 | M4 | 0.1135 | -18.9001 | 0.1039 | -19.6677 | 0.107 | -19.4123 | M4 | EM4 - HM4 | -8.7486 | -9.5001 | OK |  | -8.5973 | -10.0123 | 389.7196 | 7.4027 | -44.4123
26 |  |  | 1900 |  |  | 25 | 34.85 | 30.8441 | 34.3 | 30.7059 | 36.7 | 31.2933 | M4 | 0.1372 | -17.2529 | 0.098 | -20.1755 | 0.099 | -20.0873 | M4 | EM4 - HM4 | -10.1559 | -7.8529 | Not OK | 2.303 | -9.7067 | -10.6873 | 370.7071 | 6.2933 | -45.0873
27 |  |  | 1700 |  |  | 25 | 69 | 36.777 | 68.68 | 36.7366 | 72.8 | 37.2426 | M3 | 0.1972 | -14.1019 | 0.1743 | -15.1741 | 0.176 | -15.0897 | M4 | EM3 - HM4 | -4.223 | -4.7019 | OK |  | -3.7574 | -5.6897 | 413.6364 | 12.2426 | -40.0897
[6 empty] | 25 | 55.34 | 34.8608 | 53.36 | 34.5443 | 57.1 | 35.1327 | M4 | 0.1699 | -15.3961 | 0.1257 | -18.0133 | 0.127 | -17.9239 | M4 | EM4 - HM4 | -16.1392 | -15.9961 | Not OK | 0.14308 | -15.8673 | -18.5239 | 449.6063 | 10.1327 | -42.9239
28 |  |  | 1900 |  |  | 24.5 | 60.12 | 35.5804 | 60.09 | 35.576 | 64.3 | 36.1642 | M3 | 0.1729 | -15.2441 | 0.1525 | -16.3346 | 0.154 | -16.2496 | M4 | EM3 - HM4 | -5.4196 | -5.8441 | OK |  | -4.8358 | -6.8496 | 417.5325 | 11.6642 | -40.7496
29 |  |  | 1900 |  |  | 25 | 52.4 | 34.3866 | 47.29 | 33.4954 | 50.6 | 34.083 | M4 | 0.1361 | -17.3228 | 0.1297 | -17.7412 | 0.131 | -17.6546 | M4 | EM4 - HM4 | -6.6134 | -7.9228 | OK |  | -6.917 | -8.2546 | 386.2595 | 9.083 | -42.6546
29 |  |  | 1700 |  |  | 25 | 64.47 | 36.1872 | 58.5 | 35.3431 | 62.6 | 35.9315 | M4 | 0.1448 | -16.7846 | 0.1298 | -17.7345 | 0.135 | -17.3933 | M4 | EM4 - HM4 | -4.8128 | -7.3846 | OK |  | -5.0685 | -7.9933 | 463.7037 | 10.9315 | -42.3933
30 |  |  | 1700 |  |  | 25 | 64.47 | 36.1872 | 58.5 | 35.3431 | 62.6 | 35.9315 | M4 | 0.1448 | -16.7846 | 0.1298 | -17.7345 | 0.135 | -17.3933 | M4 | EM4 - HM4 | -4.8128 | -7.3846 | OK |  | -5.0685 | -7.9933 | 463.7037 | 10.9315 | -42.3933
30 |  |  | 1900 |  |  | 25 | 52.4 | 34.3866 | 47.29 | 33.4954 | 50.6 | 34.083 | M4 | 0.1361 | -17.3228 | 0.1297 | -17.7412 | 0.131 | -17.6546 | M4 | EM4 - HM4 | -6.6134 | -7.9228 | OK |  | -6.917 | -8.2546 | 386.2595 | 9.083 | -42.6546
31 |  |  | 1700 |  |  | 25 | 89.15 | 39.0024 | 84.95 | 38.5833 | 90.9 | 39.1713 | M3 | 0.2023 | -13.8801 | 0.1894 | -14.4524 | 0.197 | -14.1107 | M3 | EM3 - HM3 | -1.9976 | -4.4801 | OK |  | -1.8287 | -4.7107 | 461.4213 | 14.1713 | -39.1107
31 |  |  | 1900 |  |  | 25 | 78.84 | 37.9349 | 75.43 | 37.5509 | 79.2 | 37.9745 | M3 | 0.2008 | -13.9447 | 0.1867 | -14.5771 | 0.196 | -14.1549 | M3 | EM3 - HM3 | -3.0651 | -4.5447 | OK |  | -3.0255 | -4.7549 | 404.0816 | 12.9745 | -39.1549
32 |  | GSM | 1900 |  |  | 29.5 | 24.71 | 27.8575 | 23.86 | 27.5534 | 68 | 36.6502 | M3 | 0.05911 | -24.5668 | 0.0574 | -24.8218 | 0.123 | -18.2019 | M4 | EM3 - HM4 | -10.6425 | -12.6668 | OK |  | -1.8498 | -6.3019 | 552.8455 | 7.1502 | -47.7019
32 |  | WCDMA | 1900 |  |  | 24 | 36.84 | 31.3264 | 36.2 | 31.1742 | 33.3 | 30.4489 | M4 | 0.08865 | -21.0464 | 0.0875 | -21.1598 | 0.077 | -22.2702 | M4 | EM4 - HM4 | -9.6736 | -11.6464 | OK |  | -10.5511 | -12.8702 | 432.4675 | 6.4489 | -46.2702
33 |  | GSM | 1900 |  |  | 30 | 23.41 | 27.388 | 19.83 | 25.9465 | 57.7 | 35.2235 | M3 | 0.08308 | -21.6101 | 0.0623 | -24.1102 | 0.165 | -15.6503 | M3 | EM3 - HM3 | -11.112 | -9.7101 | Not OK | 1.4019 | -3.2765 | -3.7503 | 349.697 | 5.2235 | -45.6503
33 |  | WCDMA | 1900 |  |  | 24 | 38.19 | 31.639 | 32.31 | 30.1867 | 29.4 | 29.3669 | M4 | 0.1234 | -18.1737 | 0.0933 | -20.6024 | 0.084 | -21.5144 | M4 | EM4 - HM4 | -9.361 | -8.7737 | Not OK | 0.58731 | -11.6331 | -12.1144 | 350 | 5.3669 | -45.5144
34 |  | GSM | 1900 |  |  | 30.25 | 27.48 | 28.7803 | 26.69 | 28.527 | 76.6 | 37.6846 | M3 | 0.09378 | -20.5578 | 0.0769 | -22.2815 | 0.203 | -13.8501 | M3 | EM3 - HM3 | -9.7197 | -8.6578 | Not OK | 1.0619 | -0.81542 | -1.9501 | 377.3399 | 7.4346 | -44.1001
34 |  | WCDMA | 1900 |  |  | 24 | 45.63 | 33.185 | 43.37 | 32.7438 | 39.9 | 32.0195 | M4 | 0.1494 | -16.513 | 0.1179 | -18.5697 | 0.105 | -19.5762 | M4 | EM4 - HM4 | -7.815 | -7.113 | Not OK | 0.702 | -8.9805 | -10.1762 | 380 | 8.0195 | -43.5762
35 |  | GSM | 1900 |  |  | 30.25 | 27.48 | 28.7803 | 26.69 | 28.527 | 76.6 | 37.6846 | M3 | 0.09378 | -20.5578 | 0.0769 | -22.2815 | 0.203 | -13.8501 | M3 | EM3 - HM3 | -9.7197 | -8.6578 | Not OK | 1.0619 | -0.81542 | -1.9501 | 377.3399 | 7.4346 | -44.1001
35 |  | WCDMA | 1900 |  |  | 24 | 45.63 | 33.185 | 43.37 | 32.7438 | 39.9 | 32.0195 | M4 | 0.1494 | -16.513 | 0.1179 | -18.5697 | 0.105 | -19.5762 | M4 | EM4 - HM4 | -7.815 | -7.113 | Not OK | 0.702 | -8.9805 | -10.1762 | 380 | 8.0195 | -43.5762
36 |  |  | 1900 |  |  | 30 | 33.04 | 30.3808 | 27.17 | 28.6818 | 81.5 | 38.2232 | M3 | 0.1013 | -19.8878 | 0.0785 | -22.1026 | 0.179 | -14.9429 | M3 | EM3 - HM3 | -8.1192 | -7.9878 | Not OK | 0.13139 | -0.27685 | -3.0429 | 455.3073 | 8.2232 | -44.9429
37 |  |  | 1900 |  |  | 30 | 27.62 | 28.8245 | 23.81 | 27.5352 | 67 | 36.5215 | M3 | 0.08198 | -21.7258 | 0.0603 | -24.3937 | 0.129 | -17.7882 | M4 | EM3 - HM4 | -9.6755 | -9.8258 | OK | FALSE | -1.9785 | -5.8882 | 519.3798 | 6.5215 | -47.7882
38 |  |  | 1900 |  |  | 30.5 | 32.13 | 30.1382 | 26.79 | 28.5595 | 76.9 | 37.7185 | M3 | 0.085 | -21.4116 | 0.0743 | -22.5802 | 0.197 | -14.1107 | M3 | EM3 - HM3 | -8.3618 | -9.5116 | OK | FALSE | -0.78147 | -2.2107 | 390.3553 | 7.2185 | -44.6107
39 |  |  | 1900 |  |  | 29.5 | 27.29 | 28.7201 | 23.41 | 27.388 | 67.2 | 36.5474 | M3 | 0.09441 | -20.4996 | 0.0638 | -23.9036 | 0.169 | -15.4423 | M3 | EM3 - HM3 | -9.7799 | -8.5996 | Not OK | 1.1803 | -1.9526 | -3.5423 | 397.6331 | 7.0474 | -44.9423
40 |  |  | 1900 |  |  | 30 | 29.87 | 29.5047 | 28.05 | 28.9587 | 79.1 | 37.9635 | M3 | 0.1227 | -18.2231 | 0.0775 | -22.214 | 0.2 | -13.9794 | M3 | EM3 - HM3 | -8.9953 | -6.3231 | Not OK | 2.6722 | -0.53647 | -2.0794 | 395.5 | 7.9635 | -43.9794
41 |  |  | 1900 |  |  | 30 | 29.28 | 29.3314 | 28.05 | 28.9587 | 79.1 | 37.9635 | M3 | 0.1091 | -19.2435 | 0.0826 | -21.6604 | 0.213 | -13.4324 | M3 | EM3 - HM3 | -9.1686 | -7.3435 | Not OK | 1.8251 | -0.53647 | -1.5324 | 371.3615 | 7.9635 | -43.4324
42 |  |  | 1900 |  |  | 30.5 | 27.35 | 28.7391 | 25.09 | 27.99 | 73 | 37.2665 | M3 | 0.0782 | -22.1359 | 0.0691 | -23.2104 | 0.183 | -14.751 | M3 | EM3 - HM3 | -9.7609 | -10.2359 | OK | FALSE | -1.2335 | -2.851 | 398.9071 | 6.7665 | -45.251
43 |  |  | 1900 |  |  | 30 | 29.15 | 29.2928 | 28.58 | 29.1212 | 80.6 | 38.1267 | M3 | 0.1083 | -19.3074 | 0.0721 | -22.8413 | 0.186 | -14.6097 | M3 | EM3 - HM3 | -9.2072 | -7.4074 | Not OK | 1.7998 | -0.3733 | -2.7097 | 433.3333 | 8.1267 | -44.6097
44 |  |  | 1900 |  |  | 30.5 | 29.29 | 29.3344 | 28.89 | 29.215 | 82.9 | 38.3711 | M3 | 0.08879 | -21.0327 | 0.0803 | -21.9057 | 0.212 | -13.4733 | M3 | EM3 - HM3 | -9.1656 | -9.1327 | Not OK | 0.032894 | -0.12891 | -1.5733 | 391.0377 | 7.8711 | -43.9733
45 |  |  | 1900 |  |  | 30.5 | 19.24 | 25.6841 | 15.99 | 24.077 | 45.9 | 33.2363 | M4 | 0.05063 | -25.9118 | 0.0481 | -26.3571 | 0.129 | -17.7882 | M4 | EM4 - HM4 | -12.8159 | -14.0118 | OK | FALSE | -5.2637 | -5.8882 | 355.814 | 2.7363 | -48.2882
46 |  |  | 1900 |  |  | 30 | 24.69 | 27.8504 | 23.76 | 27.5169 | 68.2 | 36.6757 | M3 | 0.0811 | -21.8196 | 0.0638 | -23.9036 | 0.171 | -15.3401 | M3 | EM3 - HM3 | -10.6496 | -9.9196 | Not OK | 0.73 | -1.8243 | -3.4401 | 398.8304 | 6.6757 | -45.3401
47 |  |  | 1900 |  |  | 30 | 21.38 | 26.6002 | 19.41 | 25.7605 | 55.7 | 34.9171 | M3 | 0.07417 | -22.5954 | 0.0496 | -26.0904 | 0.131 | -17.6546 | M4 | EM3 - HM4 | -11.8998 | -10.6954 | Not OK | 1.2044 | -3.5829 | -5.7546 | 425.1908 | 4.9171 | -47.6546
54 |  |  | 1900 |  |  | 29.5 | 40.27 | 32.0996 | 32.02 | 30.1084 | 91.9 | 39.2663 | M2 | 0.1215 | -18.3085 | 0.0875 | -21.1598 | 0.231 | -12.7278 | M3 | EM2 - HM3 | -6.4004 | -6.4085 | OK | FALSE | 0.76631 | -0.82776 | 397.8355 | 9.7663 | -42.2278
54 |  |  | 1900 |  |  | 24 | 61.2 | 35.735 | 42.17 | 32.5001 | 38.8 | 31.7766 | M4 | 0.1607 | -15.8797 | 0.1213 | -18.3228 | 0.108 | -19.3315 | M4 | EM4 - HM4 | -5.265 | -6.4797 | OK | FALSE | -9.2234 | -9.9315 | 359.2593 | 7.7766 | -43.3315
55 |  |  | 1900 |  |  | 30 | 44.22 | 32.9124 | 44.22 | 32.9124 | 126 | 42.0074 | M2 | 0.1334 | -17.4969 | 0.1159 | -18.7183 | 0.248 | -12.111 | M2 | EM2 - HM2 | -5.5876 | -5.5969 | OK | FALSE | 3.5074 | -0.21097 | 508.0645 | 12.0074 | -42.111
56 |  |  | 1900 |  |  | 30.4 | 52.83 | 34.4576 | 52.36 | 34.38 | 150.8 | 43.568 | M1 | 0.166 | -15.5978 | 0.1436 | -16.8569 | 0.359 | -8.8981 | M2 | EM1 - HM2 | -4.0424 | -3.6978 | Not OK | 0.34455 | 5.068 | 3.0019 | 420.0557 | 13.168 | -39.2981
57 |  |  | 1900 |  |  | 30.2 | 40.64 | 32.1791 | 40.63 | 32.1769 | 117 | 41.3637 | M2 | 0.1305 | -17.6878 | 0.1065 | -19.453 | 0.247 | -12.1461 | M3 | EM2 - HM3 | -6.3209 | -5.7878 | Not OK | 0.53314 | 2.8637 | -0.24606 | 473.6842 | 11.1637 | -42.3461
58 |  |  | 1900 |  |  | 30 | 44.91 | 33.0469 | 44.7 | 33.0062 | 134.1 | 42.5486 | M2 | 0.1486 | -16.5596 | 0.1263 | -17.9719 | 0.288 | -10.8122 | M2 | EM2 - HM2 | -5.4531 | -4.6596 | Not OK | 0.79352 | 4.0486 | 1.0878 | 465.625 | 12.5486 | -40.8122
59 |  |  | 1900 |  |  | 30.2 | 55.73 | 34.9218 | 50.49 | 34.0641 | 145.4 | 43.2513 | M2 | 0.1509 | -16.4262 | 0.1339 | -17.4644 | 0.3 | -10.4576 | M2 | EM2 - HM2 | -3.5782 | -4.5262 | OK | FALSE | 4.7513 | 1.4424 | 484.6667 | 13.0513 | -40.6576
60 |  |  | 1900 |  |  | 30 | 36.05 | 31.1381 | 26.74 | 28.5432 | 76.2 | 37.6391 | M3 | 0.1174 | -18.6066 | 0.0832 | -21.5975 | 0.178 | -14.9916 | M3 | EM3 - HM3 | -7.3619 | -6.7066 | Not OK | 0.65526 | -0.8609 | -3.0916 | 428.0899 | 7.6391 | -44.9916
61 |  |  | 1900 |  |  | 30.3 | 48.72 | 33.7541 | 30.99 | 29.8244 | 90.2 | 39.1041 | M2 | 0.1296 | -17.7479 | 0.0966 | -20.3005 | 0.256 | -11.8352 | M2 | EM2 - HM2 | -4.7459 | -5.8479 | OK | FALSE | 0.60413 | 0.064799 | 352.3438 | 8.8041 | -42.1352
62 |  |  | 1900 |  |  | 30 | 38.38 | 31.6821 | 30.13 | 29.58 | 90.4 | 39.1234 | M2 | 0.08927 | -20.9859 | 0.0711 | -22.9626 | 0.162 | -15.8097 | M3 | EM2 - HM3 | -6.8179 | -9.0859 | OK | FALSE | 0.62337 | -3.9097 | 558.0247 | 9.1234 | -45.8097
63 |  |  | 1900 |  |  | 30.5 | 38.03 | 31.6025 | 37.38 | 31.4528 | 108.8 | 40.7326 | M2 | 0.1125 | -18.9769 | 0.0928 | -20.649 | 0.246 | -12.1813 | M3 | EM2 - HM3 | -6.8975 | -7.0769 | OK | FALSE | 2.2326 | -0.2813 | 442.2764 | 10.2326 | -42.6813
64 |  |  | 1900 |  |  | 30.5 | 37.1 | 31.3875 | 34.7 | 30.8066 | 104.7 | 40.3989 | M2 | 0.1236 | -18.1596 | 0.0947 | -20.473 | 0.216 | -13.3109 | M3 | EM2 - HM3 | -7.1125 | -6.2596 | Not OK | 0.85289 | 1.8989 | -1.4109 | 484.7222 | 9.8989 | -43.8109
65 |  |  | 1900 |  |  | 30 | 50.61 | 34.0847 | 44.95 | 33.0546 | 130.8 | 42.3322 | M2 | 0.1581 | -16.0214 | 0.1577 | -16.0434 | 0.418 | -7.5765 | M2 | EM2 - HM2 | -4.4153 | -4.1214 | Not OK | 0.29391 | 3.8322 | 4.3235 | 312.9187 | 12.3322 | -37.5765
66 |  |  | 1900 |  |  | 30 | 28.01 | 28.9463 | 22 | 26.8485 | 66 | 36.3909 | M3 | 0.07776 | -22.1849 | 0.057 | -24.8825 | 0.13 | -17.7211 | M4 | EM3 - HM4 | -9.5537 | -10.2849 | OK | FALSE | -2.1091 | -5.8211 | 507.6923 | 6.3909 | -47.7211
67 |  |  | 1900 |  |  |  | 35.79 | 31.0752 | 27.72 | 28.8559 | 79 | 37.9525 | M3 | 0.1202 | -18.4019 | 0.0902 | -20.8959 | 0.193 | -14.2889 | M3 | EM3 - HM3 | -7.4248 | -6.5019 | Not OK | 0.92286 | -0.54746 | -2.3889 | 409.3264 | 37.9525 | -14.2889
68 |  |  | 1900 |  |  | 30 | 31.75 | 30.0349 | 30.52 | 29.6917 | 87.9 | 38.8798 | M2 | 0.1058 | -19.5103 | 0.1058 | -19.5103 | 0.237 | -12.505 | M3 | EM2 - HM3 | -8.4651 | -7.6103 | Not OK | 0.85484 | 0.37978 | -0.60503 | 370.8861 | 8.8798 | -42.505
69 |  |  | 1900 |  |  | 30 | 61.75 | 35.8127 | 40.88 | 32.2302 | 116.5 | 41.3265 | M2 | 0.15 | -16.4782 | 0.1495 | -16.5072 | 0.32 | -9.897 | M2 | EM2 - HM2 | -2.6873 | -4.5782 | OK | FALSE | 2.8265 | 2.003 | 364.0625 | 11.3265 | -39.897
70 |  |  | 1900 |  |  | 30 | 24.9 | 27.924 | 24.04 | 27.6187 | 69 | 36.777 | M3 | 0.09421 | -20.5181 | 0.0649 | -23.7551 | 0.174 | -15.189 | M3 | EM3 - HM3 | -10.576 | -8.6181 | Not OK | 1.958 | -1.723 | -3.289 | 396.5517 | 6.777 | -45.189
71 |  |  | 1900 |  |  | 29.7 | 24.55 | 27.801 | 23.89 | 27.5643 | 68.8 | 36.7518 | M3 | 0.08254 | -21.6667 | 0.0645 | -23.8088 | 0.169 | -15.4423 | M3 | EM3 - HM3 | -10.699 | -9.7667 | Not OK | 0.93226 | -1.7482 | -3.5423 | 407.1006 | 7.0518 | -45.1423
72 |  |  | 1900 |  |  | 30 | 43.79 | 32.8275 | 39.96 | 32.0325 | 112.7 | 41.0385 | M2 | 0.1416 | -16.9787 | 0.1078 | -19.3476 | 0.278 | -11.1191 | M2 | EM2 - HM2 | -5.6725 | -5.0787 | Not OK | 0.59377 | 2.5385 | 0.7809 | 405.3957 | 11.0385 | -41.1191
73 |  |  | 1900 |  |  | 30.5 | 49.14 | 33.8287 | 47.7 | 33.5704 | 136.9 | 42.7281 | M2 | 0.1734 | -15.219 | 0.1348 | -17.4062 | 0.356 | -8.971 | M2 | EM2 - HM2 | -4.6713 | -3.319 | Not OK | 1.3523 | 4.2281 | 2.929 | 384.5506 | 12.2281 | -39.471
[6 empty] | Peak / pulsed averaged channel power (dBm) | [15 empty] | Delta (dB) E-Field relative to M3 (pass) | Delta (dB) H-Field relative to M3 (pass)
103 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 96.6 | 39.6995 | 2 |  |  |  |  | 0.281 | -11.0259 | 2 |  | 1.1995 | 0.87413 | OK |  |  |  | 343.7722 | 9.6995 | -41.0259
104 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 112.9 | 41.0539 | 2 |  |  |  |  | 0.308 | -10.229 | 2 |  | 2.5539 | 1.671 | OK |  |  |  | 366.5584 | 11.0539 | -40.229
105 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 114.9 | 41.2064 | 2 |  |  |  |  | 0.303 | -10.3711 | 2 |  | 2.7064 | 1.5289 | OK |  |  |  | 379.2079 | 11.2064 | -40.3711
106 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 147.4 | 43.3699 | 2 |  |  |  |  | 0.37 | -8.636 | 2 |  | 4.8699 | 3.264 | OK |  |  |  | 398.3784 | 13.3699 | -38.636
107 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 82.68 | 38.348 | 3 |  |  |  |  | 0.23 | -12.7654 | 3 |  | -0.15199 | -0.86544 | OK |  |  |  | 359.4783 | 8.348 | -42.7654
108 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 82.68 | 38.348 | 3 |  |  |  |  | 0.23 | -12.7654 | 3 |  | -0.15199 | -0.86544 | OK |  |  |  | 359.4783 | 8.348 | -42.7654
109 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 132.1 | 42.4181 | 2 |  |  |  |  | 0.327 | -9.709 | 2 |  | 3.9181 | 2.191 | OK |  |  |  | 403.9755 | 12.4181 | -39.709
110 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 105.6 | 40.4733 | 2 |  |  |  |  | 0.284 | -10.9336 | 2 |  | 1.9733 | 0.96637 | OK |  |  |  | 371.831 | 10.4733 | -40.9336
111 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 128.6 | 42.1848 | 2 |  |  |  |  | 0.32 | -9.897 | 2 |  | 3.6848 | 2.003 | OK |  |  |  | 401.875 | 12.1848 | -39.897
112 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 158.2 | 43.9841 | 1 |  |  |  |  | 0.43 | -7.3306 | 2 |  | 5.4841 | 4.5694 | OK |  |  |  | 367.907 | 13.9841 | -37.3306
113 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 67.6 | 36.5989 | 3 |  |  |  |  | 0.17 | -15.391 | 3 |  | -1.9011 | -3.491 | OK |  |  |  | 397.6471 | 6.5989 | -45.391
114 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 103.2 | 40.2736 | 2 |  |  |  |  | 0.29 | -10.752 | 2 |  | 1.7736 | 1.148 | OK |  |  |  | 355.8621 | 10.2736 | -40.752
115 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 147.6 | 43.3817 | 2 |  |  |  |  | 0.39 | -8.1787 | 2 |  | 4.8817 | 3.7213 | OK |  |  |  | 378.4615 | 13.3817 | -38.1787
116 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 154.5 | 43.7786 | 1 |  |  |  |  | 0.46 | -6.7448 | 1 |  | 5.2786 | 5.1552 | OK |  |  |  | 335.8696 | 13.7786 | -36.7448
117 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 153 | 43.6938 | 1 |  |  |  |  | 0.448 | -6.9744 | 2 |  | 5.1938 | 4.9256 | OK |  |  |  | 341.5179 | 13.6938 | -36.9744
118 | Pass | GSM 1900 |  | Candy Bar PDA | 2006 | 30 |  |  |  |  | 79.1 | 37.9635 | 3 |  |  |  |  | 0.22 | -13.1515 | 3 | 3 | -0.53647 | -1.2515 | OK |  |  |  | 359.5455 | 7.9635 | -43.1515
119 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 97.2 | 39.7533 | 2 |  |  |  |  | 0.287 | -10.8424 | 2 |  | 1.2533 | 1.0576 | OK |  |  |  | 338.676 | 9.7533 | -40.8424
120 |  | GSM 1900 |  |  | 2006 | 29 |  |  |  |  | 81.6 | 38.2338 | 3 |  |  |  |  | 0.242 | -12.3237 | 3 |  | -0.2662 | -0.42369 | OK |  |  |  | 337.1901 | 9.2338 | -41.3237
121 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 81.3 | 38.2018 | 3 |  |  |  |  | 0.247 | -12.1461 | 3 |  | -0.29819 | -0.24606 | Not OK | 0.052128 |  |  | 329.1498 | 8.2018 | -42.1461
122 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 86.9 | 38.7804 | 2 |  |  |  |  | 0.287 | -10.8424 | 2 |  | 0.2804 | 1.0576 | Not OK | 0.77724 |  |  | 302.7875 | 8.7804 | -40.8424
123 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 91.52 | 39.2303 | 2 |  |  |  |  | 0.26 | -11.7005 | 2 |  | 0.73032 | 0.19947 | OK |  |  |  | 352 | 9.2303 | -41.7005
124 | Fail | GSM 1900 |  | Candy Bar PDA | 2006 | 30 |  |  |  |  | 108.8 | 40.7326 | 2 |  |  |  |  | 0.343 | -9.2941 | 2 | 2 | 2.2326 | 2.6059 | Not OK | 0.3733 |  |  | 317.2012 | 10.7326 | -39.2941
125 | Fail | GSM 1900 |  | Candy Bar PDA | 2006 | 30 |  |  |  |  | 96.4 | 39.6815 | 2 |  |  |  |  | 0.319 | -9.9242 | 2 | 2 | 1.1815 | 1.9758 | Not OK | 0.79427 |  |  | 302.1944 | 9.6815 | -39.9242
126 |  | GSM 1900 |  |  | 2006 | 30 |  |  |  |  | 145.1 | 43.2333 | 2 |  |  |  |  | 0.39 | -8.1787 | 2 |  | 4.7333 | 3.7213 | OK |  |  |  | 372.0513 | 13.2333 | -38.1787
133 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 36.2 | 31.1742 | 4 |  |  |  |  | 0.1 | -20 | 4 |  | -9.8258 | -10.6 | OK |  |  |  | 362 | 7.6742 | -43.5
134 |  | WCDMA FDD II 1900 |  |  | 2006 | 24 |  |  |  |  | 47.9 | 33.6067 | 4 |  |  |  |  | 0.117 | -18.6363 | 4 |  | -7.3933 | -9.2363 | OK |  |  |  | 409.4017 | 9.6067 | -42.6363
135 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 62.8 | 35.9592 | 4 |  |  |  |  | 0.161 | -15.8635 | 4 |  | -5.0408 | -6.4635 | OK |  |  |  | 390.0621 | 12.4592 | -39.3635
136 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 68.2 | 36.6757 | 3 |  |  |  |  | 0.15 | -16.4782 | 4 |  | -4.3243 | -7.0782 | OK |  |  |  | 454.6667 | 13.1757 | -39.9782
137 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 48.7 | 33.7506 | 4 |  |  |  |  | 0.12 | -18.4164 | 4 |  | -7.2494 | -9.0164 | OK |  |  |  | 405.8333 | 10.2506 | -41.9164
138 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 72.2 | 37.1707 | 3 |  |  |  |  | 0.191 | -14.3793 | 3 |  | -3.8293 | -4.9793 | OK |  |  |  | 378.0105 | 13.6707 | -37.8793
139 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 34.9 | 30.8565 | 4 |  |  |  |  | 0.092 | -20.7242 | 4 |  | -10.1435 | -11.3242 | OK |  |  |  | 379.3478 | 7.3565 | -44.2242
140 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 38.8 | 31.7766 | 4 |  |  |  |  | 0.133 | -17.523 | 4 |  | -9.2234 | -8.123 | Not OK | 1.1004 |  |  | 291.7293 | 8.2766 | -41.023
141 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 36.2 | 31.1742 | 4 |  |  |  |  | 0.128 | -17.8558 | 4 |  | -9.8258 | -8.4558 | Not OK | 1.37 |  |  | 282.8125 | 7.6742 | -41.3558
142 |  | CDMA 1900 |  |  | 2006 | 23.5 |  |  |  |  | 33.9 | 30.604 | 4 |  |  |  |  | 0.119 | -18.4891 | 4 |  | -10.396 | -9.0891 | Not OK | 1.3069 |  |  | 284.8739 | 7.104 | -41.9891
201 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 128.4 | 42.1713 | 2 |  |  |  |  | 0.302 | -10.3999 | 2 |  | 3.6713 | 1.5001 | OK |  |  |  | 425.1656 | 12.1713 | -40.3999
202 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 75.98 | 37.614 | 3 |  |  |  |  | 0.207 | -13.6806 | 3 |  | -0.88601 | -1.7806 | OK |  |  |  | 367.0531 | 7.614 | -43.6806
203 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 67.1 | 36.5345 | 3 |  |  |  |  | 0.216 | -13.3109 | 3 |  | -1.9655 | -1.4109 | Not OK | 0.55462 |  |  | 310.6481 | 6.5345 | -43.3109
203 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 25.4 | 28.0967 | 4 |  |  |  |  | 0.089 | -21.0122 | 4 |  | -12.9033 | -11.6122 | Not OK | 1.2911 |  |  | 285.3933 | 4.5967 | -44.5122
204 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 72.7 | 37.2307 | 3 |  |  |  |  | 0.225 | -12.9563 | 3 |  | -1.2693 | -1.0563 | Not OK | 0.21296 |  |  | 323.1111 | 7.2307 | -42.9563
205 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 75.6 | 37.5704 | 3 |  |  |  |  | 0.228 | -12.8413 | 3 |  | -0.92956 | -0.9413 | OK |  |  |  | 331.5789 | 7.5704 | -42.8413
205 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 36.7 | 31.2933 | 4 |  |  |  |  | 0.104 | -19.6593 | 4 |  | -9.7067 | -10.2593 | OK |  |  |  | 352.8846 | 7.7933 | -43.1593
206 |  | GSM 1900 |  |  | 2007 | 30 |  |  |  |  | 70.4 | 36.9515 | 3 |  |  |  |  | 0.175 | -15.1392 | 3 |  | -1.5485 | -3.2392 | OK |  |  |  | 402.2857 | 6.9515 | -45.1392
208 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 31.6 | 29.9937 | 4 |  |  |  |  | 0.085 | -21.4116 | 4 |  | -11.0063 | -12.0116 | OK |  |  |  | 371.7647 | 6.4937 | -44.9116
209 |  | CDMA 1900 |  |  | 2007 | 23.5 |  |  |  |  | 49.8 | 33.9446 | 4 |  |  |  |  | 0.142 | -16.9542 | 4 |  | -7.0554 | -7.5542 | OK |  |  |  | 350.7042 | 10.4446 | -40.4542
300 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 104.5 | 40.3823 | [5 empty] | 0.34 | -9.3704 |  |  | 1.8823 | 2.5296 | Not OK | 0.64725 |  |  | 307.3529 | 10.3823 | -39.3704
301 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 79.8 | 38.0401 | [5 empty] | 0.155 | -16.1934 |  |  | -0.45994 | -4.2934 | OK |  |  |  | 514.8387 | 8.0401 | -46.1934
302 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 133.5 | 42.5096 | [5 empty] | 0.376 | -8.4962 |  |  | 4.0096 | 3.4038 | OK |  |  |  | 355.0532 | 12.5096 | -38.4962
303 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 106.8 | 40.5714 | [5 empty] | 0.401 | -7.9371 |  |  | 2.0714 | 3.9629 | Not OK | 1.8915 |  |  | 266.3342 | 10.5714 | -37.9371
305 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 118.3 | 41.4597 | [5 empty] | 0.34 | -9.3704 |  |  | 2.9597 | 2.5296 | OK |  |  |  | 347.9412 | 11.4597 | -39.3704
307 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 119.3 | 41.5328 | [5 empty] | 0.289 | -10.782 |  |  | 3.0328 | 1.118 | OK |  |  |  | 412.8028 | 11.5328 | -40.782
311 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 101.4 | 40.1208 | [5 empty] | 0.164 | -15.7031 |  |  | 1.6208 | -3.8031 | OK |  |  |  | 618.2927 | 10.1208 | -45.7031
312 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 117.5 | 41.4008 | [5 empty] | 0.288 | -10.8122 |  |  | 2.9008 | 1.0878 | OK |  |  |  | 407.9861 | 11.4008 | -40.8122
313 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 107.6 | 40.6362 | [5 empty] | 0.218 | -13.2309 |  |  | 2.1362 | -1.3309 | OK |  |  |  | 493.578 | 10.6362 | -43.2309
314 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 105.4 | 40.4568 | [5 empty] | 0.317 | -9.9788 |  |  | 1.9568 | 1.9212 | OK |  |  |  | 332.4921 | 10.4568 | -39.9788
315 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 109.6 | 40.7962 | [5 empty] | 0.335 | -9.4991 |  |  | 2.2962 | 2.4009 | Not OK | 0.10469 |  |  | 327.1642 | 10.7962 | -39.4991
316 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 139.2 | 42.8728 | [5 empty] | 0.211 | -13.5144 |  |  | 4.3728 | -1.6144 | OK |  |  |  | 659.7156 | 12.8728 | -43.5144
317 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 138.6 | 42.8353 | [5 empty] | 0.45 | -6.9357 |  |  | 4.3353 | 4.9643 | Not OK | 0.62899 |  |  | 308 | 12.8353 | -36.9357
318 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 49 | 33.8039 | [5 empty] | 0.172 | -15.2894 |  |  | -4.6961 | -3.3894 | Not OK | 1.3066 |  |  | 284.8837 | 3.8039 | -45.2894
319 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 171 | 44.6599 | [5 empty] | 0.364 | -8.778 |  |  | 6.1599 | 3.122 | OK | FALSE |  |  | 469.7802 | 14.6599 | -38.778
320 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 108.8 | 40.7326 | [5 empty] | 0.26 | -11.7005 |  |  | 2.2326 | 0.19947 | OK |  |  |  | 418.4615 | 10.7326 | -41.7005
321 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 91.1 | 39.1904 | [5 empty] | 0.242 | -12.3237 |  |  | 0.69037 | -0.42369 | OK |  |  |  | 376.4463 | 9.1904 | -42.3237
322 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 73.9 | 37.3729 | [5 empty] | 0.213 | -13.4324 |  |  | -1.1271 | -1.5324 | OK |  |  |  | 346.9484 | 7.3729 | -43.4324
323 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 86.8 | 38.7704 | [5 empty] | 0.278 | -11.1191 |  |  | 0.27039 | 0.7809 | Not OK | 0.5105 |  |  | 312.2302 | 8.7704 | -41.1191
324 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 72 | 37.1466 | [5 empty] | 0.179 | -14.9429 |  |  | -1.3534 | -3.0429 | OK |  |  |  | 402.2346 | 7.1466 | -44.9429
325 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 124.5 | 41.9034 | [5 empty] | 0.308 | -10.229 |  |  | 3.4034 | 1.671 | OK |  |  |  | 404.2208 | 11.9034 | -40.229
326 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 108.2 | 40.6845 | [5 empty] | 0.317 | -9.9788 |  |  | 2.1845 | 1.9212 | OK |  |  |  | 341.3249 | 10.6845 | -39.9788
327 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 86.7 | 38.7604 | [5 empty] | 0.24 | -12.3958 |  |  | 0.26038 | -0.49578 | OK |  |  |  | 361.25 | 8.7604 | -42.3958
328 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 95.6 | 39.6092 | [5 empty] | 0.24 | -12.3958 |  |  | 1.1092 | -0.49578 | OK |  |  |  | 398.3333 | 9.6092 | -42.3958
329 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 77.1 | 37.7411 | [5 empty] | 0.221 | -13.1122 |  |  | -0.75891 | -1.2122 | OK |  |  |  | 348.8688 | 7.7411 | -43.1122
330 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 79.7 | 38.0292 | [5 empty] | 0.185 | -14.6566 |  |  | -0.47083 | -2.7566 | OK |  |  |  | 430.8108 | 8.0292 | -44.6566
331 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 83.2 | 38.4025 | [5 empty] | 0.168 | -15.4938 |  | 3 | -0.097533 | -3.5938 | OK |  |  |  | 495.2381 | 8.4025 | -45.4938
332 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 83.2 | 38.4025 | [5 empty] | 0.168 | -15.4938 |  |  | -0.097533 | -3.5938 | OK |  |  |  | 495.2381 | 8.4025 | -45.4938
333 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 73.4 | 37.3139 | [5 empty] | 0.206 | -13.7227 |  |  | -1.1861 | -1.8227 | OK |  |  |  | 356.3107 | 7.3139 | -43.7227
334 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 78.4 | 37.8863 | [5 empty] | 0.215 | -13.3512 |  |  | -0.61368 | -1.4512 | OK |  |  |  | 364.6512 | 7.8863 | -43.3512
335 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 92.3 | 39.304 | [5 empty] | 0.39 | -8.1787 |  |  | 0.80403 | 3.7213 | Not OK | 2.9173 |  |  | 236.6667 | 9.304 | -38.1787
336 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 96.5 | 39.6905 | [5 empty] | 0.27 | -11.3727 |  |  | 1.1905 | 0.52728 | OK | FALSE |  |  | 357.4074 | 9.6905 | -41.3727
337 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 65.2 | 36.285 | [5 empty] | 0.29 | -10.752 |  |  | -2.215 | 1.148 | Not OK | 3.363 |  |  | 224.8276 | 6.285 | -40.752
338 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 97.9 | 39.8157 | [5 empty] | 0.29 | -10.752 |  |  | 1.3157 | 1.148 | OK |  |  |  | 337.5862 | 9.8157 | -40.752
339 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 112.3 | 41.0076 | [5 empty] | 0.47 | -6.558 |  |  | 2.5076 | 5.342 | Not OK | 2.8344 |  |  | 238.9362 | 11.0076 | -36.558
340 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 118.8 | 41.4963 | [5 empty] | 0.41 | -7.7443 |  |  | 2.9963 | 4.1557 | Not OK | 1.1593 |  |  | 289.7561 | 11.4963 | -37.7443
341 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 93.5 | 39.4162 | [5 empty] | 0.44 | -7.1309 |  |  | 0.91623 | 4.7691 | Not OK | 3.8528 |  |  | 212.5 | 9.4162 | -37.1309
342 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 150 | 43.5218 | [5 empty] | 0.649 | -3.7551 |  |  | 5.0218 | 8.1449 | Not OK | 3.1231 |  |  | 231.1248 | 13.5218 | -33.7551
345 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 130.8 | 42.3322 | [5 empty] | 1.15 | 1.214 |  |  | -6.1678 | 3.114 | Not OK | 9.2818 |  |  | 113.7391 | 12.3322 | -28.786
347 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 102.5 | 40.2145 | [5 empty] | 0.542 | -5.32 |  |  | 1.7145 | 6.58 | Not OK | 4.8655 |  |  | 189.1144 | 10.2145 | -35.32
348 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 130.1 | 42.2855 | [5 empty] | 0.385 | -8.2908 |  |  | 3.7855 | 3.6092 | OK | FALSE |  |  | 337.9221 | 12.2855 | -38.2908
349 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 56.3 | 35.0102 | [5 empty] | 0.109 | -19.2515 |  |  | -3.4898 | -7.3515 | OK |  |  |  | 516.5138 | 5.0102 | -49.2515
352 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 135.1 | 42.6131 | [5 empty] | 0.361 | -8.8499 |  |  | 4.1131 | 3.0501 | OK |  |  |  | 374.2382 | 12.6131 | -38.8499
353 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 95 | 39.5545 | [5 empty] | 0.34 | -9.3704 |  |  | 1.0545 | 2.5296 | Not OK | 1.4751 |  |  | 279.4118 | 9.5545 | -39.3704
354 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 75.3 | 37.5359 | [5 empty] | 0.176 | -15.0897 |  | 2 | -0.9641 | -3.1897 | OK |  |  |  | 427.8409 | 7.5359 | -45.0897
356 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 74.5 | 37.4431 | [5 empty] | 0.2 | -13.9794 |  |  | -1.0569 | -2.0794 | OK |  |  |  | 372.5 | 7.4431 | -43.9794
357 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 95 | 39.5545 | [5 empty] | 0.16 | -15.9176 |  |  | 1.0545 | -4.0176 | OK |  |  |  | 593.75 | 9.5545 | -45.9176
358 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 150.3 | 43.5392 | [5 empty] | 0.442 | -7.0916 |  |  | 5.0392 | 4.8084 | OK |  |  |  | 340.0452 | 13.5392 | -37.0916
359 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 125.9 | 42.0005 | [5 empty] | 0.343 | -9.2941 |  | 1 | 3.5005 | 2.6059 | OK |  |  |  | 367.0554 | 12.0005 | -39.2941
360 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 142.2 | 43.058 | [5 empty] | 0.366 | -8.7304 |  |  | 2.3436 | 3.1696 | Not OK | 0.82599 |  |  | 388.5246 | 13.058 | -38.7304
361 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 110.2 | 40.8436 | [5 empty] | 0.32 | -9.897 |  | 2 | -0.24493 | 2.003 | Not OK | 2.2479 |  |  | 344.375 | 10.8436 | -39.897
362 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 81.8 | 38.2551 | [5 empty] | 0.272 | -11.3086 |  |  | -0.24493 | 0.59138 | Not OK | 0.83631 |  |  | 300.7353 | 8.2551 | -41.3086
363 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 134.8 | 42.5938 | [5 empty] | 0.428 | -7.3711 |  |  | 4.0938 | 4.5289 | Not OK | 0.43508 |  |  | 314.9533 | 12.5938 | -37.3711
364 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 93.2 | 39.3883 | [5 empty] | 0.305 | -10.314 |  |  | 0.88832 | 1.586 | Not OK | 0.69768 |  |  | 305.5738 | 9.3883 | -40.314
365 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 133.4 | 42.5031 | [5 empty] | 0.548 | -5.2244 |  |  | 4.0031 | 6.6756 | Not OK | 2.6725 |  |  | 243.4307 | 12.5031 | -35.2244
366 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 55.5 | 34.8859 | [5 empty] | 0.203 | -13.8501 |  |  | -3.6141 | -1.9501 | Not OK | 1.6641 |  |  | 273.399 | 4.8859 | -43.8501
367 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 110.7 | 40.883 | [5 empty] | 0.393 | -8.1121 |  |  | 2.383 | 3.7879 | Not OK | 1.4049 |  |  | 281.6794 | 10.883 | -38.1121
368 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 133 | 42.477 | [5 empty] | 0.459 | -6.7637 |  |  | 3.977 | 5.1363 | Not OK | 1.1592 |  |  | 289.7603 | 12.477 | -36.7637
369 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 144.1 | 43.1733 | [5 empty] | 0.505 | -5.9342 |  |  | 4.6733 | 5.9658 | Not OK | 1.2925 |  |  | 285.3465 | 13.1733 | -35.9342
370 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 155.2 | 43.8178 | [5 empty] | 0.662 | -3.5828 |  |  | 5.3178 | 8.3172 | Not OK | 2.9993 |  |  | 234.4411 | 13.8178 | -33.5828
371 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 149.6 | 43.4986 | [5 empty] | 0.617 | -4.1943 |  |  | 4.9986 | 7.7057 | Not OK | 2.7071 |  |  | 242.4635 | 13.4986 | -34.1943
372 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 151.1 | 43.5853 | [5 empty] | 0.581 | -4.7165 |  |  | 5.0853 | 7.1835 | Not OK | 2.0982 |  |  | 260.0688 | 13.5853 | -34.7165
373 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 136.1 | 42.6772 | [5 empty] | 0.601 | -4.4225 |  |  | 4.1772 | 7.4775 | Not OK | 3.3003 |  |  | 226.4559 | 12.6772 | -34.4225
374 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 164.6 | 44.3286 | [5 empty] | 0.614 | -4.2366 |  |  | 5.8286 | 7.6634 | Not OK | 1.8348 |  |  | 268.0782 | 14.3286 | -34.2366
375 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 133.1 | 42.4836 | [5 empty] | 0.479 | -6.3933 |  |  | 3.9836 | 5.5067 | Not OK | 1.5231 |  |  | 277.8706 | 12.4836 | -36.3933
376 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 112.1 | 40.9921 | [5 empty] | 0.29 | -10.752 |  |  | 2.4921 | 1.148 | OK |  |  |  | 386.5517 | 10.9921 | -40.752
377 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 86.7 | 38.7604 | [5 empty] | 0.23 | -12.7654 |  |  | 0.26038 | -0.86544 | OK |  |  |  | 376.9565 | 8.7604 | -42.7654
378 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 102.4 | 40.206 | [5 empty] | 0.223 | -13.0339 |  |  | 1.706 | -1.1339 | OK |  |  |  | 459.1928 | 10.206 | -43.0339
379 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 88.6 | 38.9487 | [5 empty] | 0.223 | -13.0339 |  |  | 0.44867 | -1.1339 | OK |  |  |  | 397.3094 | 8.9487 | -43.0339
380 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 114.8 | 41.1988 | [5 empty] | 0.326 | -9.7356 |  | 2 | 2.6988 | 2.1644 | OK |  |  |  | 352.1472 | 11.1988 | -39.7356
381 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 85 | 38.5884 | [5 empty] | 0.239 | -12.432 |  |  | 0.088379 | -0.53204 | OK |  |  |  | 355.6485 | 8.5884 | -42.432
382 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 83.5 | 38.4337 | [5 empty] | 0.152 | -16.3631 |  |  | -0.06627 | -4.4631 | OK |  |  |  | 549.3421 | 8.4337 | -46.3631
383 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 74.2 | 37.4081 | [5 empty] | 0.147 | -16.6537 |  |  | -1.0919 | -4.7537 | OK |  |  |  | 504.7619 | 7.4081 | -46.6537
384 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 94 | 39.4626 | [5 empty] | 0.263 | -11.6009 |  |  | 0.96256 | 0.29911 | OK |  |  |  | 357.4144 | 9.4626 | -41.6009
385 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 76.6 | 37.6846 | [5 empty] | 0.293 | -10.6626 |  |  | -0.81542 | 1.2374 | Not OK | 2.0528 |  |  | 261.4334 | 7.6846 | -40.6626
386 |  | GSM1900 |  |  | 2006 | 30 |  |  |  |  | 91.8 | 39.2569 | [5 empty] | 0.251 | -12.0065 |  |  | 0.75685 | -0.10653 | OK |  |  |  | 365.7371 | 9.2569 | -42.0065
387 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 77.6 | 37.7972 | [5 empty] | 0.206 | -13.7227 |  |  | -0.70277 | -1.8227 | OK |  |  |  | 376.699 | 7.7972 | -43.7227
388 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 65.6 | 36.3381 | [5 empty] | 0.262 | -11.634 |  |  | -2.1619 | 0.26603 | Not OK | 2.4279 |  |  | 250.3817 | 6.3381 | -41.634
389 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 102.3 | 40.1975 | [5 empty] | 0.193 | -14.2889 |  |  | 1.6975 | -2.3889 | OK |  |  |  | 530.0518 | 10.1975 | -44.2889
390 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 117.7 | 41.4155 | [5 empty] | 0.297 | -10.5449 |  |  | 2.9155 | 1.3551 | OK |  |  |  | 396.2963 | 11.4155 | -40.5449
391 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 131.8 | 42.3983 | [5 empty] | 0.342 | -9.3195 |  |  | 3.8983 | 2.5805 | OK |  |  |  | 385.3801 | 12.3983 | -39.3195
392 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 85 | 38.5884 | [5 empty] | 0.231 | -12.7278 |  |  | 0.088379 | -0.82776 | OK |  |  |  | 367.9654 | 8.5884 | -42.7278
393 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 86.3 | 38.7202 | [5 empty] | 0.246 | -12.1813 |  |  | 0.22022 | -0.2813 | OK |  |  |  | 350.813 | 8.7202 | -42.1813
394 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 52.7 | 34.4362 | [5 empty] | 0.141 | -17.0156 |  |  | -4.0638 | -5.1156 | OK |  |  |  | 373.7589 | 4.4362 | -47.0156
395 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 124.3 | 41.8894 | [5 empty] | 0.414 | -7.66 |  |  | 3.3894 | 4.24 | Not OK | 0.85058 |  |  | 300.2415 | 11.8894 | -37.66
396 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 85.4 | 38.6292 | [5 empty] | 0.195 | -14.1993 |  |  | 0.12916 | -2.2993 | OK |  |  |  | 437.9487 | 8.6292 | -44.1993
397 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 65.6 | 36.3381 | [5 empty] | 0.176 | -15.0897 |  |  | -2.1619 | -3.1897 | OK |  |  |  | 372.7273 | 6.3381 | -45.0897
398 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 68.5 | 36.7138 | [5 empty] | 0.243 | -12.2879 |  |  | -1.7862 | -0.38787 | Not OK | 1.3983 |  |  | 281.893 | 6.7138 | -42.2879
399 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 106.2 | 40.5225 | [5 empty] | 0.238 | -12.4685 |  |  | 2.0225 | -0.56846 | OK |  |  |  | 446.2185 | 10.5225 | -42.4685
400 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 67.8 | 36.6246 | [5 empty] | 0.223 | -13.0339 |  |  | -1.8754 | -1.1339 | Not OK | 0.7415 |  |  | 304.0359 | 6.6246 | -43.0339
401 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 68.7 | 36.7391 | [5 empty] | 0.201 | -13.9361 |  |  | -1.7609 | -2.0361 | OK |  |  |  | 341.791 | 6.7391 | -43.9361
403 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 98.5 | 39.8687 | [5 empty] | 0.231 | -12.7278 |  |  | 1.3687 | -0.82776 | OK |  |  |  | 426.4069 | 9.8687 | -42.7278
404 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 134.5 | 42.5744 | [5 empty] | 0.038 | -28.4043 |  |  | 4.0744 | -16.5043 | OK |  |  |  | 3539.4737 | 12.5744 | -58.4043
405 |  | GSM1900 |  |  | 2007 | 30 |  |  |  |  | 71 | 37.0252 | [5 empty] | 0.204 | -13.8074 |  |  | -1.4748 | -1.9074 | OK |  |  |  | 348.0392 | 7.0252 | -43.8074
assumed numbers

## Data Wi-Fi
Product Number | HAC Compliant? | Air Interface | Bands | Form Factor | Standard Version | Max Power dBm | Max Avg. E Field (V/m) | Max Avg. E Field dB(V/m) | After Exclusion E field (V/m) | After Exclusion E field dB(V/m) | PMF Scaled E field (V/m) | PMF Scaled E field dB(V/m) | Rating E field | Max Avg. H Field (A/m) | Max Avg. H Field dB(A/m) | After Exclusion H field (A/m) | After Exclusion H field dB(A/m) | PMF Scaled H field (A/m) | PMF Scaled H field dB(A/m) | Rating H field | E H Ratings | Delta (dB) E-Field relative to M3 (pass) | Delta (dB) H-Field relative to M3 (pass) | OK E-Field Dominate | Amount H-Field is Dominate in dB | Corrected Delta (dB) E-Field relative to M3 (pass) | Corrected Delta (dB) H-Field relative to M3 (pass) | E/H ratio (Ohms) | peak dB(V/m) - peak dBm | peak dB(A/m) - peak dBm
127 | Pass | 802.11b/g |  | Candy Bar PDA | 2006 | 18 |  |  |  |  | 14.4 | 23.1672 | 4 |  |  |  |  | 0.037 | -28.636 | 4 | 4 | -17.8328 | -19.236 | OK |  |  |  | 389.1892 | 5.1672 | -46.636
128 | Pass | 802.11b./g |  |  | 2006 | 18 |  |  |  |  | 29.1 | 29.2779 | 4 |  |  |  |  | 0.074 | -22.6154 | 4 | 4 | -11.7221 | -13.2154 | OK |  |  |  | 393.2432 | 11.2779 | -40.6154
129 | Pass | 802.11b/g |  | Candy Bar PDA | 2006 | 18 |  |  |  |  | 25.6 | 28.1648 | 4 |  |  |  |  | 0.05 | -26.0206 | 4 | 4 | -12.8352 | -16.6206 | OK |  |  |  | 512 | 10.1648 | -44.0206
130 | Pass | 802.11a |  | Candy Bar PDA | 2006 | 17 |  |  |  |  | 13.1 | 22.3454 | 4 | [7 empty] | 4 | -18.6546 | [6 empty] | 5.3454
131 | Pass | 802.11b/g |  | Candy Bar PDA | 2006 | 18 |  |  |  |  | 29.4 | 29.3669 | 4 |  |  |  |  | 0.09 | -20.9151 |  | 4 | -11.6331 | -11.5151 | Not OK | 0.1179 |  |  | 326.6667 | 11.3669 | -38.9151
132 | Pass | 802.11a |  | Candy Bar PDA | 2006 | 17 |  |  |  |  | 13.1 | 22.3454 | 4 | [7 empty] | 4 | -18.6546 | [6 empty] | 5.3454

## Wi-Fi
HAC Near-Field RF Emissions Test Log | [5 empty] | E-Field |  |  |  |  | H-Field
Technology | Freq. (MHz) | Conducted Tx power (dBm) | Device Type, Model | Device Type | PMF | Peak E-Field (V/m) | Peak E-Field dB(V/m) | M rating | Verdict | Date | Peak H-Field (A/m) | Peak H-Field dB(A/m) | M rating | Verdict | Date | E/H ratio (Ohms) | peak dB(V/m) - peak dBm | peak dB(A/m) - peak dBm
802.11b | 2412 | 18 | Device #1 | Brick | 1.58 | 39.8 | 31.9977 | 4 | P | 10/14/09 | 0.174 | -15.189 | 4 | P | 10/15/09 | 228.7356 | 13.9977 | -33.189
802.11b | 2437 | 18 | Device #1 | Brick | 1.58 | 34.5 | 30.7564 | 4 | P | 10/14/09 | 0.155 | -16.1934 | 4 | P | 10/15/09 | 222.5806 | 12.7564 | -34.1934
802.11b | 2462 | 18 | Device #1 | Brick | 1.58 | 27.6 | 28.8182 | 4 | P | 10/14/09 | 0.124 | -18.1316 | 4 | P | 10/15/09 | 222.5806 | 10.8182 | -36.1316
802.11b | 2412 | 17.5 | Device #2 | Brick | 1.58 | 24.1 | 27.6403 | 4 | P | 10/14/09 | 0.07 | -23.098 | 4 | P | 10/15/09 | 344.2857 | 10.1403 | -40.598
802.11b | 2437 | 17.5 | Device #2 | Brick | 1.58 | 29.6 | 29.4258 | 4 | P | 10/14/09 | 0.095 | -20.4455 | 4 | P | 10/15/09 | 311.5789 | 11.9258 | -37.9455
802.11b | 2462 | 17.5 | Device #2 | Brick | 1.58 | 32.3 | 30.1841 | 4 | P | 10/14/09 | 0.091 | -20.8192 | 4 | P | 10/15/09 | 354.9451 | 12.6841 | -38.3192
802.11b | 2412 | 17.5 | Device #3 | Brick | 1.58 | 52.3 | 34.37 | 4 | P | 10/14/09 | 0.086 | -21.31 | 4 | P | 10/15/09 | 608.1395 | 16.87 | -38.81
802.11b | 2437 | 17.5 | Device #3 | Brick | 1.58 | 42.5 | 32.5678 | 4 | P | 10/14/09 | 0.098 | -20.1755 | 4 | P | 10/15/09 | 433.6735 | 15.0678 | -37.6755
802.11b | 2462 | 17.5 | Device #3 | Brick | 1.58 | 49.1 | 33.8216 | 4 | P | 10/14/09 | 0.08 | -21.9382 | 4 | P | 10/15/09 | 613.75 | 16.3216 | -39.4382
802.11b | 2412 | 17.5 | Device #4 | Brick | 1.58 | 36.7 | 31.2933 | 4 | P | 10/14/09 | 0.132 | -17.5885 | 4 | P | 10/15/09 | 278.0303 | 13.7933 | -35.0885
802.11b | 2437 | 17.5 | Device #4 | Brick | 1.58 | 33.8 | 30.5783 | 4 | P | 10/14/09 | 0.121 | -18.3443 | 4 | P | 10/15/09 | 279.3388 | 13.0783 | -35.8443
802.11b | 2462 | 17.5 | Device #4 | Brick | 1.58 | 36.9 | 31.3405 | 4 | P | 10/14/09 | 0.122 | -18.2728 | 4 | P | 10/15/09 | 302.459 | 13.8405 | -35.7728
802.11b | 2412 | 17.5 | Device #5 | Brick | 1.58 | 28.1 | 28.9741 | 4 | P | 10/14/09 | 0.088 | -21.1103 | 4 | P | 10/15/09 | 319.3182 | 11.4741 | -38.6103
802.11b | 2437 | 17.5 | Device #5 | Brick | 1.58 | 32 | 30.103 | 4 | P | 10/14/09 | 0.104 | -19.6593 | 4 | P | 10/15/09 | 307.6923 | 12.603 | -37.1593
802.11b | 2462 | 17.5 | Device #5 | Brick | 1.58 | 34.2 | 30.6805 | 4 | P | 10/14/09 | 0.102 | -19.828 | 4 | P | 10/15/09 | 335.2941 | 13.1805 | -37.328
802.11b | 2412 | 17.5 | Device #6 | Brick | 1.58 |  |  | 4 | P | 10/14/09 |  |  | 4 | P | 10/16/09
802.11b | 2437 | 17.5 | Device #6 | Brick | 1.58 |  |  | 4 | P | 10/14/09 |  |  | 4 | P | 10/16/09
802.11b | 2462 | 17.5 | Device #6 | Brick | 1.58 |  |  | 4 | P | 10/14/09 |  |  | 4 | P | 10/16/09
802.11b | 2412 | 17.5 | Device #7 | Brick | 1.58 | 27.5 | 28.7867 | 4 | P | 10/14/09 | 0.112 | -19.0156 | 4 | P | 10/16/09 | 245.5357 | 11.2867 | -36.5156
802.11b | 2437 | 17.5 | Device #7 | Brick | 1.58 | 25.7 | 28.1987 | 4 | P | 10/14/09 | 0.111 | -19.0935 | 4 | P | 10/16/09 | 231.5315 | 10.6987 | -36.5935
802.11b | 2462 | 17.5 | Device #7 | Brick | 1.58 | 27.8 | 28.8809 | 4 | P | 10/14/09 | 0.119 | -18.4891 | 4 | P | 10/16/09 | 233.6134 | 11.3809 | -35.9891
802.11b | 2412 | 18 | Device #8 | Brick | 1.58 | 48.6 | 33.7327 | 4 | P | 10/14/09 | 0.133 | -17.523 | 4 | P | 10/16/09 | 365.4135 | 15.7327 | -35.523
802.11b | 2437 | 18 | Device #8 | Brick | 1.58 | 54.8 | 34.7756 | 4 | P | 10/14/09 | 0.14 | -17.0774 | 4 | P | 10/16/09 | 391.4286 | 16.7756 | -35.0774
802.11b | 2462 | 18 | Device #8 | Brick | 1.58 | 55 | 34.8073 | 4 | P | 10/14/09 | 0.137 | -17.2656 | 4 | P | 10/16/09 | 401.4599 | 16.8073 | -35.2656
assumed peak power

## Bluetooth
HAC Near-Field RF Emissions Test Log | [5 empty] | E-Field |  |  |  |  | H-Field
Technology | Freq. (MHz) | Conducted Tx power (dBm) | Device Type, Model | Device Type | PMF | Peak E-Field (V/m) | Peak E-Field dB(V/m) | M rating | Verdict | Date | Peak H-Field (A/m) | Peak H-Field dB(A/m) | M rating | Verdict | Date | E/H (Ohms) | peak dB(V/m) - peak dBm | peak dB(A/m) - peak dBm
Bluetooth | 2402 | 8.5 | Device #1 | Brick | 1.53 | 3.8 | 11.5957 | 4 | P | 2009-05-10 | 0.012 | -38.4164 | 4 | P | 2009-06-10 | 316.6667 | 3.0957 | -46.9164
Bluetooth | 2441 | 8.5 | Device #1 | Brick | 1.53 | 3.71 | 11.3875 | 4 | P | 2009-05-10 | 0.011 | -39.1721 | 4 | P | 2009-06-10 | 337.2727 | 2.8875 | -47.6721
Bluetooth | 2480 | 8.5 | Device #1 | Brick | 1.53 | 3.45 | 10.7564 | 4 | P | 2009-05-10 | 0.01 | -40 | 4 | P | 2009-06-10 | 345 | 2.2564 | -48.5
Bluetooth | 2402 | 5.33 | Device #2 | Brick | 1.53 | 4.05 | 12.1491 | 4 | P | 2009-05-10 | 0.006 | -44.437 | 4 | P | 2009-06-10 | 675 | 6.8191 | -49.767
Bluetooth | 2441 | 5.33 | Device #2 | Brick | 1.53 | 2.91 | 9.2779 | 4 | P | 2009-05-10 | 0.007 | -43.098 | 4 | P | 2009-06-10 | 415.7143 | 3.9479 | -48.428
Bluetooth | 2480 | 5.33 | Device #2 | Brick | 1.53 | 3.89 | 11.799 | 4 | P | 2009-05-10 | 0.009 | -40.9151 | 4 | P | 2009-06-10 | 432.2222 | 6.469 | -46.2451
Bluetooth | 2402 | 7.67 | Device #3 | Brick | 1.53 | 4.73 | 13.4972 | 4 | P | 2009-05-10 | 0.017 | -35.391 | 4 | P | 2009-06-10 | 278.2353 | 5.8272 | -43.061
Bluetooth | 2441 | 7.67 | Device #3 | Brick | 1.53 | 3.86 | 11.7317 | 4 | P | 2009-05-10 | 0.014 | -37.0774 | 4 | P | 2009-06-10 | 275.7143 | 4.0617 | -44.7474
Bluetooth | 2480 | 7.67 | Device #3 | Brick | 1.53 | 3.66 | 11.2696 | 4 | P | 2009-05-10 | 0.011 | -39.1721 | 4 | P | 2009-06-10 | 332.7273 | 3.5996 | -46.8421
Bluetooth | 2402 | 8.67 | Device #4 | Brick | 1.53 | 6.81 | 16.6629 | 4 | P | 2009-05-10 | 0.021 | -33.5556 | 4 | P | 2009-06-10 | 324.2857 | 7.9929 | -42.2256
Bluetooth | 2441 | 8.67 | Device #4 | Brick | 1.53 | 6.63 | 16.4303 | 4 | P | 2009-05-10 | 0.017 | -35.391 | 4 | P | 2009-06-10 | 390 | 7.7603 | -44.061
Bluetooth | 2480 | 8.67 | Device #4 | Brick | 1.53 | 6.05 | 15.6351 | 4 | P | 2009-05-10 | 0.016 | -35.9176 | 4 | P | 2009-06-10 | 378.125 | 6.9651 | -44.5876
Bluetooth | 2402 | 7.8 | Device #5 | Brick | 1.53 | 5.33 | 14.5345 | 4 | P | 2009-05-10 | 0.01 | -40 | 4 | P | 2009-06-10 | 533 | 6.7345 | -47.8
Bluetooth | 2441 | 7.8 | Device #5 | Brick | 1.53 | 3.17 | 10.0212 | 4 | P | 2009-05-10 | 0.01 | -40 | 4 | P | 2009-06-10 | 317 | 2.2212 | -47.8
Bluetooth | 2480 | 7.8 | Device #5 | Brick | 1.53 | 3.87 | 11.7542 | 4 | P | 2009-05-10 | 0.009 | -40.9151 | 4 | P | 2009-06-10 | 430 | 3.9542 | -48.7151
Bluetooth | 2402 | 7.8 | Device #6 | Brick | 1.53 | 4.01 | 12.0629 | 4 | P | 2009-05-10 | 0.006 | -44.437 | 4 | P | 2009-06-10 | 668.3333 | 4.2629 | -52.237
Bluetooth | 2441 | 7.8 | Device #6 | Brick | 1.53 | 4.18 | 12.4235 | 4 | P | 2009-05-10 | 0.006 | -44.437 | 4 | P | 2009-06-10 | 696.6667 | 4.6235 | -52.237
Bluetooth | 2480 | 7.8 | Device #6 | Brick | 1.53 | 3.09 | 9.7992 | 4 | P | 2009-05-10 | 0.006 | -44.437 | 4 | P | 2009-06-10 | 515 | 1.9992 | -52.237
Bluetooth | 2402 | 8 | Device #7 | Brick | 1.53 | 5.02 | 14.0141 | 4 | P | 2009-05-10 | 0.018 | -34.8945 | 4 | P | 2009-06-10 | 278.8889 | 6.0141 | -42.8945
Bluetooth | 2441 | 8 | Device #7 | Brick | 1.53 | 4.7 | 13.442 | 4 | P | 2009-05-10 | 0.015 | -36.4782 | 4 | P | 2009-06-10 | 313.3333 | 5.442 | -44.4782
Bluetooth | 2480 | 8 | Device #7 | Brick | 1.53 | 4.63 | 13.3116 | 4 | P | 2009-05-10 | 0.014 | -37.0774 | 4 | P | 2009-06-10 | 330.7143 | 5.3116 | -45.0774
Bluetooth | 2402 | 5.5 | Device #8 | Brick | 1.53 | 3.29 | 10.3439 | 4 | P | 2009-05-10 | 0.01 | -40 | 4 | P | 2009-06-10 | 329 | 4.8439 | -45.5
Bluetooth | 2441 | 5.5 | Device #8 | Brick | 1.53 | 3.1 | 9.8272 | 4 | P | 2009-05-10 | 0.009 | -40.9151 | 4 | P | 2009-06-10 | 344.4444 | 4.3272 | -46.4151
Bluetooth | 2480 | 5.5 | Device #8 | Brick | 1.53 | 3.53 | 10.9555 | 4 | P | 2009-05-10 | 0.006 | -44.437 | 4 | P | 2009-06-10 | 588.3333 | 5.4555 | -49.937
Bluetooth | 2402 | 5.33 | Device #9 | Brick | 1.53 | 9.09 | 19.1713 | 4 | P | 2009-05-10 | 0.007 | -43.098 | 4 | P | 2009-06-10 | 1298.5714 | 13.8413 | -48.428
Bluetooth | 2441 | 5.33 | Device #9 | Brick | 1.53 | 9.11 | 19.1904 | 4 | P | 2009-05-10 | 0.009 | -40.9151 | 4 | P | 2009-06-10 | 1012.2222 | 13.8604 | -46.2451
Bluetooth | 2480 | 5.33 | Device #9 | Brick | 1.53 | 9.28 | 19.351 | 4 | P | 2009-05-10 | 0.01 | -40 | 4 | P | 2009-06-10 | 928 | 14.021 | -45.33
Bluetooth | 2402 | 8 | Device #10 | Brick | 1.53 | 3.37 | 10.5526 | 4 | P | 2009-05-10 | 0.008 | -41.9382 | 4 | P | 2009-06-10 | 421.25 | 2.5526 | -49.9382
Bluetooth | 2441 | 8 | Device #10 | Brick | 1.53 | 3.27 | 10.291 | 4 | P | 2009-05-10 | 0.006 | -44.437 | 4 | P | 2009-06-10 | 545 | 2.291 | -52.437
Bluetooth | 2480 | 8 | Device #10 | Brick | 1.53 | 4.62 | 13.2928 | 4 | P | 2009-05-10 | 0.01 | -40 | 4 | P | 2009-06-10 | 462 | 5.2928 | -48
assumed peak power

## E-H Stats
linear |  |  |  | "dB" (20*log10(linear)) |  |  |  | cumulative dist. plotting indices | [5 empty] | low band | high band | 802.11b | Bluetooth |  |  | H-field dominant ref: |  |  | 50.4 | H-field dominant ref: |  |  | 49.4 | H-field dominant ref: |  |  | 48.4 | H-field dominant ref: |  |  | 47.4 | H-field dominant ref: |  |  | 46.4 | H-field dominant ref: |  |  | 45.4
E/H ratio (Ohms) GSM, CDMA, WCDMA, iDEN low band | E/H ratio (Ohms) GSM, CDMA, WCDMA high band | E/H ratio (Ohms) 802.11b (Wi-Fi) | E/H ratio (Ohms) Bluetooth | E/H ratio (dBOhms) GSM, CDMA, WCDMA, iDEN low band | E/H ratio (dBOhms) GSM, CDMA, WCDMA high band | E/H ratio (dBOhms) 802.11b (Wi-Fi) | E/H ratio (dBOhms) Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | bin labels for plotting | bins | Frequency | Frequency | Frequency | Frequency |  |  | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth
203.8462 | 113.7391 | 222.5806 | 275.7143 | 46.1861 | 41.1182 | 46.9497 | 48.8092 | 0.0046296 | 0.0047619 | 0.04 | 0.033333 | 35 | 35.5 | 0 | 0 | 0 | 0 |  |  | 4.2139 | 9.2818 | 3.4503 | 1.5908 | 3.2139 | 8.2818 | 2.4503 | 0.59081 | 2.2139 | 7.2818 | 1.4503 |  | 1.2139 | 6.2818 | 0.45025 |  | 0.21395 | 5.2818 |  |  |  | 4.2818
206.4481 | 189.1144 | 222.5806 | 278.2353 | 46.2962 | 45.5345 | 46.9497 | 48.8882 | 0.0092593 | 0.0095238 | 0.08 | 0.066667 | 36 | 36.5 | 0 | 0 | 0 | 0 |  |  | 4.1038 | 4.8655 | 3.4503 | 1.5118 | 3.1038 | 3.8655 | 2.4503 | 0.51176 | 2.1038 | 2.8655 | 1.4503 |  | 1.1038 | 1.8655 | 0.45025 |  | 0.10378 | 0.86551
213.0392 | 212.5 | 228.7356 | 278.8889 | 46.5692 | 46.5472 | 47.1867 | 48.9086 | 0.013889 | 0.014286 | 0.12 | 0.1 | 37 | 37.5 | 0 | 0 | 0 | 0 |  |  | 3.8308 | 3.8528 | 3.2133 | 1.4914 | 2.8308 | 2.8528 | 2.2133 | 0.49138 | 1.8308 | 1.8528 | 1.2133 |  | 0.83081 | 0.85282 | 0.21332
232.4594 | 224.8276 | 231.5315 | 313.3333 | 47.3269 | 47.037 | 47.2922 | 49.9201 | 0.018519 | 0.019048 | 0.16 | 0.13333 | 38 | 38.5 | 0 | 0 | 0 | 0 | for plotting reference ratios |  | 3.0731 | 3.363 | 3.1078 | 0.47987 | 2.0731 | 2.363 | 2.1078 |  | 1.0731 | 1.363 | 1.1078 |  | 0.073057 | 0.36301 | 0.1078
234.5982 | 226.4559 | 233.6134 | 316.6667 | 47.4065 | 47.0997 | 47.37 | 50.012 | 0.023148 | 0.02381 | 0.2 | 0.16667 | 39 | 39.5 | 0 | 0 | 0 | 0 |  |  | 2.9935 | 3.3003 | 3.03 | 0.38795 | 1.9935 | 2.3003 | 2.03 |  | 0.99351 | 1.3003 | 1.03 |  |  | 0.30033 | 0.030043
235.5164 | 231.1248 | 245.5357 | 317 | 47.4404 | 47.2769 | 47.8023 | 50.0212 | 0.027778 | 0.028571 | 0.24 | 0.2 | 40 | 40.5 | 0 | 0 | 0 | 0 | 0 | C63.19 WD | 2.9596 | 3.1231 | 2.5977 | 0.37881 | 1.9596 | 2.1231 | 1.5977 |  | 0.95958 | 1.1231 | 0.59771 |  |  | 0.12307
242.8224 | 234.4411 | 278.0303 | 324.2857 | 47.7058 | 47.4007 | 48.8818 | 50.2186 | 0.032407 | 0.033333 | 0.28 | 0.23333 | 41 | 41.5 | 0 | 1 | 0 | 0 | 1 | 50.4 | 2.6942 | 2.9993 | 1.5182 | 0.18144 | 1.6942 | 1.9993 | 0.51816 |  | 0.69423 | 0.99933
245.1436 | 236.6667 | 279.3388 | 329 | 47.7884 | 47.4827 | 48.9226 | 50.3439 | 0.037037 | 0.038095 | 0.32 | 0.26667 | 42 | 42.5 | 0 | 0 | 0 | 0 |  | 50.4 | 2.6116 | 2.9173 | 1.4774 | 0.056082 | 1.6116 | 1.9173 | 0.47737 |  | 0.61159 | 0.91726
272.3958 | 238.9362 | 302.459 | 330.7143 | 48.704 | 47.5656 | 49.6133 | 50.3891 | 0.041667 | 0.042857 | 0.36 | 0.3 | 43 | 43.5 | 0 | 0 | 0 | 0 |  | HA dipole | 1.696 | 2.8344 | 0.78667 | 0.010941 | 0.69599 | 1.8344 |  |  |  | 0.83436
273.065 | 242.4635 | 307.6923 | 332.7273 | 48.7253 | 47.6929 | 49.7623 | 50.4418 | 0.046296 | 0.047619 | 0.4 | 0.33333 | 44 | 44.5 | 0 | 0 | 0 | 0 |  | 53 | 1.6747 | 2.7071 | 0.63767 |  | 0.67468 | 1.7071 |  |  |  | 0.70707
282.6797 | 243.4307 | 311.5789 | 337.2727 | 49.0259 | 47.7275 | 49.8714 | 50.5596 | 0.050926 | 0.052381 | 0.44 | 0.36667 | 45 | 45.5 | 0 | 0 | 0 | 0 |  | 53 | 1.3741 | 2.6725 | 0.52864 |  | 0.37411 | 1.6725 |  |  |  | 0.67249
292.0455 | 250.3817 | 319.3182 | 344.4444 | 49.309 | 47.9721 | 50.0845 | 50.7424 | 0.055556 | 0.057143 | 0.48 | 0.4 | 46 | 46.5 | 2 | 1 | 0 | 0 |  | far-field | 1.091 | 2.4279 | 0.31553 |  | 0.090991 | 1.4279 |  |  |  | 0.42795
296.7181 | 260.0688 | 326.6667 | 345 | 49.4469 | 48.3018 | 50.2821 | 50.7564 | 0.060185 | 0.061905 | 0.52 | 0.43333 | 47 | 47.5 | 4 | 6 | 5 | 0 |  | 51.52 | 0.95312 | 2.0982 | 0.1179 |  |  | 1.0982 |  |  |  | 0.098233
314.5205 | 261.4334 | 335.2941 | 378.125 | 49.953 | 48.3472 | 50.5085 | 51.5527 | 0.064815 | 0.066667 | 0.56 | 0.46667 | 48 | 48.5 | 2 | 6 | 1 | 0 |  | 51.52 | 0.44702 | 2.0528 |  |  |  | 1.0528 |  |  |  | 0.052777
318.3721 | 266.3342 | 344.2857 | 390 | 50.0587 | 48.5085 | 50.7384 | 51.8213 | 0.069444 | 0.071429 | 0.6 | 0.5 | 49 | 49.5 | 5 | 16 | 2 | 3 |  |  | 0.3413 | 1.8915 |  |  |  | 0.89146
319.0994 | 268.0782 | 354.9451 | 415.7143 | 50.0785 | 48.5652 | 51.0032 | 52.3759 | 0.074074 | 0.07619 | 0.64 | 0.53333 | 50 | 50.5 | 7 | 18 | 5 | 7 |  |  | 0.32148 | 1.8348 |  |  |  | 0.83477
320 | 273.399 | 365.4135 | 421.25 | 50.103 | 48.7359 | 51.2557 | 52.4908 | 0.078704 | 0.080952 | 0.68 | 0.56667 | 51 | 51.5 | 9 | 59 | 4 | 3 |  |  | 0.297 | 1.6641 |  |  |  | 0.66406
328.6076 | 277.8706 | 389.1892 | 430 | 50.3336 | 48.8769 | 51.8032 | 52.6694 | 0.083333 | 0.085714 | 0.72 | 0.6 | 52 | 52.5 | 7 | 52 | 4 | 4 |  |  | 0.066448 | 1.5231 |  |  |  | 0.52315
330.9259 | 279.4118 | 391.4286 | 432.2222 | 50.3946 | 48.9249 | 51.8531 | 52.7141 | 0.087963 | 0.090476 | 0.76 | 0.63333 | 53 | 53.5 | 9 | 26 | 1 | 3 |  |  | 0.0053841 | 1.4751 |  |  |  | 0.47511
333.5484 | 281.6794 | 393.2432 | 462 | 50.4632 | 48.9951 | 51.8932 | 53.2928 | 0.092593 | 0.095238 | 0.8 | 0.66667 | 54 | 54.5 | 16 | 17 | 1 | 1 |  |  |  | 1.4049 |  |  |  | 0.4049
338.5586 | 281.893 | 401.4599 | 515 | 50.5927 | 49.0017 | 52.0728 | 54.2361 | 0.097222 | 0.1 | 0.84 | 0.7 | 55 | 55.5 | 25 | 5 | 0 | 3 |  |  |  | 1.3983 |  |  |  | 0.39831
340 | 282.8125 | 433.6735 | 533 | 50.6296 | 49.03 | 52.7433 | 54.5345 | 0.10185 | 0.10476 | 0.88 | 0.73333 | 56 | 56.5 | 16 | 3 | 2 | 1 |  |  |  | 1.37 |  |  |  | 0.37003
346.0581 | 283.2765 | 512 | 545 | 50.783 | 49.0442 | 54.1854 | 54.7279 | 0.10648 | 0.10952 | 0.92 | 0.76667 | 57 | 57.5 | 40 | 0 | 0 | 2 |  |  |  | 1.3558 |  |  |  | 0.35579
349.1031 | 284.8739 | 608.1395 | 588.3333 | 50.8591 | 49.0931 | 55.6801 | 55.3925 | 0.11111 | 0.11429 | 0.96 | 0.8 | 58 | 58.5 | 45 | 0 | 0 | 0 |  |  |  | 1.3069 |  |  |  | 0.30695
350 | 284.8837 | 613.75 | 668.3333 | 50.8814 | 49.0934 | 55.7598 | 56.4999 | 0.11574 | 0.11905 | 1 | 0.83333 | 59 | 59.5 | 20 | 0 | 0 | 1 |  |  |  | 1.3066 |  |  |  | 0.30665
356.3433 | 285.3465 |  | 675 | 51.0374 | 49.1075 |  | 56.5861 | 0.12037 | 0.12381 |  | 0.86667 | 60 | 60.5 | 8 | 0 | 0 | 1 |  |  |  | 1.2925 |  |  |  | 0.29255
356.8807 | 285.3933 |  | 696.6667 | 51.0505 | 49.1089 |  | 56.8605 | 0.125 | 0.12857 |  | 0.9 | 61 | 61.5 | 1 | 0 | 0 | 0 |  |  |  | 1.2911 |  |  |  | 0.29113
358.7387 | 289.7561 |  | 928 | 51.0956 | 49.2407 |  | 59.351 | 0.12963 | 0.13333 |  | 0.93333 | 62 | 62.5 | 0 | 0 | 0 | 1 |  |  |  | 1.1593 |  |  |  | 0.15935
372.3776 | 289.7603 |  | 1012.2222 | 51.4197 | 49.2408 |  | 60.1055 | 0.13426 | 0.1381 |  | 0.96667 | 63 | 63.5 | 0 | 0 | 0 | 0 |  |  |  | 1.1592 |  |  |  | 0.15922
378.267 | 291.7293 |  | 1298.5714 | 51.556 | 49.2996 |  | 62.2693 | 0.13889 | 0.14286 |  | 1 | 64 | 64.5 | 0 | 0 | 0 | 0 |  |  |  | 1.1004 |  |  |  | 0.1004
382.2504 | 300.2415 |  |  | 51.647 | 49.5494 |  |  | 0.14352 | 0.14762 |  |  | 65 | 65.5 | 0 | 0 | 0 | 0 |  |  |  | 0.85058
391.1321 | 300.7353 |  |  | 51.8465 | 49.5637 |  |  | 0.14815 | 0.15238 |  |  | 66 | 66.5 | 0 | 0 | 0 | 0 |  |  |  | 0.83631
398.5714 | 302.1944 |  |  | 52.0101 | 49.6057 |  |  | 0.15278 | 0.15714 |  |  | 67 | 67.5 | 0 | 0 | 0 | 0 |  |  |  | 0.79427
406.0109 | 302.7875 |  |  | 52.1708 | 49.6228 |  |  | 0.15741 | 0.1619 |  |  | 68 | 68.5 | 1 | 0 | 0 | 0 | (questionable; not included in cum. dist.) |  |  | 0.77724
413.6364 | 304.0359 |  |  | 52.3324 | 49.6585 |  |  | 0.16204 | 0.16667 |  |  | 69 | 69.5 | 0 | 0 | 0 | 0 |  |  |  | 0.7415
416.3478 | 305.5738 |  |  | 52.3891 | 49.7023 |  |  | 0.16667 | 0.17143 |  |  | 70 | 70.5 | 0 | 0 | 0 | 0 |  |  |  | 0.69768
422.4 | 307.3529 |  |  | 52.5145 | 49.7527 |  |  | 0.1713 | 0.17619 |  |  | 71 | 71.5 | 0 | 1 | 0 | 0 | (questionable; not included in cum. dist.) |  |  | 0.64725
429.249 | 308 |  |  | 52.6542 | 49.771 |  |  | 0.17593 | 0.18095 |  |  | 72 | 72.5 | 1 | 0 | 0 | 0 | (questionable; not included in cum. dist.) |  |  | 0.62899
436.3636 | 310.6481 |  |  | 52.797 | 49.8454 |  |  | 0.18056 | 0.18571 |  |  | 73 | 73.5 | 0 | 0 | 0 | 0 |  |  |  | 0.55462
442.1348 | 312.2302 |  |  | 52.9111 | 49.8895 |  |  | 0.18519 | 0.19048 |  |  | 74 | 74.5 | 0 | 0 | 0 | 0 |  |  |  | 0.5105
452.6539 | 312.9187 |  |  | 53.1153 | 49.9086 |  |  | 0.18981 | 0.19524 |  |  | 75 | 75.5 | 0 | 0 | 0 | 0 |  |  |  | 0.49137
458.8745 | 314.9533 |  |  | 53.2339 | 49.9649 |  |  | 0.19444 | 0.2 |  |  |  | more | 0 | 0 | 0 | 0 |  |  |  | 0.43508
463.0435 | 317.2012 |  |  | 53.3124 | 50.0267 |  |  | 0.19907 | 0.20476 | [11 empty] | 0.3733
464.4689 | 323.1111 |  |  | 53.3391 | 50.187 |  |  | 0.2037 | 0.20952 | [11 empty] | 0.21296
467.0077 | 327.1642 |  |  | 53.3865 | 50.2953 |  |  | 0.20833 | 0.21429 | [11 empty] | 0.10469
474.9747 | 329.1498 |  |  | 53.5334 | 50.3479 |  |  | 0.21296 | 0.21905 | [11 empty] | 0.052128
477.1543 | 331.5789 |  |  | 53.5732 | 50.4117 |  |  | 0.21759 | 0.22381 | [10 empty] | 0.032174 | 0.075941 | 0.19385 | 0.040594 | 0.018811 | 0.039076 | 0.11076 | 0.010626 | 0.0097042 | 0.01952 | 0.054795 | 0 | 0.002983 | 0.0093205 | 0.010013 | 0 | 0.0002942 | 0.0058546 | 0 | 0 | 0 | 0.0040779 | 0 | 0
483.1373 | 332.4921 |  |  | 53.6814 | 50.4356 |  |  | 0.22222 | 0.22857 | [10 empty] | probabilities of changing M category if H-field dropped |  |  |  | probabilities of changing M category if H-field dropped |  |  |  | probabilities of changing M category if H-field dropped |  |  |  | probabilities of changing M category if H-field dropped |  |  |  | probabilities of changing M category if H-field dropped |  |  |  | probabilities of changing M category if H-field dropped
484.3177 | 335.8696 |  |  | 53.7026 | 50.5234 |  |  | 0.22685 | 0.23333
484.3243 | 337.1901 |  |  | 53.7027 | 50.5575 |  |  | 0.23148 | 0.2381 | [10 empty] | % H-field dominant (ref C63.19 rating diff.) |  |  |  | % H-field dominant (ref C63.19 rating diff.) |  |  |  | % H-field dominant (ref C63.19 rating diff.) |  |  |  | % H-field dominant (ref C63.19 rating diff.) |  |  |  | % H-field dominant (ref C63.19 rating diff.) |  |  |  | % H-field dominant (ref C63.19 rating diff.)
484.6154 | 337.5862 |  |  | 53.7079 | 50.5677 |  |  | 0.23611 | 0.24286 | [10 empty] | 0.087963 | 0.17143 | 0.52 | 0.3 | 0.055556 | 0.14286 | 0.32 | 0.1 | 0.037037 | 0.066667 | 0.24 | 0 | 0.018519 | 0.07619 | 0.2 | 0 | 0.0092593 | 0.0095238 | 0 | 0 | 0 | 0.0047619 | 0 | 0
485.3846 | 337.9221 |  |  | 53.7217 | 50.5763 |  |  | 0.24074 | 0.24762
488.1188 | 338.676 |  |  | 53.7705 | 50.5957 |  |  | 0.24537 | 0.25238
502.3981 | 340.0452 |  |  | 54.021 | 50.6307 |  |  | 0.25 | 0.25714
502.4793 | 341.3249 |  |  | 54.0224 | 50.6634 |  |  | 0.25463 | 0.2619
502.5641 | 341.5179 |  |  | 54.0238 | 50.6683 |  |  | 0.25926 | 0.26667
504 | 341.791 |  |  | 54.0486 | 50.6752 |  |  | 0.26389 | 0.27143
505.7778 | 343.7722 |  |  | 54.0792 | 50.7254 |  |  | 0.26852 | 0.27619
512.0316 | 344.375 |  |  | 54.1859 | 50.7406 |  |  | 0.27315 | 0.28095
513.1707 | 346.9484 |  |  | 54.2052 | 50.8053 |  |  | 0.27778 | 0.28571
521.2598 | 347.9412 |  |  | 54.3411 | 50.8301 |  |  | 0.28241 | 0.29048
533.4594 | 348.0392 |  |  | 54.542 | 50.8326 |  |  | 0.28704 | 0.29524
539.2523 | 348.8688 |  |  | 54.6358 | 50.8532 |  |  | 0.29167 | 0.3
541.3919 | 349.697 |  |  | 54.6702 | 50.8738 |  |  | 0.2963 | 0.30476
542 | 350 |  |  | 54.68 | 50.8814 |  |  | 0.30093 | 0.30952
542.4528 | 350.2591 |  |  | 54.6872 | 50.8878 |  |  | 0.30556 | 0.31429
542.9907 | 350.7042 |  |  | 54.6958 | 50.8988 |  |  | 0.31019 | 0.31905
548.2927 | 350.813 |  |  | 54.7802 | 50.9015 |  |  | 0.31481 | 0.32381
548.3333 | 352 |  |  | 54.7809 | 50.9309 |  |  | 0.31944 | 0.32857
548.5366 | 352.1472 |  |  | 54.7841 | 50.9345 |  |  | 0.32407 | 0.33333
556.6586 | 352.3438 |  |  | 54.9118 | 50.9393 |  |  | 0.3287 | 0.3381
557.6923 | 352.8846 |  |  | 54.9279 | 50.9527 |  |  | 0.33333 | 0.34286
558.1818 | 355.0532 |  |  | 54.9355 | 51.0059 |  |  | 0.33796 | 0.34762
558.5551 | 355.6485 |  |  | 54.9413 | 51.0204 |  |  | 0.34259 | 0.35238
559.1716 | 355.814 |  |  | 54.9509 | 51.0245 |  |  | 0.34722 | 0.35714
559.3186 | 355.8621 |  |  | 54.9532 | 51.0256 |  |  | 0.35185 | 0.3619
561.3861 | 356.3107 |  |  | 54.9852 | 51.0366 |  |  | 0.35648 | 0.36667
562.2905 | 356.9343 |  |  | 54.9992 | 51.0518 |  |  | 0.36111 | 0.37143
562.4528 | 357.4074 |  |  | 55.0017 | 51.0633 |  |  | 0.36574 | 0.37619
564.8387 | 357.4144 |  |  | 55.0385 | 51.0634 |  |  | 0.37037 | 0.38095
569.4915 | 359.2593 |  |  | 55.1097 | 51.1082 |  |  | 0.375 | 0.38571
572.7273 | 359.4783 |  |  | 55.159 | 51.1135 |  |  | 0.37963 | 0.39048
586.2805 | 359.4783 |  |  | 55.3621 | 51.1135 |  |  | 0.38426 | 0.39524
588.2784 | 359.5455 |  |  | 55.3917 | 51.1151 |  |  | 0.38889 | 0.4
589.196 | 361.25 |  |  | 55.4052 | 51.1562 |  |  | 0.39352 | 0.40476
594.1296 | 362 |  |  | 55.4776 | 51.1742 |  |  | 0.39815 | 0.40952
597.6923 | 364.0625 |  |  | 55.5296 | 51.2235 |  |  | 0.40278 | 0.41429
604.2254 | 364.6512 |  |  | 55.624 | 51.2376 |  |  | 0.40741 | 0.41905
604.646 | 365.7371 |  |  | 55.63 | 51.2634 |  |  | 0.41204 | 0.42381
607.8189 | 366.5584 |  |  | 55.6755 | 51.2829 |  |  | 0.41667 | 0.42857
618.2186 | 366.8874 |  |  | 55.8228 | 51.2907 |  |  | 0.4213 | 0.43333
627.027 | 367.0531 |  |  | 55.9457 | 51.2946 |  |  | 0.42593 | 0.4381
627.1111 | 367.0554 |  |  | 55.9469 | 51.2946 |  |  | 0.43056 | 0.44286
627.8788 | 367.8571 |  |  | 55.9575 | 51.3136 |  |  | 0.43519 | 0.44762
632.5513 | 367.907 |  |  | 56.0219 | 51.3148 |  |  | 0.43981 | 0.45238
651.495 | 367.9654 |  |  | 56.2782 | 51.3161 |  |  | 0.44444 | 0.45714
651.6949 | 370.7071 |  |  | 56.2809 | 51.3806 |  |  | 0.44907 | 0.4619
654.6296 | 370.8861 |  |  | 56.3199 | 51.3848 |  |  | 0.4537 | 0.46667
654.7619 | 371.3615 |  |  | 56.3217 | 51.3959 |  |  | 0.45833 | 0.47143
658.2609 | 371.7647 |  |  | 56.368 | 51.4054 |  |  | 0.46296 | 0.47619
666.3793 | 371.831 |  |  | 56.4744 | 51.4069 |  |  | 0.46759 | 0.48095
667.9894 | 372.0513 |  |  | 56.4954 | 51.4121 |  |  | 0.47222 | 0.48571
672.1763 | 372.5 |  |  | 56.5497 | 51.4225 |  |  | 0.47685 | 0.49048
673.913 | 372.7273 |  |  | 56.5721 | 51.4278 |  |  | 0.48148 | 0.49524
677.8761 | 372.9927 |  |  | 56.623 | 51.434 |  |  | 0.48611 | 0.5
679.1005 | 373.7589 |  |  | 56.6387 | 51.4518 |  |  | 0.49074 | 0.50476
679.8507 | 374.2382 |  |  | 56.6483 | 51.463 |  |  | 0.49537 | 0.50952
683.4016 | 376.4463 |  |  | 56.6935 | 51.5141 |  |  | 0.5 | 0.51429
685.4701 | 376.699 |  |  | 56.7198 | 51.5199 |  |  | 0.50463 | 0.51905
685.9935 | 376.9565 |  |  | 56.7264 | 51.5258 |  |  | 0.50926 | 0.52381
686.5815 | 377.1812 |  |  | 56.7338 | 51.531 |  |  | 0.51389 | 0.52857
688.8889 | 377.3399 |  |  | 56.763 | 51.5347 |  |  | 0.51852 | 0.53333
688.8889 | 377.3399 |  |  | 56.763 | 51.5347 |  |  | 0.52315 | 0.5381
696.99 | 378.0105 |  |  | 56.8645 | 51.5501 |  |  | 0.52778 | 0.54286
697.546 | 378.4615 |  |  | 56.8715 | 51.5604 |  |  | 0.53241 | 0.54762
698.4436 | 379.2079 |  |  | 56.8826 | 51.5775 |  |  | 0.53704 | 0.55238
698.5258 | 379.3478 |  |  | 56.8836 | 51.5808 |  |  | 0.54167 | 0.55714
702.2556 | 380 |  |  | 56.9299 | 51.5957 |  |  | 0.5463 | 0.5619
703.7037 | 380 |  |  | 56.9478 | 51.5957 |  |  | 0.55093 | 0.56667
703.9146 | 384.5506 |  |  | 56.9504 | 51.6991 |  |  | 0.55556 | 0.57143
704.3379 | 384.9624 |  |  | 56.9556 | 51.7084 |  |  | 0.56019 | 0.57619
709.375 | 385.3801 |  |  | 57.0175 | 51.7178 |  |  | 0.56481 | 0.58095
712.5 | 386.2595 |  |  | 57.0557 | 51.7376 |  |  | 0.56944 | 0.58571
714.7727 | 386.2595 |  |  | 57.0834 | 51.7376 |  |  | 0.57407 | 0.59048
714.9351 | 386.5517 |  |  | 57.0853 | 51.7442 |  |  | 0.5787 | 0.59524
716.0714 | 388.5246 |  |  | 57.0991 | 51.7884 |  |  | 0.58333 | 0.6
721.0526 | 389.7196 |  |  | 57.1593 | 51.815 |  |  | 0.58796 | 0.60476
722.0339 | 390.0621 |  |  | 57.1712 | 51.8227 |  |  | 0.59259 | 0.60952
725.3521 | 390.3553 |  |  | 57.211 | 51.8292 |  |  | 0.59722 | 0.61429
728.9655 | 391.0377 |  |  | 57.2541 | 51.8444 |  |  | 0.60185 | 0.61905
735.6495 | 392.7419 |  |  | 57.3334 | 51.8821 |  |  | 0.60648 | 0.62381
736.6242 | 395.5 |  |  | 57.3449 | 51.9429 |  |  | 0.61111 | 0.62857
738.835 | 396.2963 |  |  | 57.3709 | 51.9604 |  |  | 0.61574 | 0.63333
739.5833 | 396.5517 |  |  | 57.3797 | 51.966 |  |  | 0.62037 | 0.6381
739.6135 | 397.3094 |  |  | 57.3801 | 51.9826 |  |  | 0.625 | 0.64286
740.4624 | 397.6331 |  |  | 57.3901 | 51.9897 |  |  | 0.62963 | 0.64762
741.3717 | 397.6471 |  |  | 57.4007 | 51.99 |  |  | 0.63426 | 0.65238
742.4855 | 397.8355 |  |  | 57.4138 | 51.9941 |  |  | 0.63889 | 0.65714
747.9021 | 398.3333 |  |  | 57.4769 | 52.0049 |  |  | 0.64352 | 0.6619
748.4185 | 398.3784 |  |  | 57.4829 | 52.0059 |  |  | 0.64815 | 0.66667
749.6575 | 398.8304 |  |  | 57.4973 | 52.0158 |  |  | 0.65278 | 0.67143
749.7175 | 398.9071 |  |  | 57.498 | 52.0174 |  |  | 0.65741 | 0.67619
750.5618 | 401.875 |  |  | 57.5077 | 52.0818 |  |  | 0.66204 | 0.68095
750.5618 | 402.2346 |  |  | 57.5077 | 52.0896 |  |  | 0.66667 | 0.68571
750.6369 | 402.2857 |  |  | 57.5086 | 52.0907 |  |  | 0.6713 | 0.69048
753.0201 | 403.9755 |  |  | 57.5361 | 52.1271 |  |  | 0.67593 | 0.69524
753.0387 | 404.0816 |  |  | 57.5363 | 52.1294 |  |  | 0.68056 | 0.7
753.0675 | 404.2208 |  |  | 57.5367 | 52.1324 |  |  | 0.68519 | 0.70476
753.2895 | 405.3957 |  |  | 57.5392 | 52.1576 |  |  | 0.68981 | 0.70952
753.2895 | 405.8333 |  |  | 57.5392 | 52.167 |  |  | 0.69444 | 0.71429
755.1122 | 407.1006 |  |  | 57.5602 | 52.194 |  |  | 0.69907 | 0.71905
757.3529 | 407.9861 |  |  | 57.586 | 52.2129 |  |  | 0.7037 | 0.72381
757.4163 | 409.3264 |  |  | 57.5867 | 52.2414 |  |  | 0.70833 | 0.72857
766.3004 | 409.4017 |  |  | 57.688 | 52.243 |  |  | 0.71296 | 0.73333
769.3878 | 412.8028 |  |  | 57.7229 | 52.3149 |  |  | 0.71759 | 0.7381
773.3871 | 413.6364 |  |  | 57.7679 | 52.3324 |  |  | 0.72222 | 0.74286
773.8636 | 417.5325 |  |  | 57.7733 | 52.4138 |  |  | 0.72685 | 0.74762
775.3731 | 418.4615 |  |  | 57.7902 | 52.4331 |  |  | 0.73148 | 0.75238
776.1092 | 420.0557 |  |  | 57.7985 | 52.4661 |  |  | 0.73611 | 0.75714
777.8846 | 423.1884 |  |  | 57.8183 | 52.5307 |  |  | 0.74074 | 0.7619
780.1587 | 425.1656 |  |  | 57.8437 | 52.5712 |  |  | 0.74537 | 0.76667
781.1388 | 425.1908 |  |  | 57.8546 | 52.5717 |  |  | 0.75 | 0.77143
782.7485 | 426.4069 |  |  | 57.8724 | 52.5965 |  |  | 0.75463 | 0.77619
782.9787 | 427.8409 |  |  | 57.875 | 52.6256 |  |  | 0.75926 | 0.78095
792.5481 | 428.0899 |  |  | 57.9805 | 52.6307 |  |  | 0.76389 | 0.78571
796.6903 | 430.8108 |  |  | 58.0258 | 52.6857 |  |  | 0.76852 | 0.79048
802.5478 | 431.8182 |  |  | 58.0894 | 52.706 |  |  | 0.77315 | 0.79524
803.0488 | 432.4675 |  |  | 58.0948 | 52.7191 |  |  | 0.77778 | 0.8
804.2254 | 433.3333 |  |  | 58.1076 | 52.7364 |  |  | 0.78241 | 0.80476
806.6946 | 437.9487 |  |  | 58.1342 | 52.8285 |  |  | 0.78704 | 0.80952
806.9841 | 438.7755 |  |  | 58.1373 | 52.8448 |  |  | 0.79167 | 0.81429
808.6835 | 442.2764 |  |  | 58.1556 | 52.9139 |  |  | 0.7963 | 0.81905
810.9244 | 446.2185 |  |  | 58.1796 | 52.991 |  |  | 0.80093 | 0.82381
815.4971 | 449.6063 |  |  | 58.2284 | 53.0566 |  |  | 0.80556 | 0.82857
816.2921 | 454.6667 |  |  | 58.2369 | 53.1539 |  |  | 0.81019 | 0.83333
817.301 | 455.3073 |  |  | 58.2476 | 53.1661 |  |  | 0.81481 | 0.8381
819.8347 | 459.1928 |  |  | 58.2745 | 53.2399 |  |  | 0.81944 | 0.84286
821.1454 | 460.1852 |  |  | 58.2884 | 53.2587 |  |  | 0.82407 | 0.84762
823.8434 | 461.4213 |  |  | 58.3169 | 53.282 |  |  | 0.8287 | 0.85238
826.0223 | 463.7037 |  |  | 58.3398 | 53.3248 |  |  | 0.83333 | 0.85714
826.4 | 463.7037 |  |  | 58.3438 | 53.3248 |  |  | 0.83796 | 0.8619
826.8145 | 465.625 |  |  | 58.3482 | 53.3607 |  |  | 0.84259 | 0.86667
828.1879 | 465.9722 |  |  | 58.3626 | 53.3672 |  |  | 0.84722 | 0.87143
828.4866 | 468.8525 |  |  | 58.3657 | 53.4207 |  |  | 0.85185 | 0.87619
831.9783 | 469.7802 |  |  | 58.4022 | 53.4379 |  |  | 0.85648 | 0.88095
832.4324 | 473.6842 |  |  | 58.407 | 53.5098 |  |  | 0.86111 | 0.88571
838.7755 | 484.6667 |  |  | 58.4729 | 53.7089 |  |  | 0.86574 | 0.89048
845.1777 | 484.7222 |  |  | 58.539 | 53.7099 |  |  | 0.87037 | 0.89524
845.6446 | 488.2812 |  |  | 58.5438 | 53.7734 |  |  | 0.875 | 0.9
853.3333 | 493.578 |  |  | 58.6224 | 53.8671 |  |  | 0.87963 | 0.90476
853.6232 | 495.2381 |  |  | 58.6253 | 53.8963 |  |  | 0.88426 | 0.90952
855.8282 | 495.2381 |  |  | 58.6477 | 53.8963 |  |  | 0.88889 | 0.91429
862.7706 | 504.7619 |  |  | 58.7179 | 54.0617 |  |  | 0.89352 | 0.91905
863.3218 | 507.6923 |  |  | 58.7235 | 54.112 |  |  | 0.89815 | 0.92381
863.9004 | 508.0645 |  |  | 58.7293 | 54.1184 |  |  | 0.90278 | 0.92857
869.5906 | 510.5263 |  |  | 58.7863 | 54.1604 |  |  | 0.90741 | 0.93333
880 | 514.0625 |  |  | 58.8897 | 54.2203 |  |  | 0.91204 | 0.9381
883.3935 | 514.8387 |  |  | 58.9231 | 54.2334 |  |  | 0.91667 | 0.94286
885.0962 | 516.5138 |  |  | 58.9398 | 54.2616 |  |  | 0.9213 | 0.94762
885.4415 | 519.3798 |  |  | 58.9432 | 54.3097 |  |  | 0.92593 | 0.95238
886.8687 | 521.875 |  |  | 58.9572 | 54.3513 |  |  | 0.93056 | 0.95714
890.3846 | 530.0518 |  |  | 58.9916 | 54.4864 |  |  | 0.93519 | 0.9619
893.6364 | 539.1753 |  |  | 59.0232 | 54.6346 |  |  | 0.93981 | 0.96667
896.4286 | 549.3421 |  |  | 59.0503 | 54.7969 |  |  | 0.94444 | 0.97143
904 | 552.8455 |  |  | 59.1234 | 54.8521 |  |  | 0.94907 | 0.97619
930.8772 | 558.0247 |  |  | 59.3778 | 54.9331 |  |  | 0.9537 | 0.98095
938.1503 | 593.75 |  |  | 59.4454 | 55.4721 |  |  | 0.95833 | 0.98571
957.2193 | 609 |  |  | 59.6202 | 55.6923 |  |  | 0.96296 | 0.99048
972.5888 | 618.2927 |  |  | 59.7586 | 55.8239 |  |  | 0.96759 | 0.99524
988.8525 | 659.7156 |  |  | 59.9026 | 56.3871 |  |  | 0.97222 | 1
994.902 | 3539.4737 |  |  | 59.9556 | 70.9788 |  |  | 0.97685
1002.834 |  |  |  | 60.0246 |  |  |  | 0.98148
1004.0956 |  |  |  | 60.0355 |  |  |  | 0.98611
1059.1176 |  |  |  | 60.4989 |  |  |  | 0.99074
1059.2275 |  |  |  | 60.4998 |  |  |  | 0.99537
1084.2308 |  |  |  | 60.7024 |  |  |  | 1
2586.6856 |  |  |  | 68.2549
3968.1818 |  |  |  | 71.9718

## E-H graphs
probability of over-rating category by: | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | % H-field dominant (ref: C63.19 rating diff.) by: | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth
> 0 dB | 3.217% | 7.594% | 19.39% | 4.059% | > 0 dB | 8.796% | 17.14% | 52% | 30%
> 1 dB | 1.881% | 3.908% | 11.08% | 1.063% | > 1 dB | 5.556% | 14.29% | 32% | 10%
> 2 dB | 0.9704% | 1.952% | 5.479% | 0% | > 2 dB | 3.704% | 6.667% | 24% | 0%
> 3 dB | 0.2983% | 0.9321% | 1.001% | 0% | > 3 dB | 1.852% | 7.619% | 20% | 0%
> 4 dB | 0.02942% | 0.5855% | 0% | 0% | > 4 dB | 0.9259% | 0.9524% | 0% | 0%
> 5 dB | 0% | 0.4078% | 0% | 0% | > 5 dB | 0% | 0.4762% | 0% | 0%

## E-field gain
dB(V/m) - dBm |  |  |  | cumulative dist. plotting indices | [5 empty] | low band | high band | 802.11b | Bluetooth
GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | bin labels for plotting | bins | Frequency | Frequency | Frequency | Frequency
3.0474 | 2.7363 | 5.1672 | 1.9992 | 0.0045872 | 0.0047619 | 0.037037 | 0.033333 | 0 | 0.5 | 0 | 0 | 0 | 0
4.8084 | 3.8039 | 5.3454 | 2.2212 | 0.0091743 | 0.0095238 | 0.074074 | 0.066667 | 1 | 1.5 | 0 | 0 | 0 | 0
6.6001 | 4.4362 | 5.3454 | 2.2564 | 0.013761 | 0.014286 | 0.11111 | 0.1 | 2 | 2.5 | 0 | 0 | 0 | 4
7.547 | 4.5967 | 10.1403 | 2.291 | 0.018349 | 0.019048 | 0.14815 | 0.13333 | 3 | 3.5 | 1 | 1 | 0 | 3
7.7166 | 4.8859 | 10.1648 | 2.5526 | 0.022936 | 0.02381 | 0.18519 | 0.16667 | 4 | 4.5 | 0 | 2 | 0 | 6
8.2441 | 4.9171 | 10.6987 | 2.8875 | 0.027523 | 0.028571 | 0.22222 | 0.2 | 5 | 5.5 | 1 | 6 | 3 | 6
8.3414 | 5.0102 | 10.8182 | 3.0957 | 0.03211 | 0.033333 | 0.25926 | 0.23333 | 6 | 6.5 | 0 | 7 | 0 | 3
8.7411 | 5.2235 | 11.2779 | 3.5996 | 0.036697 | 0.038095 | 0.2963 | 0.26667 | 7 | 7.5 | 1 | 27 | 0 | 3
8.9936 | 5.3669 | 11.2867 | 3.9479 | 0.041284 | 0.042857 | 0.33333 | 0.3 | 8 | 8.5 | 4 | 31 | 0 | 2
9.2655 | 6.285 | 11.3669 | 3.9542 | 0.045872 | 0.047619 | 0.37037 | 0.33333 | 9 | 9.5 | 5 | 29 | 0 | 0
9.3491 | 6.2933 | 11.3809 | 4.0617 | 0.050459 | 0.052381 | 0.40741 | 0.36667 | 10 | 10.5 | 8 | 29 | 2 | 0
9.351 | 6.3381 | 11.4741 | 4.2629 | 0.055046 | 0.057143 | 0.44444 | 0.4 | 11 | 11.5 | 10 | 27 | 7 | 0
9.7281 | 6.3381 | 11.9258 | 4.3272 | 0.059633 | 0.061905 | 0.48148 | 0.43333 | 12 | 12.5 | 22 | 20 | 1 | 0
9.7403 | 6.3909 | 12.603 | 4.6235 | 0.06422 | 0.066667 | 0.51852 | 0.46667 | 13 | 13.5 | 33 | 20 | 5 | 0
9.8233 | 6.4489 | 12.6841 | 4.8439 | 0.068807 | 0.071429 | 0.55556 | 0.5 | 14 | 14.5 | 31 | 10 | 3 | 3
9.87 | 6.4937 | 12.7564 | 5.2928 | 0.073394 | 0.07619 | 0.59259 | 0.53333 | 15 | 15.5 | 38 | 1 | 1 | 0
9.9094 | 6.5215 | 13.0783 | 5.3116 | 0.077982 | 0.080952 | 0.62963 | 0.56667 | 16 | 16.5 | 21 | 0 | 2 | 0
9.9905 | 6.5345 | 13.1805 | 5.442 | 0.082569 | 0.085714 | 0.66667 | 0.6 | 17 | 17.5 | 11 | 0 | 3 | 0
10.137 | 6.5989 | 13.7933 | 5.4555 | 0.087156 | 0.090476 | 0.7037 | 0.63333 | 18 | 18.5 | 5 | 0 | 0 | 0
10.3876 | 6.6246 | 13.8405 | 5.8272 | 0.091743 | 0.095238 | 0.74074 | 0.66667 | 19 | 19.5 | 7 | 0 | 0 | 0
10.6025 | 6.6757 | 13.9977 | 6.0141 | 0.09633 | 0.1 | 0.77778 | 0.7 | 20 | 20.5 | 3 | 0 | 0 | 0
10.6768 | 6.7138 | 15.0678 | 6.469 | 0.10092 | 0.10476 | 0.81481 | 0.73333 | 21 | 21.5 | 3 | 0 | 0 | 0
10.6896 | 6.7391 | 15.7327 | 6.7345 | 0.1055 | 0.10952 | 0.85185 | 0.76667 | 22 | 22.5 | 5 | 0 | 0 | 0
10.7335 | 6.7665 | 16.3216 | 6.8191 | 0.11009 | 0.11429 | 0.88889 | 0.8 | 23 | 23.5 | 5 | 0 | 0 | 0
10.7835 | 6.777 | 16.7756 | 6.9651 | 0.11468 | 0.11905 | 0.92593 | 0.83333 | 24 | 24.5 | 1 | 0 | 0 | 0
10.7842 | 6.9515 | 16.8073 | 7.7603 | 0.11927 | 0.12381 | 0.96296 | 0.86667 | 25 | 25.5 | 1 | 0 | 0 | 0
10.8458 | 7.0252 | 16.87 | 7.9929 | 0.12385 | 0.12857 | 1 | 0.9 | 26 | 26.5 | 2 | 0 | 0 | 0
10.8995 | 7.0474 |  | 13.8413 | 0.12844 | 0.13333 |  | 0.93333 | 27 | 27.5 | 0 | 0 | 0 | 0
11.1149 | 7.0518 |  | 13.8604 | 0.13303 | 0.1381 |  | 0.96667 | 28 | 28.5 | 0 | 0 | 0 | 0
11.1781 | 7.104 |  | 14.021 | 0.13761 | 0.14286 |  | 1 | 29 | 29.5 | 0 | 0 | 0 | 0
11.5006 | 7.1466 |  |  | 0.1422 | 0.14762 |  |  | 30 | 30.5 | 0 | 0 | 0 | 0
11.6141 | 7.1502 |  |  | 0.14679 | 0.15238 |  |  | 31 | 31.5 | 0 | 0 | 0 | 0
11.7358 | 7.2185 |  |  | 0.15138 | 0.15714 |  |  | 32 | 32.5 | 0 | 0 | 0 | 0
11.8147 | 7.2307 |  |  | 0.15596 | 0.1619 |  |  | 33 | 33.5 | 0 | 0 | 0 | 0
11.8409 | 7.2979 |  |  | 0.16055 | 0.16667 |  |  | 34 | 34.5 | 0 | 0 | 0 | 0
11.8657 | 7.3139 |  |  | 0.16514 | 0.17143 |  |  | 35 | 35.5 | 0 | 0 | 0 | 0
11.9172 | 7.3565 |  |  | 0.16972 | 0.17619 |  |  | 36 | 36.5 | 0 | 0 | 0 | 0
11.93 | 7.3729 |  |  | 0.17431 | 0.18095 |  |  | 37 | 37.5 | 0 | 0 | 0 | 0
12.0252 | 7.4027 |  |  | 0.1789 | 0.18571 |  |  | 38 | 38.5 | 0 | 1 | 0 | 0
12.0425 | 7.4081 |  |  | 0.18349 | 0.19048 |  |  | 39 | 39.5 | 0 | 0 | 0 | 0
12.0668 | 7.4346 |  |  | 0.18807 | 0.19524 |  |  | 40 | 40.5 | 0 | 0 | 0 | 0
12.0813 | 7.4346 |  |  | 0.19266 | 0.2 |  |  |  | more | 0 | 0 | 0 | 0
12.23 | 7.4431 |  |  | 0.19725 | 0.20476
12.2668 | 7.5359 |  |  | 0.20183 | 0.20952
12.3011 | 7.5704 |  |  | 0.20642 | 0.21429
12.3716 | 7.614 |  |  | 0.21101 | 0.21905
12.3996 | 7.6391 |  |  | 0.2156 | 0.22381
12.4089 | 7.6694 |  |  | 0.22018 | 0.22857
12.4368 | 7.6742 |  |  | 0.22477 | 0.23333
12.4785 | 7.6742 |  |  | 0.22936 | 0.2381
12.4868 | 7.6846 |  |  | 0.23394 | 0.24286
12.4873 | 7.7411 |  |  | 0.23853 | 0.24762
12.515 | 7.7766 |  |  | 0.24312 | 0.25238
12.5246 | 7.7933 |  |  | 0.24771 | 0.25714
12.6252 | 7.7972 |  |  | 0.25229 | 0.2619
12.6277 | 7.8711 |  |  | 0.25688 | 0.26667
12.6629 | 7.8863 |  |  | 0.26147 | 0.27143
12.6796 | 7.9635 |  |  | 0.26606 | 0.27619
12.7021 | 7.9635 |  |  | 0.27064 | 0.28095
12.8096 | 7.9635 |  |  | 0.27523 | 0.28571
12.8496 | 8.0195 |  |  | 0.27982 | 0.29048
12.8981 | 8.0195 |  |  | 0.2844 | 0.29524
12.9198 | 8.0292 |  |  | 0.28899 | 0.3
12.9245 | 8.0401 |  |  | 0.29358 | 0.30476
12.9901 | 8.1267 |  |  | 0.29817 | 0.30952
13.0032 | 8.2018 |  |  | 0.30275 | 0.31429
13.0423 | 8.2232 |  |  | 0.30734 | 0.31905
13.0769 | 8.2551 |  |  | 0.31193 | 0.32381
13.1479 | 8.2766 |  |  | 0.31651 | 0.32857
13.159 | 8.348 |  |  | 0.3211 | 0.33333
13.2562 | 8.348 |  |  | 0.32569 | 0.3381
13.2814 | 8.4025 |  |  | 0.33028 | 0.34286
13.311 | 8.4025 |  |  | 0.33486 | 0.34762
13.311 | 8.4337 |  |  | 0.33945 | 0.35238
13.332 | 8.5339 |  |  | 0.34404 | 0.35714
13.3696 | 8.5884 |  |  | 0.34862 | 0.3619
13.378 | 8.5884 |  |  | 0.35321 | 0.36667
13.4112 | 8.6292 |  |  | 0.3578 | 0.37143
13.4195 | 8.7148 |  |  | 0.36239 | 0.37619
13.4237 | 8.7202 |  |  | 0.36697 | 0.38095
13.4527 | 8.7506 |  |  | 0.37156 | 0.38571
13.4609 | 8.7604 |  |  | 0.37615 | 0.39048
13.4692 | 8.7604 |  |  | 0.38073 | 0.39524
13.4815 | 8.7704 |  |  | 0.38532 | 0.4
13.4939 | 8.7804 |  |  | 0.38991 | 0.40476
13.5021 | 8.7862 |  |  | 0.3945 | 0.40952
13.5267 | 8.8041 |  |  | 0.39908 | 0.41429
13.5267 | 8.8798 |  |  | 0.40367 | 0.41905
13.5676 | 8.9271 |  |  | 0.40826 | 0.42381
13.5717 | 8.9487 |  |  | 0.41284 | 0.42857
13.5757 | 9.083 |  |  | 0.41743 | 0.43333
13.6163 | 9.083 |  |  | 0.42202 | 0.4381
13.677 | 9.1234 |  |  | 0.42661 | 0.44286
13.7412 | 9.1684 |  |  | 0.43119 | 0.44762
13.8049 | 9.1854 |  |  | 0.43578 | 0.45238
13.8168 | 9.1904 |  |  | 0.44037 | 0.45714
13.8287 | 9.2303 |  |  | 0.44495 | 0.4619
13.8563 | 9.2338 |  |  | 0.44954 | 0.46667
13.8863 | 9.2569 |  |  | 0.45413 | 0.47143
14.0359 | 9.304 |  |  | 0.45872 | 0.47619
14.0398 | 9.3883 |  |  | 0.4633 | 0.48095
14.0835 | 9.4162 |  |  | 0.46789 | 0.48571
14.1358 | 9.4626 |  |  | 0.47248 | 0.49048
14.1434 | 9.5545 |  |  | 0.47706 | 0.49524
14.1447 | 9.5545 |  |  | 0.48165 | 0.5
14.2194 | 9.6067 |  |  | 0.48624 | 0.50476
14.227 | 9.6092 |  |  | 0.49083 | 0.50952
14.2308 | 9.67 |  |  | 0.49541 | 0.51429
14.261 | 9.6815 |  |  | 0.5 | 0.51905
14.2835 | 9.6905 |  |  | 0.50459 | 0.52381
14.3134 | 9.6995 |  |  | 0.50917 | 0.52857
14.4029 | 9.7533 |  |  | 0.51376 | 0.53333
14.4094 | 9.7663 |  |  | 0.51835 | 0.5381
14.4349 | 9.8157 |  |  | 0.52294 | 0.54286
14.4472 | 9.8687 |  |  | 0.52752 | 0.54762
14.4656 | 9.8702 |  |  | 0.53211 | 0.55238
14.5242 | 9.8989 |  |  | 0.5367 | 0.55714
14.5352 | 9.9947 |  |  | 0.54128 | 0.5619
14.6364 | 10.1208 |  |  | 0.54587 | 0.56667
14.6475 | 10.1327 |  |  | 0.55046 | 0.57143
14.6617 | 10.1975 |  |  | 0.55505 | 0.57619
14.6739 | 10.206 |  |  | 0.55963 | 0.58095
14.6904 | 10.2145 |  |  | 0.56422 | 0.58571
14.6905 | 10.2326 |  |  | 0.56881 | 0.59048
14.7014 | 10.2506 |  |  | 0.57339 | 0.59524
14.73 | 10.2736 |  |  | 0.57798 | 0.6
14.7478 | 10.3283 |  |  | 0.58257 | 0.60476
14.7852 | 10.3823 |  |  | 0.58716 | 0.60952
14.7935 | 10.4446 |  |  | 0.59174 | 0.61429
14.801 | 10.4479 |  |  | 0.59633 | 0.61905
14.8328 | 10.4568 |  |  | 0.60092 | 0.62381
14.889 | 10.4733 |  |  | 0.6055 | 0.62857
14.9302 | 10.5225 |  |  | 0.61009 | 0.63333
14.9414 | 10.5714 |  |  | 0.61468 | 0.6381
14.991 | 10.6362 |  |  | 0.61927 | 0.64286
15.0086 | 10.6845 |  |  | 0.62385 | 0.64762
15.0452 | 10.6923 |  |  | 0.62844 | 0.65238
15.0864 | 10.7326 |  |  | 0.63303 | 0.65714
15.1885 | 10.7326 |  |  | 0.63761 | 0.6619
15.1953 | 10.7962 |  |  | 0.6422 | 0.66667
15.2122 | 10.8436 |  |  | 0.64679 | 0.67143
15.2155 | 10.883 |  |  | 0.65138 | 0.67619
15.2727 | 10.9176 |  |  | 0.65596 | 0.68095
15.2804 | 10.9315 |  |  | 0.66055 | 0.68571
15.282 | 10.9315 |  |  | 0.66514 | 0.69048
15.2861 | 10.9921 |  |  | 0.66972 | 0.69524
15.3072 | 11.0076 |  |  | 0.67431 | 0.7
15.3323 | 11.0385 |  |  | 0.6789 | 0.70476
15.3469 | 11.0539 |  |  | 0.68349 | 0.70952
15.4123 | 11.1637 |  |  | 0.68807 | 0.71429
15.4123 | 11.1988 |  |  | 0.69266 | 0.71905
15.4229 | 11.2064 |  |  | 0.69725 | 0.72381
15.4568 | 11.3265 |  |  | 0.70183 | 0.72857
15.4747 | 11.3645 |  |  | 0.70642 | 0.73333
15.553 | 11.4008 |  |  | 0.71101 | 0.7381
15.5785 | 11.4155 |  |  | 0.7156 | 0.74286
15.6035 | 11.4597 |  |  | 0.72018 | 0.74762
15.6594 | 11.4955 |  |  | 0.72477 | 0.75238
15.7328 | 11.4963 |  |  | 0.72936 | 0.75714
15.8279 | 11.5328 |  |  | 0.73394 | 0.7619
15.8594 | 11.5345 |  |  | 0.73853 | 0.76667
15.909 | 11.5989 |  |  | 0.74312 | 0.77143
15.9121 | 11.6642 |  |  | 0.74771 | 0.77619
15.9152 | 11.8894 |  |  | 0.75229 | 0.78095
15.9183 | 11.9034 |  |  | 0.75688 | 0.78571
15.9494 | 12.0005 |  |  | 0.76147 | 0.79048
15.9533 | 12.0074 |  |  | 0.76606 | 0.79524
15.9999 | 12.1713 |  |  | 0.77064 | 0.8
16.0019 | 12.1848 |  |  | 0.77523 | 0.80476
16.0111 | 12.2281 |  |  | 0.77982 | 0.80952
16.047 | 12.2426 |  |  | 0.7844 | 0.81429
16.2089 | 12.2855 |  |  | 0.78899 | 0.81905
16.2729 | 12.3322 |  |  | 0.79358 | 0.82381
16.4668 | 12.3322 |  |  | 0.79817 | 0.82857
16.4872 | 12.3983 |  |  | 0.80275 | 0.83333
16.5771 | 12.4181 |  |  | 0.80734 | 0.8381
16.5886 | 12.4592 |  |  | 0.81193 | 0.84286
16.6231 | 12.477 |  |  | 0.81651 | 0.84762
16.7428 | 12.4836 |  |  | 0.8211 | 0.85238
16.7659 | 12.5031 |  |  | 0.82569 | 0.85714
16.8193 | 12.5096 |  |  | 0.83028 | 0.8619
16.8522 | 12.5486 |  |  | 0.83486 | 0.86667
17.0074 | 12.5744 |  |  | 0.83945 | 0.87143
17.3624 | 12.5938 |  |  | 0.84404 | 0.87619
17.4489 | 12.6131 |  |  | 0.84862 | 0.88095
17.4574 | 12.6772 |  |  | 0.85321 | 0.88571
17.5035 | 12.8353 |  |  | 0.8578 | 0.89048
17.5526 | 12.8728 |  |  | 0.86239 | 0.89524
17.6899 | 12.9745 |  |  | 0.86697 | 0.9
18.0874 | 13.0513 |  |  | 0.87156 | 0.90476
18.3285 | 13.058 |  |  | 0.87615 | 0.90952
18.5442 | 13.168 |  |  | 0.88073 | 0.91429
18.6678 | 13.1733 |  |  | 0.88532 | 0.91905
18.7274 | 13.1757 |  |  | 0.88991 | 0.92381
18.7734 | 13.2333 |  |  | 0.8945 | 0.92857
18.9841 | 13.3699 |  |  | 0.89908 | 0.93333
19.2064 | 13.3816 |  |  | 0.90367 | 0.9381
19.4442 | 13.3817 |  |  | 0.90826 | 0.94286
19.5668 | 13.4986 |  |  | 0.91284 | 0.94762
19.5825 | 13.5218 |  |  | 0.91743 | 0.95238
20.0571 | 13.5392 |  |  | 0.92202 | 0.95714
20.5074 | 13.5853 |  |  | 0.92661 | 0.9619
20.583 | 13.6707 |  |  | 0.93119 | 0.96667
20.9755 | 13.6938 |  |  | 0.93578 | 0.97143
21.6597 | 13.7786 |  |  | 0.94037 | 0.97619
21.6851 | 13.8178 |  |  | 0.94495 | 0.98095
22.3749 | 13.9841 |  |  | 0.94954 | 0.98571
22.4019 | 14.1713 |  |  | 0.95413 | 0.99048
22.4149 | 14.3286 |  |  | 0.95872 | 0.99524
22.6185 | 14.6599 |  |  | 0.9633 | 1
22.6585 | 37.9525 |  |  | 0.96789
23.2687 |  |  |  | 0.97248
23.2875 |  |  |  | 0.97706
23.4268 |  |  |  | 0.98165
24.1978 |  |  |  | 0.98624
25.0964 |  |  |  | 0.99083
25.7705 |  |  |  | 0.99541
26.2104 |  |  |  | 1

## H-field gain
dB(A/m) - dBm |  |  |  | cumulative dist. plotting indices | [5 empty] | low band | high band | 802.11b | Bluetooth
GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth | bin labels for plotting | bins | Frequency | Frequency | Frequency | Frequency
-60.1309 | -58.4043 | -46.636 | -52.437 | 0.0046083 | 0.0047847 | 0.04 | 0.033333 | -60 | -59.5 | 1 | 0 | 0 | 0
-52.0624 | -49.2515 | -44.0206 | -52.237 | 0.0092166 | 0.0095694 | 0.08 | 0.066667 | -59 | -58.5 | 0 | 0 | 0 | 0
-50.7211 | -48.2882 | -40.6154 | -52.237 | 0.013825 | 0.014354 | 0.12 | 0.1 | -58 | -57.5 | 0 | 1 | 0 | 0
-48.0897 | -47.7882 | -40.598 | -52.237 | 0.018433 | 0.019139 | 0.16 | 0.13333 | -57 | -56.5 | 0 | 0 | 0 | 0
-48.0391 | -47.7211 | -39.4382 | -49.9382 | 0.023041 | 0.023923 | 0.2 | 0.16667 | -56 | -55.5 | 0 | 0 | 0 | 0
-47.223 | -47.7019 | -38.9151 | -49.937 | 0.02765 | 0.028708 | 0.24 | 0.2 | -55 | -54.5 | 0 | 0 | 0 | 0
-46.8971 | -47.6546 | -38.81 | -49.767 | 0.032258 | 0.033493 | 0.28 | 0.23333 | -54 | -53.5 | 0 | 0 | 0 | 0
-46.6387 | -47.0156 | -38.6103 | -48.7151 | 0.036866 | 0.038278 | 0.32 | 0.26667 | -53 | -52.5 | 0 | 0 | 0 | 0
-46.6107 | -46.6537 | -38.3192 | -48.5 | 0.041475 | 0.043062 | 0.36 | 0.3 | -52 | -51.5 | 1 | 0 | 0 | 4
-46.4806 | -46.3631 | -37.9455 | -48.428 | 0.046083 | 0.047847 | 0.4 | 0.33333 | -51 | -50.5 | 1 | 0 | 0 | 0
-46.3917 | -46.2702 | -37.6755 | -48.428 | 0.050691 | 0.052632 | 0.44 | 0.36667 | -50 | -49.5 | 0 | 0 | 0 | 3
-45.9563 | -46.1934 | -37.328 | -48 | 0.0553 | 0.057416 | 0.48 | 0.4 | -49 | -48.5 | 0 | 1 | 0 | 2
-45.8795 | -45.9176 | -37.1593 | -47.8 | 0.059908 | 0.062201 | 0.52 | 0.43333 | -48 | -47.5 | 2 | 5 | 0 | 6
-45.7654 | -45.8097 | -36.5935 | -47.8 | 0.064516 | 0.066986 | 0.56 | 0.46667 | -47 | -46.5 | 4 | 2 | 1 | 2
-45.7278 | -45.7031 | -36.5156 | -47.6721 | 0.069124 | 0.07177 | 0.6 | 0.5 | -46 | -45.5 | 8 | 8 | 0 | 4
-45.6902 | -45.6503 | -36.1316 | -46.9164 | 0.073733 | 0.076555 | 0.64 | 0.53333 | -45 | -44.5 | 15 | 25 | 0 | 4
-45.5786 | -45.5144 | -35.9891 | -46.8421 | 0.078341 | 0.08134 | 0.68 | 0.56667 | -44 | -43.5 | 23 | 23 | 1 | 2
-45.4515 | -45.4938 | -35.8443 | -46.4151 | 0.082949 | 0.086124 | 0.72 | 0.6 | -43 | -42.5 | 31 | 31 | 0 | 2
-45.432 | -45.4938 | -35.7728 | -46.2451 | 0.087558 | 0.090909 | 0.76 | 0.63333 | -42 | -41.5 | 27 | 26 | 0 | 1
-45.3597 | -45.4455 | -35.523 | -46.2451 | 0.092166 | 0.095694 | 0.8 | 0.66667 | -41 | -40.5 | 26 | 23 | 2 | 0
-45.3546 | -45.391 | -35.2656 | -45.5 | 0.096774 | 0.10048 | 0.84 | 0.7 | -40 | -39.5 | 22 | 16 | 0 | 0
-45.2879 | -45.3401 | -35.0885 | -45.33 | 0.10138 | 0.10526 | 0.88 | 0.73333 | -39 | -38.5 | 18 | 17 | 4 | 0
-45.1529 | -45.2894 | -35.0774 | -45.0774 | 0.10599 | 0.11005 | 0.92 | 0.76667 | -38 | -37.5 | 4 | 11 | 3 | 0
-45.1461 | -45.251 | -34.1934 | -44.7474 | 0.1106 | 0.11483 | 0.96 | 0.8 | -37 | -36.5 | 5 | 9 | 4 | 0
-45.1103 | -45.189 | -33.189 | -44.5876 | 0.11521 | 0.11962 | 1 | 0.83333 | -36 | -35.5 | 7 | 3 | 5 | 0
-44.9376 | -45.1755 |  | -44.4782 | 0.11982 | 0.1244 |  | 0.86667 | -35 | -34.5 | 9 | 3 | 3 | 0
-44.8692 | -45.1423 |  | -44.061 | 0.12442 | 0.12919 |  | 0.9 | -34 | -33.5 | 5 | 5 | 1 | 0
-44.8013 | -45.1392 |  | -43.061 | 0.12903 | 0.13397 |  | 0.93333 | -33 | -32.5 | 6 | 0 | 1 | 0
-44.7433 | -45.0897 |  | -42.8945 | 0.13364 | 0.13876 |  | 0.96667 | -32 | -31.5 | 1 | 0 | 0 | 0
-44.7005 | -45.0897 |  | -42.2256 | 0.13825 | 0.14354 |  | 1 | -31 | -30.5 | 2 | 0 | 0 | 0
-44.696 | -45.0873 |  |  | 0.14286 | 0.14833 |  |  | -30 | -29.5 | 0 | 0 | 0 | 0
-44.6593 | -45 |  |  | 0.14747 | 0.15311 |  |  | -29 | -28.5 | 0 | 1 | 0 | 0
-44.4461 | -44.9916 |  |  | 0.15207 | 0.15789 |  |  | -28 | -27.5 | 0 | 0 | 0 | 0
-44.4373 | -44.9646 |  |  | 0.15668 | 0.16268 |  |  | -27 | -26.5 | 0 | 0 | 0 | 0
-44.2767 | -44.9429 |  |  | 0.16129 | 0.16746 |  |  | -26 | -25.5 | 0 | 0 | 0 | 0
-44.2767 | -44.9429 |  |  | 0.1659 | 0.17225 |  |  | -25 | -24.5 | 0 | 0 | 0 | 0
-44.0935 | -44.9423 |  |  | 0.17051 | 0.17703 |  |  | -24 | -23.5 | 0 | 0 | 0 | 0
-44.0568 | -44.9116 |  |  | 0.17512 | 0.18182 |  |  | -23 | -22.5 | 0 | 0 | 0 | 0
-44.0259 | -44.6566 |  |  | 0.17972 | 0.1866 |  |  | -22 | -21.5 | 0 | 0 | 0 | 0
-44.0259 | -44.6107 |  |  | 0.18433 | 0.19139 |  |  | -21 | -20.5 | 0 | 0 | 0 | 0
-43.9123 | -44.6097 |  |  | 0.18894 | 0.19617 |  |  | -20 | -19.5 | 0 | 0 | 0 | 0
-43.905 | -44.5122 |  |  | 0.19355 | 0.20096 |  |  | -19 | -18.5 | 0 | 0 | 0 | 0
-43.9031 | -44.4123 |  |  | 0.19816 | 0.20574 |  |  | -18 | -17.5 | 0 | 0 | 0 | 0
-43.8727 | -44.3315 |  |  | 0.20276 | 0.21053 |  |  | -17 | -16.5 | 0 | 0 | 0 | 0
-43.8424 | -44.2889 |  |  | 0.20737 | 0.21531 |  |  | -16 | -15.5 | 0 | 0 | 0 | 0
-43.782 | -44.2242 |  |  | 0.21198 | 0.2201 |  |  | -15 | -14.5 | 0 | 0 | 0 | 0
-43.782 | -44.1993 |  |  | 0.21659 | 0.22488 |  |  | -14 | -13.5 | 0 | 1 | 0 | 0
-43.7626 | -44.1721 |  |  | 0.2212 | 0.22967 |  |  | -13 | -12.5 | 0 | 0 | 0 | 0
-43.752 | -44.1001 |  |  | 0.22581 | 0.23445 |  |  | -12 | -11.5 | 0 | 0 | 0 | 0
-43.6923 | -44.1001 |  |  | 0.23041 | 0.23923 |  |  | -11 | -10.5 | 0 | 0 | 0 | 0
-43.6504 | -44.0156 |  |  | 0.23502 | 0.24402 |  |  | -10 | -9.5 | 0 | 0 | 0 | 0
-43.6036 | -43.9794 |  |  | 0.23963 | 0.2488 |  |  |  | more | 0 | 0 | 0 | 0
-43.5742 | -43.9794 |  |  | 0.24424 | 0.25359
-43.5742 | -43.9733 |  |  | 0.24885 | 0.25837
-43.5449 | -43.9361 |  |  | 0.25346 | 0.26316
-43.4891 | -43.8501 |  |  | 0.25806 | 0.26794
-43.4866 | -43.8109 |  |  | 0.26267 | 0.27273
-43.4336 | -43.8074 |  |  | 0.26728 | 0.27751
-43.4287 | -43.7227 |  |  | 0.27189 | 0.2823
-43.3711 | -43.7227 |  |  | 0.2765 | 0.28708
-43.3443 | -43.6806 |  |  | 0.28111 | 0.29187
-43.3443 | -43.5762 |  |  | 0.28571 | 0.29665
-43.3425 | -43.5762 |  |  | 0.29032 | 0.30144
-43.3315 | -43.5144 |  |  | 0.29493 | 0.30622
-43.3259 | -43.5 |  |  | 0.29954 | 0.311
-43.314 | -43.4324 |  |  | 0.30415 | 0.31579
-43.2572 | -43.4324 |  |  | 0.30876 | 0.32057
-43.229 | -43.3512 |  |  | 0.31336 | 0.32536
-43.2024 | -43.3315 |  |  | 0.31797 | 0.33014
-43.1728 | -43.3109 |  |  | 0.32258 | 0.33493
-43.1316 | -43.2309 |  |  | 0.32719 | 0.33971
-43.0618 | -43.1593 |  |  | 0.3318 | 0.3445
-43.0614 | -43.1515 |  |  | 0.33641 | 0.34928
-43.0614 | -43.1316 |  |  | 0.34101 | 0.35407
-42.9515 | -43.1122 |  |  | 0.34562 | 0.35885
-42.8443 | -43.0339 |  |  | 0.35023 | 0.36364
-42.7356 | -43.0339 |  |  | 0.35484 | 0.36842
-42.7356 | -43.0339 |  |  | 0.35945 | 0.37321
-42.7356 | -42.9728 |  |  | 0.36406 | 0.37799
-42.709 | -42.9563 |  |  | 0.36866 | 0.38278
-42.6825 | -42.9239 |  |  | 0.37327 | 0.38756
-42.6825 | -42.8558 |  |  | 0.37788 | 0.39234
-42.6363 | -42.8558 |  |  | 0.38249 | 0.39713
-42.6034 | -42.8558 |  |  | 0.3871 | 0.40191
-42.5891 | -42.8413 |  |  | 0.39171 | 0.4067
-42.5338 | -42.7654 |  |  | 0.39631 | 0.41148
-42.4732 | -42.7654 |  |  | 0.40092 | 0.41627
-42.4579 | -42.7654 |  |  | 0.40553 | 0.42105
-42.4474 | -42.7278 |  |  | 0.41014 | 0.42584
-42.4239 | -42.7278 |  |  | 0.41475 | 0.43062
-42.396 | -42.6813 |  |  | 0.41935 | 0.43541
-42.3704 | -42.6546 |  |  | 0.42396 | 0.44019
-42.3558 | -42.6546 |  |  | 0.42857 | 0.44498
-42.3449 | -42.6363 |  |  | 0.43318 | 0.44976
-42.3195 | -42.523 |  |  | 0.43779 | 0.45455
-42.3195 | -42.505 |  |  | 0.4424 | 0.45933
-42.3195 | -42.4685 |  |  | 0.447 | 0.46411
-42.2211 | -42.432 |  |  | 0.45161 | 0.4689
-42.2185 | -42.3958 |  |  | 0.45622 | 0.47368
-42.1704 | -42.3958 |  |  | 0.46083 | 0.47847
-42.0939 | -42.3933 |  |  | 0.46544 | 0.48325
-42.0445 | -42.3933 |  |  | 0.47005 | 0.48804
-42.0199 | -42.3461 |  |  | 0.47465 | 0.49282
-42.0156 | -42.3237 |  |  | 0.47926 | 0.49761
-41.9579 | -42.2879 |  |  | 0.48387 | 0.50239
-41.9466 | -42.2656 |  |  | 0.48848 | 0.50718
-41.9223 | -42.2656 |  |  | 0.49309 | 0.51196
-41.8739 | -42.2278 |  |  | 0.4977 | 0.51675
-41.8019 | -42.2024 |  |  | 0.5023 | 0.52153
-41.6595 | -42.1813 |  |  | 0.50691 | 0.52632
-41.6537 | -42.1461 |  |  | 0.51152 | 0.5311
-41.636 | -42.1352 |  |  | 0.51613 | 0.53589
-41.5363 | -42.111 |  |  | 0.52074 | 0.54067
-41.471 | -42.0065 |  |  | 0.52535 | 0.54545
-41.4502 | -41.9891 |  |  | 0.52995 | 0.55024
-41.4502 | -41.9164 |  |  | 0.53456 | 0.55502
-41.3062 | -41.8328 |  |  | 0.53917 | 0.55981
-41.2211 | -41.7005 |  |  | 0.54378 | 0.56459
-41.1565 | -41.7005 |  |  | 0.54839 | 0.56938
-41.0901 | -41.634 |  |  | 0.553 | 0.57416
-41.082 | -41.6009 |  |  | 0.5576 | 0.57895
-41.0681 | -41.5363 |  |  | 0.56221 | 0.58373
-41.0461 | -41.4205 |  |  | 0.56682 | 0.58852
-41.0023 | -41.3727 |  |  | 0.57143 | 0.5933
-40.9805 | -41.3558 |  |  | 0.57604 | 0.59809
-40.9597 | -41.3237 |  |  | 0.58065 | 0.60287
-40.9371 | -41.3086 |  |  | 0.58525 | 0.60766
-40.8945 | -41.1191 |  |  | 0.58986 | 0.61244
-40.8724 | -41.1191 |  |  | 0.59447 | 0.61722
-40.7443 | -41.0259 |  |  | 0.59908 | 0.62201
-40.7443 | -41.023 |  |  | 0.60369 | 0.62679
-40.7443 | -40.9336 |  |  | 0.60829 | 0.63158
-40.7443 | -40.8424 |  |  | 0.6129 | 0.63636
-40.681 | -40.8424 |  |  | 0.61751 | 0.64115
-40.6503 | -40.8122 |  |  | 0.62212 | 0.64593
-40.6181 | -40.8122 |  |  | 0.62673 | 0.65072
-40.5973 | -40.782 |  |  | 0.63134 | 0.6555
-40.5679 | -40.752 |  |  | 0.63594 | 0.66029
-40.535 | -40.752 |  |  | 0.64055 | 0.66507
-40.4938 | -40.752 |  |  | 0.64516 | 0.66986
-40.4732 | -40.752 |  |  | 0.64977 | 0.67464
-40.4118 | -40.7496 |  |  | 0.65438 | 0.67943
-40.3306 | -40.6626 |  |  | 0.65899 | 0.68421
-40.2391 | -40.6576 |  |  | 0.66359 | 0.689
-40.2181 | -40.5449 |  |  | 0.6682 | 0.69378
-40.1503 | -40.4542 |  |  | 0.67281 | 0.69856
-40.1392 | -40.3999 |  |  | 0.67742 | 0.70335
-40.0405 | -40.3711 |  |  | 0.68203 | 0.70813
-40.0133 | -40.314 |  |  | 0.68664 | 0.71292
-39.9423 | -40.229 |  |  | 0.69124 | 0.7177
-39.9357 | -40.229 |  |  | 0.69585 | 0.72249
-39.8972 | -40.0897 |  |  | 0.70046 | 0.72727
-39.8972 | -39.9788 |  |  | 0.70507 | 0.73206
-39.8767 | -39.9788 |  |  | 0.70968 | 0.73684
-39.8464 | -39.9782 |  |  | 0.71429 | 0.74163
-39.7072 | -39.9242 |  |  | 0.71889 | 0.74641
-39.7072 | -39.897 |  |  | 0.7235 | 0.7512
-39.6609 | -39.897 |  |  | 0.72811 | 0.75598
-39.6392 | -39.897 |  |  | 0.73272 | 0.76077
-39.5632 | -39.7356 |  |  | 0.73733 | 0.76555
-39.5396 | -39.709 |  |  | 0.74194 | 0.77033
-39.4916 | -39.4991 |  |  | 0.74654 | 0.77512
-39.2856 | -39.471 |  |  | 0.75115 | 0.7799
-39.251 | -39.3704 |  |  | 0.75576 | 0.78469
-39.1784 | -39.3704 |  |  | 0.76037 | 0.78947
-39.038 | -39.3704 |  |  | 0.76498 | 0.79426
-39.038 | -39.3635 |  |  | 0.76959 | 0.79904
-39.0206 | -39.3195 |  |  | 0.77419 | 0.80383
-39.0206 | -39.2981 |  |  | 0.7788 | 0.80861
-38.9327 | -39.2941 |  |  | 0.78341 | 0.8134
-38.8998 | -39.2941 |  |  | 0.78802 | 0.81818
-38.8486 | -39.2889 |  |  | 0.79263 | 0.82297
-38.834 | -39.1549 |  |  | 0.79724 | 0.82775
-38.8159 | -39.1107 |  |  | 0.80184 | 0.83254
-38.7134 | -38.8499 |  |  | 0.80645 | 0.83732
-38.6255 | -38.778 |  |  | 0.81106 | 0.84211
-38.5309 | -38.7304 |  |  | 0.81567 | 0.84689
-38.5145 | -38.636 |  |  | 0.82028 | 0.85167
-38.5145 | -38.4962 |  |  | 0.82488 | 0.85646
-38.3521 | -38.2908 |  |  | 0.82949 | 0.86124
-38.1141 | -38.1787 |  |  | 0.8341 | 0.86603
-37.8141 | -38.1787 |  |  | 0.83871 | 0.87081
-37.6521 | -38.1787 |  |  | 0.84332 | 0.8756
-37.2934 | -38.1121 |  |  | 0.84793 | 0.88038
-37.2654 | -37.9371 |  |  | 0.85253 | 0.88517
-37.1522 | -37.8793 |  |  | 0.85714 | 0.88995
-36.9994 | -37.7443 |  |  | 0.86175 | 0.89474
-36.8223 | -37.66 |  |  | 0.86636 | 0.89952
-36.2005 | -37.5765 |  |  | 0.87097 | 0.90431
-36.0568 | -37.3711 |  |  | 0.87558 | 0.90909
-35.9081 | -37.3306 |  |  | 0.88018 | 0.91388
-35.8232 | -37.1309 |  |  | 0.88479 | 0.91866
-35.7335 | -37.0916 |  |  | 0.8894 | 0.92344
-35.6557 | -36.9744 |  |  | 0.89401 | 0.92823
-35.5217 | -36.9357 |  |  | 0.89862 | 0.93301
-35.2705 | -36.7637 |  |  | 0.90323 | 0.9378
-35.1328 | -36.7448 |  |  | 0.90783 | 0.94258
-35.0938 | -36.558 |  |  | 0.91244 | 0.94737
-35.0928 | -36.3933 |  |  | 0.91705 | 0.95215
-35.0036 | -35.9342 |  |  | 0.92166 | 0.95694
-34.9273 | -35.6626 |  |  | 0.92627 | 0.96172
-34.9207 | -35.32 |  |  | 0.93088 | 0.96651
-34.8807 | -35.2244 |  |  | 0.93548 | 0.97129
-34.7026 | -34.7165 |  |  | 0.94009 | 0.97608
-34.2961 | -34.4225 |  |  | 0.9447 | 0.98086
-34.2916 | -34.2366 |  |  | 0.94931 | 0.98565
-34.1504 | -34.1943 |  |  | 0.95392 | 0.99043
-33.9538 | -33.7551 |  |  | 0.95853 | 0.99522
-33.7716 | -33.5828 |  |  | 0.96313 | 1
-33.2302 | -28.786 |  |  | 0.96774
-32.828 | -14.2889 |  |  | 0.97235
-32.8066 |  |  |  | 0.97696
-32.7561 |  |  |  | 0.98157
-32.702 |  |  |  | 0.98618
-32.6593 |  |  |  | 0.99078
-32.0732 |  |  |  | 0.99539
-31.1853 |  |  |  | 1
-31.0485

## gain graphs
% E-field "gain" exceeds low, high band cutoff of: | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band | 802.11b (Wi-Fi) | Bluetooth
27,17 dB | 0% | 0% | 0% | 0%
26,16 dB | 0.4587% | 0% | 14.81% | 0%
25,15 dB | 1.376% | 0% | 22.22% | 0%
24,14 dB | 1.835% | 1.429% | 22.22% | 3.333%
23,13 dB | 3.211% | 10% | 40.74% | 10%
22,12 dB | 5.505% | 21.43% | 51.85% | 10%
21,11 dB | 6.881% | 30.48% | 74.07% | 10%