Minutes

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Monday, December 15, 2008

Etymotic Research

Elk Grove Village, IL

Participants

Stephen Berger

Steve Julstrom

Tom Knipple

Mead Killion

Bob Hofmann

By phone

Jim Turner

Dave Case

Joe Morissey

Dave Preves

Justin Chow

Donald Bowen

Tom Victorian

Meeting Agenda

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Monday, December 15, 2008 between 9:00AM – 5:00 PM CT

Location: 	Etymotic Research
	61 Martin Ln.

Elk Grove Village, IL

Dial-in Number: 	1-218-936-4700

Access Code: 	626642

- Call to order and introductions

- Approval of agenda

- Review and approval of past minutes

- Patent Policy

- Review of draft and status of items listed on website.

Lunch

- Demonstration – Generalized Test Method

- Continued work on draft

- Workplan & schedule

- Next meeting: TBD

- Adjourn

Minutes

- The chair called the meeting to order at 9:15 AM CT.   Participants were identified and self-introductions were made.

- The agenda was reviewed and approved.

- Minutes for the November 17 telecon were approved.

- The chair reviewed slides on the patent policy.

- A review of draft, revision 1.5 was held, with discussion on various points.  A number of decisions were made of what should be placed in the next version of the draft.  The following is a listing of the decisions made during the review:

  - Tom Victorian moved that the GTEM limits for hearing aid immunity be reduced by 7 dB from the values used for dipole illumination.  The motion was preceded by a presentation by Tom Victorian of data on the relative immunity levels between GTEM and dipole illumination.  During discussion other data from older studies was displayed and discussed.  

The motion passed unanimously.

  - Mead Killion moved to just keep the logarithmic form of the limit tables in Clause 8 and create an Annex to hold the linear version of the tables and move the linear version of the limit tables to that annex, as a convenience for users of the standard.

The motion passed unanimously.

  - Steve Julstrom moved to accept the spectral and temporal weighting values contained in draft revision 1.5 Annex D.6 and D.7 and maintain those values in the next version.

The motion passed unanimously.

  - Mead Killion moved that the table in Annex D.6 be limited to a single digit of precision.

The motion passed unanimously.

  - A discussion was held of the wording in subclause 5.3.3 “WD setup and use”.  After discussion and collaborative editing a wording was adopted for use in the next version of the draft.

  - An action item was assigned to the chair to add a footnote to Table 8-7 clarifying that the values are relative to the field strength of a CW signal that when modulated by 1 kHz 80% AM would produce the same reading as the signal under test, weighted by the spectral and temporal filters used in the standard.

  - After discussion it was decided to name the factor developed in Annex C.3.1 the Modulation Interference Factor (MIF).

  - After discussion Steve Julstrom moved to reduce the values in Table 7-7 by 6 dB.  The change was made to make the table appropriate for use with the generalized test method being adopted in the draft.  This change was made to maintain an approximate equivalence for GSM to the current table.  In discussion the expectation of the WG is that it will be approximately 2 dB easier for GSM phones to pass the new limits adopted.

Motion passed unanimously

  - A discussion was held and wording developed for a new definition of “RF Audio Interference Level” as the term to be used in describing the value being measured in the standard.  The final wording was adopted without opposition.

  - Tom Knipple moved to remove the requirement for H-Field emissions testing.

After discussion the motion passed with 1 negative vote.

  - A discussion was held and agreement reached without opposition to change Annex C.3.1 to make clear that the MIF could be done by conducted measurements so as to be usable in applying the value to Clause 4 evaluations of the potential of a technology to cause hearing aid interference.

- After the review of the draft Steve Julstrom demonstrated his implementation of the spectral and temporal weighting and measurement of the MIF.  Measurements were made on GSM, TDMA and CDMA phones.

- A WG3 telecon was planned for:
   - Wednesday Jan. 21 at 11:30 AM ET
   - Additional telecons may be set by the WG during the January telecon 

The next in-person meeting was planned for April 20, 2009 during the next C63 series of meetings at the IEEE headquarters in Piscataway, NJ, with time to be determined by the chair.  The next SC8 meeting will be the following morning on April 21, 2009.

- Meeting was adjourned at 5:30 PM

Summary of Action Items

| # | Item | Status |
| --- | --- | --- |
|  | Action Items from December 15, 2008 WG Meeting |
| 1 | Action item: Stephen Berger is to prepare a new draft, reflecting the decisions of the meeting. |  |
| 9 | Action item: |  |
| 10 | Action item: |  |