COMPATIBILITY BETWEEN

HEARING AIDS

AND

WIRELESS DEVICES

H. Stephen Berger

Siemens Business Communications Systems

COMPATIBILITY BETWEEN

HEARING AIDS

AND

WIRELESS DEVICES

H. Stephen Berger

Siemens Business Communications Systems

ABSTRACT

In late 1995 the FCC called the cellular phone industry and hearing instrument industry together to address the issue of the hearing aid users ability to use cellular phones.  At that point it was becoming increasingly apparent that cellular phones using digital transmission formats could interfere with many hearing aids.  This paper presents some of the technical challenges and measurement techniques developed to resolve this interference issue.

INTRODUCTION

Early in 1996 the FCC called together a Summit between the hearing industry, the wireless industry and consumers to resolve the compatibility issue between hearing aids and cellular phones.  As a result of those efforts, in the Spring of 1996 ANSI ASC 63 formed task group C63.19 to develop a measurement standard for hearing aid compatibility with wireless communications devices.  The goal of this task group is to develop a set of tests which would evaluate the compatibility of hearing aids with cellular phones.  The challenges presented to the task group were formidable.  In order to accomplish this task several significant technical issues had to be faced.  This technical undertaking came to include three research project, and over 60 engineers from 30 different companies working on the standard.  This paper discusses the technical challenges of this task and the rational for the resolutions which have been reached.

The essence of the problem is that the RF energy transmitted by a wireless device is received by the circuitry in many hearing aids.  Once the energy is in the hearing aid it may be audio rectified across some non-linear junction, resulting in a "buzz".  Significant effort has been invested in understanding and addressing this issue.  This mechanism of interference is well known.  The challenge in this case is that hearing aid wearers want to be able to use cellular phones.  This means that the hearing aid must be located well into the near-field region of the transmitting antenna.  Accordingly an evaluation of the immunity of the hearing aid must be of that immunity in the near-field environment.

A second challenge faced is that, in the near-field, the fields from a wireless device are highly variable in intensity and field impedance.  The quantification of the environment in which a hearing aid must operate presents a significant challenge.  Movements of only a centimeter can produce significant changes in the field magnitude or impedance.

A third challenge is presented by the presence of the hearing aid wearer. The human tissue in the head and hand has a very significant influence on the field generated by the cellular phone.  The question of how to properly account for this field deformation when evaluating a hearing aid's immunity presents special challenges.

A fourth challenge is that many hearing aids are equipped with a magnetic coupling mode, called the TeleCoil (t-coil) mode, in addition to the primary audio coupling mode.  Testing for compatibility in this mode has its own set of challenges.  For example, in this mode there is the possibility of RF interference and in addition electronic noise in the kHz region adds a second, independent, source of interference with the desired reception.

This paper will discuss how these and other measurement challenges can be meet.  The resulting measurement techniques allow the accurate evaluation of the system performance of a hearing aid used with the new generation of cellular phones or other wireless communications device.  The resulting tests present new test methodology for near-field evaluation of system immunity.  This addition brings a valuable evaluation tool to compliment the more mature far-field evaluation techniques, which are available.

QUANTIFICATION OF THE OBJECTIVE

The objective of the wireless hearing aid compatibility effort is to provide for good system performance between hearing aids and wireless devices.  This section discusses the logic by which this objective is parameterized and target values are assigned to those constituent parameters.

The first step is to normalize the test for normal speech.  65 dB SPL (sound pressure level referenced to 20 mPascals) is a nominal value often assigned to normal speech.  For the purposes of this discussion, the complexities of frequency differences within the audio frequency band will be disregarded.  So, for any given frequency in the audio frequency band a level of 65 dB SPL is assumed to deliver normal speech.  A hearing aid increases this nominal level by some gain factor in order to compensate for the user's hearing loss.  Other processing may be involved, such as noise cancellation or automatic gain control.  However, at any given point in time a hearing aid may be assumed to receive a signal, nominally at 65 dB SPL, and deliver to the user a signal of (65 + G) dB SPL, where G is the gain at the frequency of interest.

In order to be intelligible the signal delivered to the user must have an acceptable signal to noise or signal to interference ratio.  Three performance categories have come to be commonly used, which may be characterized as: usable in an emergency, acceptable for normal use, and excellent performance.  Various research studies  have indicated that a speech to interference ratio of 20-26 dB is acceptable for normal use.  Further work is underway to refine this range to a single number, defining the level acceptable for normal use.  A 10 dB degradation, to a range of 10-16 dB, yields a system which could be characterized as usable in an emergency but inadequate for regular use.  Alternatively, an improvement of 10 dB, to a range of 30-36 dB, yields an excellent performance level, where there is little discernible noise or interference.  Taking 20 dB signal to interference as a nominal target performance criterion results in a (45 + G) dB SPL level for interference, measured at the output.  Equivalent target parameters may be derived for the other performance categories.

The target performance is then parameterized and allocated between the system components, the phone and hearing aid.  The picture is further complicated by the fact that this equipment operates within the near-field region of the phone's antenna.  The picture is further complicated by the field perturbation created by the user's head and hand and other nearby objects.  Setting these complications aside for the moment, a general division must be made between the E and H field performance of the equipment.  So a target immunity for both the E and H field must be set for the hearing aid.  These immunity targets must be coordinated for both near-field E and H field of the radiating device, in the area controlled for the use of the hearing aid.  Hence, when phones provide fields at or below the emission limits and hearing aids exhibit immunities at or above the immunity targets the required audio performance will be delivered.

Next, the t-coil coupling mode must be dealt with.  In this mode a hearing aid receives its signal not by a microphone but through an inductive pickup coil.  The intended magnetic signal for the t-coil is that produced by the audio signal driving the phone receiver element or some other inductive circuit element.  In the ideal case a hearing aid delivers the same audio level whether in audio or t-coil mode.  Going back to the 65 dB SPL level for normal speech, some magnetic field emission level may be assigned to it.  The hearing aid in t-coil mode receives this signal inductively and amplifies it to the same (65 +G) dB SPL level which it would deliver if the user were using the microphone mode.  The intended advantage of this system is that any audio background noise is not received and so the user receives a better signal.

However, in the case of a wireless phone there is a lot of electronics in the handset, unlike landline phone handsets.  The electronics also emits magnetic fields from the currents in the circuitry.  These emissions are received as noise to the intended audio signal.  In order to deliver the desired performance two criterion must be met.  First the user of a hearing aid in t-coil mode should receive a signal of (65+G) dB SPL for a normal speech level emission.  Second that signal should have a noise component of no more than (45+G) dB SPL.  In paramaterizing this the phone must deliver a magnetic field of the intended audio signal within some defined target window.  The hearing aid is then responsible for receiving this signal and amplifying it to the intended (65+G) dB SPL level.

This picture is further complicated by orientation effects.  If the hearing aid's inductive coil is cross polarized to the field then a significant loss of reception will result.  This must be left for the user to adjust for optimum reception.

In the t-coil mode there are two distinct noise sources.  The currents in the phone's circuits will create magnetic emissions in the audio frequency band used by the intended signal.  Secondly, the RF fields may be audio rectified within the t-coil circuits exactly as they are for the microphone mode.  The combined effect of these independent interference sources must be below the (45+G) dB SPL level.

In order to achieve the desired target performance for t-coil mode a number of parameters must be met by the equipment involved.  For the phone a magnetic emission must be delivered which is within some defined relationship to its acoustic output.  Further, this signal must have an intended signal to interference ratio of 21 dB or better.  The change from a 20 to a 21 dB target allows for the addition of RF interference noise.  Other partitions, which sum to the same 20 dB, final performance figure are possible.  The phone must also provide RF E and H fields below set levels in the area intended for use by the hearing aid.  The hearing aid must then receive and amplify the magnetic field, provided by the phone, to the same level  it would the matching acoustic signal.  It must not add any significant noise to the signal from its own circuitry.  In addition it must provide an immunity to the RF E and H fields such that the signal received at 21 dB signal to interference is delivered at the desired 20 dB quality level.

In summary then, the following parameters must be established and observed for wireless phones and hearing aids to deliver the desired system performance:

In the Acoustic Mode:

1. Phones must provide an area for hearing aid use at or below established E and H field limits.

2. Hearing aids must provide RF immunity at or above the established RF E and H field levels.

In the T-Coil Mode:

1. Phones must provide an audio frequency magnetic field emission of the intended signal in an established relationship to its acoustic signal.

2. The phones magnetic field signal must have an acceptable signal to interference ratio, for the audio frequency signal.

3. The phone's magnetic field and the hearing aid's inductive coil must be properly aligned for maximum reception.

4. Phones must provide an area for hearing aid use at or below established RF E and H field limits.

5. Hearing aids must provide RF immunity, in t-coil mode, at or above the established RF E and H field levels.

MEASUREMENT REQUIREMENTS

These parameters result in the need for five measurements in order to assure that the intended performance will be achieved.  These measurements are:

1. Measurement of the phone's RF E and H fields in an area prescribed for hearing aid use.

2. Measurement of a hearing aid's immunity to both RF E and H field emissions, in both acoustic and t-coil mode.

3. Measurement of the phone's audio frequency magnetic field emission level.

4. Measurement of the phone's audio frequency magnetic field emission signal to noise ratio.

5. Measurement of the equivalence of the hearing aid's t-coil mode gain to its acoustic gain.

6. Measurement of the hearing aid's delivered signal to noise ratio, in t-coil mode.

TELEPHONE RF EMISSIONS TESTING

In the current draft of ANSI C63.19 an area 5 cm square is controlled for use by a hearing aid (See Figure 1).  This area is defined as being 1 cm from the surface of the phone, over the area of the receiver.  Specifically it begins at the top edge of the body of the phone and comes down over the receiver area of the phone.  It is centered, left to right, about the centerline of the phone.  The field strength in this area is to be scanned using near-field probes for both the E and H field.  Because the area is deep within the near-field of the antenna sharp field gradients are common.  Field variations of over 100 V/m can be found within this 5 cm square area.  Hence, it is vital that careful scans of the entire area be made to accurately assess the peak field potential.  Readings should be taken at increments of 2.5 cm and closer spacings are preferable.

Figure 1 - Control Area for Hearing Aid Use

The concern in assessing these fields is the interference potential of the emissions.  Therefore, it is the peak field value which is of interest.  Probes and their supporting instrumentation system must be capable of fast response.  For example, in TDMA protocols a unit transmits in a specified time slot.  It is common for the transmitter to function only one eighth of the time.  So for a 1 watt system the average power transmitted will be .125 watts.  For this reason probes which use diode detection are preferred.  Even with these systems it is common to perform averages either in the A/D conversion process or at other locations in the instrumentation software.  For the purpose of assessing the interference potential the peak reading is the desired value.

HEARING AID RF IMMUNITY MEASUREMENTS

In considering the immunity of a hearing aid to RF emissions both the near-field and the far-field immunity is of interest.  At this time there is not sufficient experience to replace one test with the other.  Far-field immunity tests are well known.  IEC 1000-4-3 gives the internationally recommended test practice for such an assessment.  In the case of hearing aids, which may be used with wireless phones, the hearing aids will be required to operate acceptably in the very near-field of the telephone's antenna.  For this reason a near-field assessment must be performed.  The recommended method for performing this assessment  is with a resonant half-wave dipole.

The near-field test is performed by feeding a resonant dipole, tuned to the prescribed frequency with a specified amount of power.  The hearing aid being evaluated is moved along the length of the dipole at a specified distance.  The fields produced by the dipole are predominately H field near the junction of the dipole and predominately E field near the tip of either element.  So by exposing a hearing aid to the field near the dipole junction and to the area near the tip both an H and E field immunity test is performed. In order to achieve good repeatability great care must be taken to control the spatial parameters involved.  The region which contains the peak emission within 10% is less than 1 cm in any direction.  In the direction tangential to the dipole the peak field falls off by more than 10% with a very few millimeters.

The test is repeated with the dipole oriented in three mutually orthogonal directions.  The surface of the hearing aid nearest the dipole during the test should be the one which will normally be closest to the phone.

AUDIO FREQUENCY EMISSIONS TESTING

Standard measurement techniques for the audio frequency magnetic emission of landline phones have existed for some time.  FCC Part 68 sets for the required limits for these signals and EIA RS-504 and IEEE 1027 set forth the specific measurement procedures.  However, these procedures, while helpful, are not sufficient for the situation with wireless telephones.  The levels established for landline telephones are not appropriate for the technologies being used for wireless devices.  Three major factors must be considered in addition to the factors involved in testing landline phones for FCC Part 68 hearing aid compatibility compliance.  First, the RF transmission of the phone may affect the test equipment being used and create inaccuracies in the testing.  This is often circumvented by redirecting the RF to an auxiliary antenna port on the telephone under test.  However, even in this case the residual RF leakage can produce fields of 30 V/m or more in the region very close to the phone.  The test procedure used must assure that any RF, including any residual RF, does not affect the test.

The second major issue is the question of signal quality.  With landline phones it is assumed that there is not a significant possibility of noise emissions adding to the intended audio magnetic emission.  So if a magnetic field of the required amplitude is measured it is assumed that this signal is produced by, and carries, the acoustic signal being delivered by the telephone receiver.  This is not the case with wireless phones.  Battery current surges, keyboard scanning and display currents can produce significant emissions of their own which will add noise to a t-coil mode hearing aid.  It is important to measure first the emission level without signal, which is the noise level from these other sources.  Then the signal is introduced at its intended level and the measurement is remade.  The second measurement is a signal plus noise measurement.  Thus a signal plus noise to noise assessment of signal quality may be made.  The division being made is that the telephone must deliver an audio frequency magnetic emission with an acceptable signal plus noise to noise figure.  The hearing aid must then deliver this signal to the user without adding significant additional noise to it, either from its own circuitry or, more probably, by audio rectification of RF emissions.

The third consideration is the assurance that the telephone's mode of operation during the test is truly representative of its normal operating mode.  The advanced signal processing algorithms used in wireless telephones often will treat a test signal in a very different manner than an actual voice signal.  This will be discussed more fully in a separate section below.  However, the simple test signals which are fully adequate for landline phones are not effective in wireless telephones.  For example, currently it is most common to induce a test signal at a specified voltage onto the tip and ring lines to a landline telephone being tested.  The signal path from that point is sufficiently standard that the acoustic signal emitted from the phone is reasonably determined.  So for the magnetic field test the test signal is put onto the telephone lines and the magnetic field at the receiver is measured.  An alternative, which does not assume this fixed relationship between input and acoustic output, calls for the receiver output to be established at a specified volume.  Then the magnetic field is measured.  Thus the relationship being sought is a fixed relationship between the acoustic output and the magnetic field output.  Care must still be taken to assure that the test signal being used has the wireless phone's various gain and other settings at similar levels to those created by an actual voice signal.

USE OF HUMAN STIMULANT

It is well known that the presence of the user perturbs the electromagnetic field in a number of ways.  For example, because human tissue is a reasonably low impedance material, being mostly conductive liquid, E-fields are quickly attenuated.  E-fields near to human tissue will change orientation to be predominately normal to the tissue.  This is because the tissue will not support the potential difference which must exist with parallel E-fields.  In contrast, H-fields are relatively unaffected by the presence of the tissue material and will propagate through it quite well.  Hence, the presence of a user will alter the E-field component of a field but leave the H-field relatively unaffected.  Further, in some units the loading affect of the user's presence may reflect back into the antenna's VSWR, substantially affecting the efficiency of energy transfer from the power amplifier into the antenna.  Thus the presence of the user complicates the electromagnetic picture substantially.

Human simulations or phantoms have been used in the field of radiation hazard assessment for some time.  However, there is not yet a generally accepted standard phantom.  A number of sizes and material types exist.  While solid materials do exist and are utilized in human phantoms liquid tissue simulants are more common.  A number of recipes exist in the literature for both liquid and solid human tissue simulants.  Typically the complexity of creating a phantom of more than one tissue type is considered unduly burdensome.  Phantoms are typically constructed of a single tissue type.  This leads to difficulties in deciding which material to use or what set of materials should be used.  In addition, the typical spacing between a user and a wireless phone is very small.  The use of a phantom presents geometric challenges as to how the measurement probes are to be inserted and moved through the small space available.

In general it is felt that the demands of using phantoms are too large for regular product testing.  Rather phantoms should be used in foundational research.  Transfer functions can be derived which allow product evaluations to be made in a more efficient test environment with appropriate adjustments having been made in the criterion used.  However, some solutions may be proposed which can only be fully assessed with the complexity of a phantom simulator.  In these cases a phantom should be used.

SPECIAL TEST SIGNAL CONSIDERATIONS

There are special considerations which must be made of the test signal used for these tests.  It is important to understand the signal processing which takes place in either a hearing aid or a wireless phone in order to correctly tests either.  If a hearing aid has automatic gain control or other signal processing, such as noise cancellation circuitry, a test signal must be provided, to assure that the hearing aid is operating in its normal mode with a typical gain setting.  If no signal is present the gain may be unrealistically high.  If there is too much signal, some hearing aids will limit or go into compression.  So a signal in the range for normal speech must be provided for these units.

Wireless telephones use very sophisticate signal processing and conditioning in their vocoders (Voice Coder-Decoder).  Complex algorithms compress the speech signal, improve the signal to noise ratio and compress the audio signal for transmission.  Increasingly hearing aids are also using similar techniques to improve their performance.  With this increasing use of signal processing care must be taken to provide a signal which does not activate any noise cancellation circuitry.  Steady tones or other signals which an algorithm may interpret as noise can result in automatic adjustments to the gain.  This effect can be hard to detect as the signal will be delivered to the output.  However, with a normal speech signal the gain would be set much differently than it is for a steady tone.  This results in misleading measurements of any noise signals produced during the test.  Typically the test signal will need to be intermittent, like real speech, and have some spread of frequency content.  Different test signals are required for different vocoder algorithms.  However, the increased complexity of these test signals increases the difficulty of properly determining the interference level created during the test.  Several test standards provide additional guidance for properly dealing with these issues.

SUMMARY

The development of cellular phones and hearing aids which can be effectively used together is being vigorously persued by both industries.  The requirements of such a system have been set forth in the measurement requirements.  In audio coupling mode the new requirements the cellular phone environment delivers from the traditional landline phone is the RF transmitted by the phone.  Hence, the requirement is to coordinate the immunity of the hearing aid with the electromagnetic environment in which it must operate.

In T-coil coupling mode the need to coordinate the RF environment with the hearing aid's immunity also exists.  In addition the phone must deliver an audio frequency magnetic signal of an amplitude and signal quality so as to be usable by the hearing aid.

Significant advances have been made on each of these fronts.  The most sensitive circuits within hearing aids have been analyzed and various techniques have proven effective in improving the immunity.  In some cases component manufacturers to the hearing aid industry are incorporating solutions at the component level.  In a similar vein much research has been focused on the nature of the electromagnetic field created around the receiver section of the phone.  Various remedies have been explored.  However, this issue is more complicated because the transmission of RF energy is necessary function of cellular phones.  This is unlike the case for hearing aids, where the reception of RF energy into the audio circuits is an unintended effect.  So anything which reduces the RF reception in hearing aids in a cost effective and producible manner can be utilized.  However, changes to the RF environment of the receiver area of the phone must be carefully coordinated with the specifications for signal transmission to the base station.

For the T-coil mode the primary focus has been on reducing the unwanted emission in the audio frequency band and increasing the desired signal, in an attempt to improve the signal to interference ratio.  Emission sources such as battery surge currents, keyboard scan and display currents are all being addressed.  Various techniques are applied such as improved layout to reduce circuit loop size and increased decoupling on primary current paths.  Considerable improvements are being reported in this area.

SUMMARY

The ever increasing use of electronics creates defacto systems which must operate acceptably.  However, these systems may not be under the control of an single company or even a single industry.  The case discussed here of wireless telephones and hearing aids is a case in point.  There is a societal desire that these devices operate well together.  However, to accomplish this significant technical barriers must be overcome.  This paper has discussed the measurement challenges which must be met to properly control the system parameters involved in this system.  The ANSI committee C63.19, having over 60 members representing more than 30 companies and organizations, has been working since April 1996 to address these issues in a consensus standard.  The primary parameters controlling this issue have been identified.  Significant efforts are underway to develop product solutions both in improved hearing aid immunity and cellular phone signal quality.  However, as with any issue of this complexity significant work remains to be done.

BIBLIOGRAPHY

[1] 	EIA RS-504-1983, "Magnetic Field Intensity Criteria for Telephone Compatibility with Hearing Aids.

[2] 	EHIMA GSM Project, Development phase, Project Report (1st part) Revision A, Technical-Audiological Laboratory and Telecom Denmark, October 1993.

[3]	EHIMA GSM Project, Development phase, Part II Project Report, Technical-Audiological Laboratory and Telecom Denmark, June 1994.

[4]	EHIMA GSM Project Final Report, Hearing Aids and GSM Mobile Telephones: Interference Problems, Methods of Measurement and Levels of Immunity, Technical-Audiological Laboratory and Telecom Denmark, 1995.

[5]	FCC 47 CFR 68, Connection of Terminal Equipment to the Telephone Network (1997).

[6]	Hearing Aids/GSM, Report from OTWIDAN, Technical-Audiological Laboratory and Telecom Denmark, April 1993.

[7]	Joyner, K.H. et al, Interference to Hearing Aids by the New Digital Mobile Telephone System, Global System for Mobile (GSM) Communication Standard, National Acoustic Laboratory, Sydney 1993.

[8]	Joyner, K.H. et al, NAL Report #131, National Acoustic Laboratory, Sydney.

[9] 	IEC 118-4: 1981, Methods of Measurement of Electroacoustical Characteristics of Hearing Aids - Part 4: Magnetic Field Strength in Audio-Frequency Induction Loops for Hearing Aid Purposes.

[10] 	IEC 118-13:199x (Draft), Methods of Measurement of Electroacoustical Characteristics of Hearing Aids - Part 13: IEC 77B/153/CD June, 1995.

[11] 	IEEE Std 1027-1996, IEEE Standard Method for Measurement of the Magnetic Field in the Vicinity of a Telephone Receiver.

[12]	McCandless, G.A. & Lyregaard, P.E., Prescription of Gain/Output (POGO) for Hearing Aids, Hearing Instruments 1:16-21, 1983.

[13]	Technical Report, GSM 05.90, GSM EMC Considerations, European Telecommunications Standards Institute, January 1993.

The three primary North American Studies have been performed by the University of Oklahoma, Dr. Levitt of CUNY with Dr. Harkin of Gaulludet University and Dr. Killion of Etomotic Research.  The research focus and protocol of each study was different.  However, the conclusions of these studies has shown remarkable consistency on the fundamental issues involved.