HEARING AID

AND

CELLULAR PHONE

COMPATIBILITY

--

WORKING TOWARD SOLUTIONS

H. Stephen Berger

Siemens Business Communications Systems

HEARING AID AND CELLULAR PHONE

COMPATIBILITY

--

WORKING TOWARD SOLUTIONS

H. Stephen Berger

Siemens Business Communications Systems

ABSTRACT

The ability of hearing aid wearers to use cellular phones using digital transmission formats has become a prominent issue.  This paper presents the requirements, the technical challenges and the emerging solutions for the combined use of hearing aids with cellular phones.

INTRODUCTION

Cellular phones increasingly utilize digital transmission formats in order to gain a variety of advantages made available by these advanced signal encoding schemes.  However, as with so many technical advances, there are also undesired side effects.  A prominent example, which is currently getting a great deal of attention is the interference potential that digital transmission formats have for many hearing aids.  This paper examines the causes of this interference and presents the solutions which are emerging to address it.

In order for a hearing aid to operate effectively with any phone, including a cellular phone it must have two things: an input signal at the required level and a signal quality which is acceptably free of noise or interference.  Some hearing aids offer a tele-coil, also called T-Coil, mode, in which the hearing aid receives and amplifies the magnetic field produced by a phone rather than the acoustic signal.   So, for each reception mode of the hearing aid the requirements of signal strength and signal quality must be met to assure proper operation of the hearing aid.  In the case of cellular phones a very important aspect of signal quality is the potential of the RF transmission to interfere with the hearing aid circuitry.  Logically this is a sub-topic of signal quality.  However, because of its importance, in this paper the coordination of phone RF emissions with hearing aid RF immunity will be treated as a third requirement.  In summary then, in order to insure proper operation for a given reception mode the requirements of signal strength, signal quality and RF emission/immunity coordination must be met.

This paper will examine the issues presented by each of these requirements, the challenges and measurement techniques required to address them and finally solutions which are emerging.

| ACOUSTIC MODE (Microphone Mode) | MAGNETIC MODE (T-Coil Mode) |
| --- | --- |
| Adequate acoustic volume. | Adequate magnetic field intensity. |
| Low background noise. | Acceptable magnetic signal to noise ratio. |
| Coordinated RF emissions and immunity. | Coordinated RF emissions and immunity. |

Table 1 Operational Requirements

ACOUSTIC MODE CHALLENGES

In the acoustic mode we begin with the requirement of acoustic volume.  Because the delivery of speech is a basic requirement of cellular phones it can generally be assumed that there will be adequate acousitc volume.  In addition cellular phones almost universally provide volume control, which allows for user control of the volume.

The second requirement is signal quality.  In general cellular phone provide superior signal quality.  Background noise elimination circuits are commonly used and provide performance which is superior to landline phones.  However, because of the mobile nature of these products, noise in the user's environment presents a real challenge in some situations.  Cellular phones will be used in a much wider variety of environments than their landline cousins.

Finally, in order to have acceptable performance the immunity of the hearing aid to RF emissions must match the emissions present in the area of the phone in which the hearing aid is being used.  This is the most significant challenge for the acoustic mode of operation.  Hearing aids do not have any legal requirement for RF immunity.  In Europe, ETSI (European Telecommunications Standards Institute) is working with EHIMA (European Hearing Instrument Manufacturers Association) to establish a 3 V/m requirement and test standard, IEC 118-13.  Because there has not been a mandatory requirement hearing aids with a very broad range of immunity to RF can be found on the market.

The emissions of the phone are largely determined by the licensed operating power.  Typically cellular phones are licensed to transmit with 1 or 2 watts of RF power.  As the antenna and receiver are typically located near the top of the phone a hearing aid will most often be used very close to the base of the antenna.  In this area intense RF fields can be expected.  It is not uncommon to find phones and hearing aids which have a 30-40 dB gap between their emissions and immunity.  It is this difference which allows for the audio rectification and the resulting 'buzz' commonly experienced by hearing aids used with cellular phones.

T-COIL MODE CHALLENGES

In the case of the T-Coil mode the challenges are similar but have their own uniqueness.  The T-Coil was originally developed for use with analogue phones.  When this mechanism was developed the receiver elements in almost all phones produced large magnetic leakage fields.  These fields could be received and translated into the intended speech signal.  However, newer technologies are much more energy efficient and produce correspondingly lower magnetic field levels.  Some technologies produce almost no magnetic field.  In order to assure adequate magnetic field signal level additional circuitry is required.  This raises issues of additional cost, complexity and reduced battery life.

Magnetic field signal quality is a much more formidable issue than acoustic signal quality.  In the traditional landline phone the only electronics in the handset capable of producing a magnetic field is the receiver element.  However, in a cellular phone there is a lot of electronics in the handset, all in very close proximity to the receiver element.  Battery surge currents, power supply components, keyboard scanning and display refresh cycles all can create magnetic fields which add noise to the desired signal.  With the increasingly small size of cellular phones all of these functions operate in very close proximity to each other.

Several research studies are currently underway which will help quantify the signal to noise requirement.  Early data indicates that a 20-26 dB signal to interference ratio is a good range for normal use.  This conclusion is supported by the primary research studies in this area.  Further work is underway to refine this range to a single number, defining the level for normal use.  A 10 dB degradation, to a range of 10-16 dB, yields a system which could be characterized as usable, but inadequate for regular use.  Alternatively, an improvement of 10 dB, to a range of 30-36 dB, yields an excellent performance level, where there is little discernible noise or interference.

The third element is the coordination of the RF emissions and immunity.  This issue is identical for the two modes except that different circuit elements are involved in the T-Coil and Acoustic modes.

MOVING TOWARD SOLUTIONS

In order to deal with this problem, two kinds of equipment, cellular phones and hearing aids, developed by two entirely different industries, must be designed as a single system.  To accomplish this the critical parameters must be identified and specifications developed which each equipment manufacturer will follow.  In this way the immunity of the hearing aids can be matched to the emissions of the phone so as to deliver to the user the intended performance.

Early in 1996 the FCC called together a Summit between the hearing industry, the wireless industry and consumers to resolve the compatibility issue between hearing aids and cellular phones.  As a result of those efforts, in the Spring of 1996 ANSI ASC 63 formed task group C63.19 to develop a measurement standard for hearing aid compatibility with wireless communications devices.  The goal of this task group is to develop a set of tests which would evaluate the compatibility of hearing aids with cellular phones.  This technical undertaking came to include three research projects, and over 60 engineers from more than 30 different companies working together on the standard.

SYSTEM COORDINATION

As ANSI C63.19 is currently drafted, the standard uses a categorization system which is based upon coordinating the characteristics of a hearing aid and cellular phone.  The users will decide the level of performance they require and then obtain the equipment which delivers that performance.  By using a categorization system improvement can come from either the phone side or the hearing aid side of the equation.

In order to be intelligible the signal delivered to the user must have an acceptable signal to noise or signal to interference ratio.  Three performance categories are being proposed for use in the current draft, which may be characterized as: usable, acceptable for normal use, and excellent performance

The target performance is then parameterized and allocated between the system components, the phone and hearing aid.  A general division must be made between the E and H field performance of the equipment.  So a target immunity for both the E and H field must be set for the hearing aid.  These immunity targets must be coordinated for both near-field E and H field of the radiating device, in the area controlled for the use of the hearing aid.

In the current draft of ANSI C63.19 an area 5 cm square, which is sub-divided into 9 sub-areas, is evaluated for use by a hearing aid (See Figure 1).  This area is defined as being 1 cm from the surface of the phone, over the area of the receiver.  The field strength in this area is to be scanned using near-field probes for both the E and H field.  Because the area is deep within the near-field of the antenna sharp field gradients and RF 'hot spots' are common.  Up to four fold field variations can be found within this 5 cm square area.  The current draft requires that the center square and 5 other connecting squares meet the performance requirements.  This allows for isolated 'hot spots' near the antenna, while requiring 75% of the area be at or below the prescribed level.

Hence, when phones provide fields at or below the emission limits and hearing aids exhibit immunities at or above the immunity targets the required audio performance will be delivered.

The following table summarizes these results and guides the user in equipment selection.  The user begins by deciding on the level of performance required.  Each category, Usable, Normal Use and Excellent Performance, is determined by the delivered signal to interference ratio.  So if the need is only to use a phone infrequently, such as for emergency situations, a phone and hearing aid of the lowest category, U1 will be adequate.  By increasing either the phone or hearing aid one performance category, to U2, a system for normal use results.  Increasing either piece of equipment by two categories, to U3, or by increasing both by one category, to U2, results in an excellent performance level.  In this way maximum flexibility is provided to the user in their equipment choices.

A similar table is used for T-Coil mode.  The U designation has a T added to it if the piece of equipment is also rated for the T-Coil mode.  So for normal use a phone or hearing aid of category U2T should be used with the other piece of equipment being U1T or higher.

Figure 1 - Control Area for Hearing Aid Use

It should be noted that the lowest category limit imposes stringent requirements upon the equipment it governs.  A U1 or U1T rating is a rigorous requirement.  No level of system performance can be assured when unrated equipment is used.  This is even true when that equipment is used with a corresponding piece of equipment of the highest rating.

The advantages of this system are numerous.  First, it deals realistically with a user's current situation.  A user who owns either a phone or hearing aid and is buying the other can get the performance they desire by obtaining a higher category in the new piece of equipment.  While it remains true that performance with unrated equipment cannot be assured in the general case, it can be tested in particular cases.  A user can simply test the new product with their current phone or hearing aid.  Care should be taken that the phone is in its maximum power transmission mode when making such a comparison. So, for example, a person who

Figure 1- System Performance Classification Table

owns a phone and is buying a new hearing aid can try a category U1, U2 or U3 hearing aid to get the performance they desire.

A second advantage is that this system allows for various hearing aids and cellular phone technologies to implement solutions on various schedules.  There are many different types of hearing aids and cellular phones.  Different technologies are involved in each industry.  In addition, various companies have different schedules for development and introduction of new products.  If a single limit approach had been adopted, in a one-size-fits-all approach, it would inevitably have been poorly adapted to all but a select group of products in each industry.  On one extreme, any given limit would have been too demanding for some technologies.  A single limit would have severely disadvantaged some technologies and product types, which might be very useful for other advantages those products offered.  On the other extreme, different technologies would not have been challenged enough.  Some product types can be brought to much higher performance levels and should be challenged to do so.  The categorization system allows this diversity to take place.

EARLY SOLUTION

To provide a complete solution a system must met the requirements listed in Table 1, of coordinated emission and immunity and an appropriate signal strength, with good signal quality.  The earliest solutions offered have been in the form of accessories to standard products and specialized products.  Several headset type devices are on the market which offer acoustic output and an enhanced magnetic field output.  These products, by their nature, place a distance between the phone and the hearing aid.  This simple mechanism effectively deals with the RF interference problem.  The RF gradient is extremely large in the area close to the antenna.  As has been said, a change of only 5 cm has been measured in some models to reduce the field strength by a factor of four.  So the normal separation distance a headset equipped user will have is more than sufficient to deal with the interference problems.

Other types of solutions are being vigorously pursued.  High immunity hearing aids are already on the market.  One of the most promising developments in this area is the introduction of hearing aid components with immunity features built in.  Early in the FCC Summit process a technology exchange began between the cellular phone industry and the hearing aid industry.  Part of this exchange was the offering of assistance by RF specialists from the phone industry.  RF protection is most effective if built directly into sensitive circuit components such as amplifiers and microphones.  Today there are commercially available components with decoupling capacitors built into them.  Some product demonstrations have shown hearing aids with only two additional capacitors built into the appropriate components which are useable with major classes of cellular phones.

USER CONTROLLABLE OPTIONS

The research performed to date demonstrates that educated users can materially improve their situation.  At least four important parameters are under the user's control: equipment selection, maximization for acoustic reception, minimization  of the RF field and cross orientation of the emitter and reception circuit.  ANSI C63.19 is being written to allow categories of equipment.  When properly matched the user will receive the desired performance.  However, the user needs to know that properly matched equipment is required and that improperly matched equipment will often fail to met expectations.  The effective dissemination of user education material will be an important part of the final solution for this effort.

The maximum acoustic field is typically at the center of the handset receiver.  For best reception this is where the hearing aid microphone should be located.  For many hearing aids, especially the behind-the-ear types, the user may not hold the equipment for best reception.  In the behind-the-ear hearing aid the microphone is either above or behind the ear.  If the user holds the handset in the normal position, with the receiver centered on the user's ear, the microphone is not receiving the best acoustic signal.  Simply holding the handset so that the receiver and microphone are centered rather than the receiver and the ear being centered, can do a lot to improve reception.

The third thing the user can do is to avoid RF 'hot spots'.  Many handsets have 'hot spots', especially around the antenna and the feed to the base of the antenna.  The RF fields change very rapidly around these areas.  By holding the phone so that a hearing aid is just a little further from a high field area can make a lot of difference.  In many cases this means having the hearing aid on the far side of the receiver from the antenna.

Orientation is a third factor.  When the antenna or other RF emitter is lined up with the sensitive circuit in a hearing aid maximum interference will occur.  On the other hand when these two elements are cross polarized then a lot of reduction in the interference can be realized.  A rotation of a few degrees can yield a significant reduction.

SUMMARY

The development of cellular phones and hearing aids which can be effectively used together is being vigorously pursued by both the hearing instrument and cellular telephone industries.  The requirements of such a system are being set forth in the measurement requirements of ANSI C63.19. One of new factors in the cellular phone environment, which is different from the traditional landline phone, is that these phones must emit RF to operate.  Hence, the requirement is to coordinate the immunity of the hearing aid with the electromagnetic environment in which it must operate.

In T-coil mode the need to coordinate the RF environment with the hearing aid's immunity also exists.  In addition the phone must deliver an audio frequency magnetic signal of an amplitude and signal quality so as to be usable by the hearing aid.

Significant advances have been made on each of these fronts.  The most sensitive circuits within hearing aids have been analyzed and various techniques have proven effective in improving the immunity.  In some cases component manufacturers to the hearing aid industry are incorporating solutions at the component level.  In a similar vein much research has been focused on the nature of the electromagnetic field created around the receiver section of the phone.  Various remedies have been explored.  Headset type devices, special antennas and other products are already on the market.  Research has been announced, investigating remedies such as special battery designs which reduce power supply noise currents, special antennas and RF shielding to create a protected area for the use of the hearing aid.  However, this issue is more complicated because the transmission of RF energy is a necessary function of cellular phones.  This is unlike the case for hearing aids, where the reception of RF energy into the audio circuits is an unintended effect.  So anything which reduces the RF reception in hearing aids in a cost effective and producible manner can be utilized.  However, changes to the RF environment of the receiver area of the phone must be carefully coordinated with the specifications for signal transmission to the base station.

For the T-coil mode the primary focus has been on reducing the unwanted emission in the audio frequency band and increasing the desired signal, in an attempt to improve the signal to interference ratio.  Emission sources such as battery surge currents, keyboard scan and display currents are all being addressed.  Various techniques are applied such as improved layout to reduce circuit loop size and increased decoupling on primary current paths.  Considerable improvements are being reported in this area.

The ever increasing use of electronics creates defacto systems which must operate acceptably.  However, these systems may not be under the control of single company or even a single industry.  The case discussed here of wireless telephones and hearing aids is a case in point.  There is a societal desire that these devices operate well together.  However, to accomplish this goal significant technical barriers must be overcome.  This paper has discussed the challenges which must be met to achieve the desired system performance.

Significant efforts are underway to develop compatible products both in improved hearing aid immunity and cellular phone signal quality.  Solutions are already beginning to appear on the market on both the hearing aid and the phone side of the equation.

BIBLIOGRAPHY

[1]	ANSI C63.19:199X, American National Standard for Methods of Measurement of Compatibility between Wireless Communications Devices and Hearing Aids.

[2]	EHIMA GSM Project Final Report, Hearing Aids and GSM Mobile Telephones: Interference Problems, Methods of Measurement and Levels of Immunity, Technical-Audiological Laboratory and Telecom Denmark, 1995.

[3]	Joyner, K.H. et al, Interference to Hearing Aids by the New Digital Mobile Telephone System, Global System for Mobile (GSM) Communication Standard, National Acoustic Laboratory, Sydney 1993.

[4] 	IEC 118-13:199x (Draft), Methods of Measurement of Electroacoustical Characteristics of Hearing Aids - Part 13: IEC 77B/153/CD June, 1995.

The three primary North American Studies have been performed by the University of Oklahoma, Dr. Levitt of CUNY with Dr. Harkin of Gaulludet University and Dr. Killion of Etomotic Research with Harry Teder.  The research focus and protocol of each study was different.  However, the conclusions of these studies has shown remarkable consistency on the fundamental issues involved.