The following is a transcript of the address given in May 1997 at Gallaudent University.

The transcript is as recorded and may contain errors and omissions.

OUR LAST BY BY NO MEANS LEAST PRESENTER IS STEPHEN

BERGER FROM SIEMENS COMMUNICATIONS SYSTEMS.

STEVE INITIATED THE WORK ON STANDARDS IN ANSI C63 WHEN

EVERYONE ELSE WAS LOOKING AT EACH OTHER AND {GRUM}

LING.

HE INITIATED THE WORK AND IT ALLOWED PEOPLE TO GET TO

WORK ON THE PROBLEM.

STEVE IS SOMEONE YOU WOULD WANT TO CLONE, NOT ONLY FOR

HIS KNOWLEDGE OF EMC, BUT I'VE WORKED WITH HIM OR THIS

ISSUE AND THE TELECOM ADVISORY COMMISSION.

HE'S A SENIOR {234ING} ENGINEERING WHERE HE PROVIDES

TECHNICAL LEADERSHIP FOR HARDWARE DESIGN {SHINS}.

HE HAS WORKED FOR EMC FOR 16 YEARS AND HAS BEEN

HEAVILY INVOLVED IN THE DEVELOPMENT OF EMC STANDARDS

AT THE IEEE ANSI ASE.

CURRENTLY HE'S A CHAIR OF THE ANSI WORKING GRAB ON TEM

DEVICES AND IS U.S. NATIONAL EXPERT.

HE HAS WRITTEN NUMEROUS TECHNICAL PAPERS AND HOLDS

SEVERAL PATENTS.

STEVE?

WELCOME.

>> THANK YOU.

GOOD AFTERNOON.

THANK YOU.

I WOULD LIKE TO START THIS AFTERNOON IN A VERY GENERAL

WAY {FAUK}ING ABOUT THE STANDARD AND WHERE IT IS AND

WHAT IT DOES, AND THEN I'LL SLOWLY MOVE TO THE MORE

TECHNICAL AND SPECIFIC DETAILS OF WHAT IS IN THE

STANDARD, AND HOPEFULLY IT WON'T GET TOO EXCRUCIATING

INTO THE MINUTIA OF IT.

LET ME SAY THAT STANDARDS BY THEIR NATURE, CAPTURE A

LOT OF EFFORT AND CODIFY IT AND HOLD IT FOR FUTURE

USE, AND THAT'S EXACTLY WHAT'S HAPPENING HERE.

ALL OF THE WORK THAT'S REPRESENTED IN THE EARLIER

TALKS, HOPEFULLY, IS BEING ACCURATELY SYNTHESIZED AND

EMBODIED IN THE STANDARD IN ORDER TO ACCURATELY HOLD

THE SOLUTION TO THE PROBLEM.

I'VE BEEN PARTICULARLY PLEASED ON THIS COMMITTEE BY

HOW MUCH EFFORT AND GOOD WILL HAS GONE INTO IT.

THIS IS A PRODUCT OF COOPERATION BY ALL INVOLVED

PARTIES.

TOM VICTORIAN IS IN THE AUDIENCE, HE'S THE CO-CHAIR

FOR THE STANDARD AS WELL A PARTICIPANTS.

THERE HAVE BEEN OVER 60 ENGINEERS WHO HAVE COOPERATED

ON THIS STANDARD, REPRESENTING 32 COMPANIES AND

ORGANIZATIONS FROM ALL SIDES OF THE ISSUE AND THIS JOB

WOULD NOT BE WHERE IT IS WITHOUT ALL OF THAT EFFORT.

IN APRIL OF 1996 THE ANSI PROJECT WAS OPENED AND

ACTUALLY TWO PROJECTS WERE INITIATED, FIRST C63.19 AS

YOU'VE HEARD TO DETERMINE WHAT THE METHOD OF

MEASUREMENTS WOULD BE, WHAT ARE THE PARAMETERS WE NEED

TO CONTROL AND TO IDENTIFY THE EXACT MEASUREMENT

PROCEDURES THAT WOULD BE USED.

SECONDLY C63.20, IT WILL DEFINE THE ACTUAL PARAMETER

NUMBERS, THE LIMITS, IF YOU WILL, THAT ARE REQUIRED IN

ORDER FOR THE HEARING AID AND A PHONE TO OPERATE

TOGETHER AS A SYSTEM AND GIVE THE USER THE DESIRED

PRODUCT THAT THEY WANT AS FAR AS THE SOUND OUTPUT.

IT'S THE RESULT OF THE F.C.C. SUMMIT AS A DIRECT

OUTPUT OF THAT, AND HAS WORKED VERY CLOSELY WITH THE

THREE RESEARCH PROJECTS THAT YOU HEARD REPORTED THIS

AFTERNOON, THAT BY HARRY LEVITT AND JUDY HARKINS, THE

PROJECT AT OU AND MEAD KILLION AND HARRY TEDER'S

STUDY.

LAST MONTH IN CHICAGO, I BELIEVE THE COMMITTEE

ACCEPTED A VERY AGGRESSIVE TED YOU'LL, WE MEET NEXT ON

JUNE 16 THROUGH THE 18TH, WE'LL BE AT THE IEEE

HEADQUARTERS AND THE COMMITTEE HAS ACCEPTED A

CHALLENGE TO HAVE A DRAFT TURNED OVER TO THE EDITORS,

THAT'S THE REASON FOR THE LOCATION.

WE'LL BE MEETING WITH THE EDITORS AND LET THEM START

THEIR JOB OF MAKING SOMETHING OF THE WRITING OF

ENGINEERS INTO VALID ENGLISH.

AT THAT POINT WE WILL NOT START BALLOTTING THE

STANDARD, THIS IS UNUSUAL BY I THINK WISE IN THIS

CASE.

THERE IS PLANNED A PERIOD OF VALIDATION AND

CONFIRMATION OF THE STANDARD FOR THREE PRIMARY

PURPOSES.

FIRST, FOR CLARITY, AND THAT'S ALREADY STARTED.

CAN WE TAKE THE STANDARD, HAND IT TO SOMEONE WHO IS

NOT ON THE COMMITTEE AND HAVE THE INTENDED TEST

PERFORMED?

SECONDLY, FOR LAB TO LAB CORRELATION.

CAN WE MOVE A TEST OBJECT BETWEEN LABS AND GET THE

SAME ANSWER WITH THE METHODS THAT WE HAVE LAID OUT?

AND THIRDLY, FOR THE EFFICACY OF THE STANDARD.

DO THESE TESTS MEASURE THE REQUIRED PARAMETERS IN

ORDER TO ASSURE THE REQUIRED OUTPUT, THAT IS DO THINGS

THAT PASS THE STANDARD TO THE LIMIT SET FORTH, RESOLVE

THE PROBLEM AND DO THINGS WHICH DO NOT FAIL TO RESOLVE

THE PROBLEM.

THAT OF COURSE, WOULD BE THE INTENT.

THAT, WE'RE ALLOWING FOR TWO CYCLES OF THAT VALIDATION

BECAUSE THAT CAN GET COMPLEX.

IF ONLY ONE CYCLE IS COMPLETE WE'LL START BALLOTTING

IN THE EARLY FALL.

IF TWO CYCLES ARE REQUIRED WE'LL START BALLOTTING IN

THE LATE FALL.

I WOULD LIKE NOW TO TALK, JUST FOR A MINUTE ABOUT THE

PHYSICS OF THE SITUATION AND THIS WILL BE REPEATING

SOME OF THE OTHER TALKS BUT THE PHYSICS OF THE

SITUATION IS REALLY FAIRLY STRAIGHTFORWARD.

AN RF FIELD IS EMITTED BY THE PHONE THAT'S A NECESSARY

PART OF ITS FUNCTION.

THE FIELD REACHES THE HEARING AID CIRCUITRY AND

THROUGH VARIOUS PARTS OF THE CIRCUITRY WHICH ACT AS

ANTENNA ELEMENTS, THE WIRING PARTICULARLY, THE RF

ENERGY IS CONDUCTED TO NONLINEAR JUNCTIONS, TYPICALLY

IN THE MICROPHONE OR PROBABLY ON MOST IMPORTANTLY, THE

MICROPHONE OR THE FIRST TRAN SIFTER IN THE FIRST

AMPLIFIER, AND THERE A DEMODULATION TAKES PLACE, WHICH

IS KNOWN AS AUDIO RECTIFICATION.

THAT MEANS THAT THE RF ENERGY IS TURNED INTO A VOLTAGE

WHICH IS IDENTICAL TO THE SIGNAL THAT'S INTENDED TO BE

PASSED ON FROM THE MICROPHONE TO THE AMPLIFIER AND ON

THROUGH, AND THEN TURNED INTO A SOUND.

NOW, IF THE RF IS CHANGING AMPLY {TUD} AT A RATE WHICH

IS IN THE AUDIO FREQUENCY BAND YOU'RE GOING TO GET AN

AUDIO OUTPUT, AND THAT'S THE MECHANISM THAT WE'RE

DEALING WITH.

THERE IS AN IMPORTANT POINT HERE, AND THAT IS THAT ANY

RF WILL, ACROSS A NONLINEAR JUNCTION, GO THROUGH THIS

PROCESS, WILL HAVE A SQUARE LAW RESPONSE AND WILL

AFFECT, WILL PUT IN A SIGNAL TO THE CIRCUITRY.

IF THE MODULATION ON THE RF IS IN THE AUDIO FREQUENCY

YOU'LL HEAR IT.

IF IT'S NOT, THERE CAN STILL BE EFFECTS, AND I'LL USE

AN ILLUSTRATION FROM ANOTHER FIELD ENTIRELY, WE HAD A

PHONE SOME YEARS AGO WHERE IN THE PRESENCE OF RF THE

PHONE WOULD TURN ITSELF OFF AND THE REASON WAS THE RF

WOULD GET IN AND PUT ACTUALLY A CONSTANT LEVEL ON A

REFERENCE LINE AND THE POWER SUPPLY THOUGHT THE POWER

TO THE PHONE WAS TOO HIGH AND IT WOULD TURN ITSELF

OFF, WHICH WAS VERY ANNOYING FOR THE USER.

AS HAS BEEN MENTIONED SEVERAL TIMES HERE, THERE CAN BE

NONAUDIO EFFECTS AND WHATEVER WE DO IN TESTING WE HAVE

TO ALLOW FOR THOSE AND REALIZE THAT THAT IS A POSSIBLE

EFFECT.

SO, WHILE WE'RE VERY FOCUSED ON THE BUZZ, WE NEED TO

MAKE SURE THAT WE DON'T DISCOUNT THE POSSIBILITY OF

OTHER EFFECTS COMING IN, AND THOSE, PERHAPS, ARE MOST

PROMINENTLY POSSIBLE WITH SIGNALS THAT DON'T MODULATE

IN THE SAME WAY AS THE TDMA OR THE PCS 1900 SIGNAL.

NOW, AS HAS BEEN SAID, THE PROBLEM IS PARTICULARLY

SIGNIFICANT WITH THE INCREASED USE OF DIGITAL

MODULATION SCHEMES AND THOSE ARE, OBVIOUSLY N. OUR

FUTURE.

THEY'RE HERE TODAY AND GROWING IN POPULARITY.

AND SO, THE FIRST QUESTION THAT THE STANDARD HAD TO

CAPTURE AND ADDRESS IS WHAT IS REQUIRED FOR A HEARING

AID AND PHONE TO WORK ACCEPTABLY AND IF I CAN HAVE MY

FIRST OVERHEAD, I WANT TO START ON A HIGH LEVEL.

SIMPLY, IN A HEARING AID WITH WE HAVE TWO MODES,

ACOUSTIC AND T-COIL.

IN THE BROADEST TERMS, ALL WE NEED IS TWO THINGS IN

ORDER FOR THESE TWO PIECES OF EQUIPMENT TO OPERATE

TOGETHER, ONE IS WE HAVE TO HAVE AN ADEQUATE SIGNAL

LEVEL, AND THE SECOND IS WE HAVE TO HAVE ADEQUATE

SIGNAL QUALITY.

BECAUSE THE RF IS SUCH A SIGNIFICANT ASPECT OF SIGNAL

QUALITY I'M GOING TO BREAK IT OUT AS A {THIRT} FACTOR

BUT LOGICALLY IT'S JUST AN ASPECT OF SIGNAL QUALITY.

AND SO, IN THE ACOUSTIC MICROPHONE MODE OF A HEARING

AID, IF THE PHONE OUTPUTS ADEQUATE VOLUME AT AN

ADEQUATE LOW BACKGROUND NOISE, AND THERE IS A

COORDINATION OF THE EMISSIONS FROM THE PHONE WITH THE

IMMUNITY OF THE HEARING AID, THEY WILL WORK TOGETHER.

IN A SIMILAR FASHION ON THE T-COIL MODE, IF THERE IS

AN ADEQUATE MAGNETIC FIELD AND AN ACCEPTABLE MAGNETIC

SIGNAL-TO-NOISE RATIO, AND THE RF EMISSIONS ARE

COORDINATED, WE'LL HAVE ADEQUATE PERFORMANCE.

AND THAT PRETTY MUCH OUTLINES WHAT THE STANDARD IS

MEASURING.

LET ME GO THROUGH KIND OF WHERE WE ARE AND DEFINE IT

MORE {QLEES}LY.

FIRST OF ALL, ON THE ACOUSTIC MODE, ACOUSTIC VOLUME I

THINK WE CAN ASSUME BECAUSE THE PHONE, AGAIN BECAUSE

IT'S PART OF ITS FUNCTION, PRETTY WELL PROVIDES THAT

AND THAT'S EMBEDDED IN THE OTHER STANDARDS FOR THE

PHONE.

AS FAR AS THE OTHER SOURCES OF NOISE, THOSE WOULD BE

ACOUSTIC SOURCES OF NOISE AND THOSE, AGAIN, ARE

CONTROLLED ELSEWHERE IN THE PHONE STANDARDS.

SO, ON THE ACOUSTIC SIDE WE FOCUS FOR THIS EFFORT ON

THE RF INTERFERENCE COORDINATIONS OF EMISSION AND

IMMUNITY.

ON THE T-COIL SIDE IT HAS BEEN SAID THAT THE MAGNETIC

FIELD IS NEEDED TO BE IN A FIXED RELATIONSHIP TO THE

ACOUSTIC VOLUME, AND THAT'S A PARTICULAR CONCERN

BECAUSE THAT WAS ORIGINALLY BROUGHT ABOUT BASED ON A

PARTICULAR RECEIVER TECHNOLOGY THAT WAS USED IN ANALOG

PHONES, AND NOW WE NEED TO LOOK AT WHAT MAKES SENSE AS

WE HAVE NEW TECHNOLOGIES ON THE HORIZON, OR IN USE

TODAY, AND SOME OF THOSE DO NOT HAVE THE MAGNETIC

LEAKAGE FIELDS THAT SOME OF THE OLD ANALOG RECEIVERS

DID.

SECONDLY, AS ALSO HAS BEEN MENTIONED IN SEVERAL OF THE

PRESENTATIONS, WHERE IN ANALOG PHONES, THE ONLY THING

IN THE PHONE THAT COULD PRODUCE IN THE RECEIVER I

SHOULD SAY, THAT COULD PRODUCE A MAGNETIC FIELD WAS

THE RECEIVER, AND SO IN F.C.C. PART 68 WE MEASURED THE

MAGNETIC FIELD AND ASSUME THAT'S ALL CONTAINING THE

DESIRED AUDIO SIGNAL.

NOW, WE HAVE A LOT OF ELECTRONICS IN THESE PHONES, AND

SOME OF THOSE ELECTRONICS CAN PRODUCE THEIR OWN NOISE

WHICH ADDS AND DISRUPTS THE INTENDED SIGNALS, THINGS

LIKE KEYBOARD SCANSES, DISPLAY REFRESH, THE SWITCHING

OF THE POWER SUPPLY, BATTERY SURGES AND THAT'S A NEW

CHALLENGE AND ANOTHER AREA OF FOCUS IN THIS EFFORT.

AND THE THIRD ONE IS VERY SIMILAR IN ACOUSTIC MODE,

THE RF COORDINATION.

LET ME GO TO THE NEXT GRAPH.

WE NOW MOVE A LITTLE MORE SPECIFICALLY INTO THE

STANDARD.

HERE THE FIRST QUESTION IS WHAT'S THE EMISSION

ENVIRONMENT THAT THE HEARING AID NEEDS TO BE IN.

THE STANDARD IS, RIGHT NOW, DEFINING A FIVE BY FIVE

SENT I METER AREA, STARTING AT THE TOP OF THE PHONE.

IT WILL BE SCANNED AND EVALUATED FOR HEARING AID USE.

THIS SURFACE IS ONE CENT METER OFF THE PHONE WHICH IS

ROUGHLY WHERE THE HEARING AID WOULD BE IN NORMAL USE.

AND, AS YOU WILL SEE LATER, IT'S, THERE IS A

TREMENDOUS FIELD RATING IN THIS AREA.

AND SO, WHAT IS CURRENTLY BEING CONCEIVED IS THAT THE

MIDDLE PORTION OF THAT AREA, YOU SEE IT'S DIVIDED INTO

NINE SUBGRIDS, AND THE MIDDLE ONE AND FIVE ADJACENT

AREAS MUST MEET THE REQUIREMENT, WHICH ALLOWS FOR A

HOT SPOT, WHICH IS VERY COMMON, UP AT THE BASE OF THE

ANTENNA.

AND YOU CAN SEE THE SIDE VIEW THERE WITH THE 1

CENTIMETER FROM THE SURFACE.

THIS IS, NOW STARTING AT THE BACK OF THE STANDARD AND

MOVING FORWARD AND WHAT WE HAVE HERE IS REALLY WHAT I

THINK THE AUD {YOL} GIST AND YOUR CLIENT NEEDS TO

UNDERSTAND.

THAT IS THE STANDARD IS LOOKING AT THE TWO {PEETION}S

OF EQUIPMENT AS A SYSTEM, AND MATCHING THEM ALLOWING

THE USER TO DEFINE WHAT THEY NEED.

SO THREE CATEGORIES OF PERFORMANCE HAVE BEEN

IDENTIFIED, USABLE AT THE LOWEST LEVEL, WHICH AS YOU

CAN SEE IS A 10 TO 16 DB RANGE, THINK THINK THAT'S

CONFIRMED BY THE RESEARCH.

THAT WILL BE CONFINED TO A SINGLE NUMBER BY THE

RESEARCHERS AND THAT COLUMN WILL CHANGE, AS WILL THE

ONES BELOW IT.

THAT RESEARCH AND DATA ANALYSIS IS STILL GOING ON.

USABLE MEANS, YOU CAN MAKE A CALL, BUT YOU WOULDN'T

WANT TO USE THIS LEVEL OF PERFORMANCE PROBABLY FOR

NORMAL USE.

SO, IF ALL YOU NEED IS A PHONE AAVAILABLE FOR

EMERGENCY USE, SOMETHING LIKE THAT, YOU CAN START AT

THE LOWEST LEVEL, THE U 1 LEVEL, AND A HEARING AID

WHICH MEETS THIS CRITERIA, MATCHED WITH THE PHONE THAT

MEETS THE US LEVEL WILL GIVE THAT SIGNAL TO

INTERFERENCE LEVEL.

HOWEVER, IF YOU WANT A PHONE OR HEARING AID SYSTEM

THAT WOULD BE JUDGED AS GOOD FOR NORMAL EVERYDAY USE,

YOU WANT ABOUT 10 DB BETTER PERFORMANCE IN THE 20 TO

26 DB RANGE, AND HERE, THE CONSUMER HAS A CHOICE.

YOU CAN EITHER HAVE A HEARING AID AT THE U1 CATEGORY

AND A PHONE AT THE U2 CATEGORY OR AT OPPOSITE, SO THIS

ALLOWS FOR SOME CONSUMER CHOICE AND IT ALSO A{LOUZ}S,

DEPENDING ON WHICH ONE WAS PURCHASED FIRST AND WHICH

ONE WAS MORE AVAILABLE.

ALSO, I SHOULD SAY ON BOTH SIDES IT ALLOWS FOR

IMPLEMENTATION INTO TECHNOLOGIES ON A LOGICAL SCALE.

BOTH HEARING AIDS AND TELEPHONES HAVE A NUMBER OF

DIFFERENT TECHNOLOGIESIES INVOLVED.

SOME ARE GOING TO BE MORE AMENABLE TO SOLUTIONS THAN

OTHERS.

BY HAVING A CATEGORIZED SYSTEMS TECHNOLOGY CAN MOVE

FORWARD AS FAST AS IT CAN AND THE MARKET CAN BRING IN

THE FORCES THAT DRIVE THE IMPLEMENTATION AND SELECT

WHICH ONE IS PREFERABLE.

TO GET UP TO 10 DB MORE YOU GET A SYSTEM THAT CAN BE

DESCRIBED AS EXCELLENT PERFORMANCE, AND HERE YOU HAVE

ALMOST NO OR PERHAPS NO DISCERNIBLE INTERFERENCE, AND

HERE, BY EITHER GETTING A PIECE OF EQUIPMENT AT THE

{T} U3 CATEGORY OR BY GETTING TWO PIECES OF EQUIPMENT

AT THE U2 CATEGORY YOU GET THAT LEX OF PERFORMANCE.

I SHOULD SAY SOMETHING AT THIS POINT ABOUT U1 AND THAT

EVEN THOUGH IT'S THE LOW {EFL} LEVEL IS NOT A TRIVIAL

LEVEL TO ACCOMPLISH.

THAT IS A STRINGENT REQUIREMENT, AND ONE THAT CANNOT

BE MET BY MUCH EQUIPMENT THAT HAS BEEN PRODUCED BEFORE

THIS.

HOWEVER, IT IS THE BEGINNING POINT AND ALLOWS THE

BUILD-UP OF THIS SYSTEM.

THE CONCEPT OF THIS SYSTEM IS TO ALLOW FOR MAXIMUM

FLEXIBILITY, FOR A SYSTEM THAT MAKES SENSE ACROSS A

BROAD SPECTRUM OF TECHNOLOGY, BUT IT WILL REQUIRE

EDUCATION OF AUDIOLOGISTS AND THEN TO HEARING AID USER

TO UNDERSTAND WHAT THEY'RE GETTING.

I'D LIKE NOW TO TURN TO THE LOGIC OF THE STANDARD FOR

A FEW MOMENTS, AND BASICALLY THE STANDARD STARTS WITH

AN ASSUMPTION OF NORMAL SPEECH FROM THE PHONE,

TYPICALLY THAT'S JUDGES TO BE A 65 D BS PL REFERENCE

TO A 20 MICRO PASS CALL LEVEL.

AND THEN THAT THE HEARING AID WILL DELIVER THAT TO THE

USER WITH THE APPROPRIATE GAIN AS NECESSARY TO THE

USER'S PRESCRIPTION.

THAT, IN ORDER FOR THAT SIGNAL TO BE ADEQUATELY

INTELLIGIBLE THE USER FOR NORMAL USE NEEDS A 20 TO 26

DB SIGNAL-TO-NOISE RATIO, AND THAT WE NEED THIS

COORDINATION OF SOUND -- OF IMMUNITY WITH A MISSION.

IN T-COIL MODE THE LOGIC IS VERY SIMPLY THAT THERE

NEEDS TO BE A SIGNAL THAT'S IN A KNOWN RELATIONSHIP TO

THE ACOUSTIC SIGNAL SO THAT THE HEARING AID CAN SET

ITS AMPLIFICATION PATH TO MATCH THAT OF THE PHONE IN

ACOUSTIC MODE.

EXCUSE ME, TO MATCH THE HEARING AID'S PATH IN ACOUSTIC

MODE.

SECONDLY, THAT THE SIGNAL QUALITY, AGAIN, IS ADEQUATE

SO THAT THERE IS GOOD INTELLIGIBILITY.

THERE IS AN ADDED COMPLICATION HERE, AND THAT IS NOW

YOU HAVE TWO POTENTIAL NOISE SOURCES AND THEY'RE

INDEPENDENT.

YOU HAVE THE BASE BAND NOISE SOURCE, AND YOU HAVE THE

RF POTENTIAL.

WHAT IS DONE IS WE DROP THE OR ACTUALLY INCREASE, DROP

IN TERMS OF DB, INCREASE IN TERMS OF SEVERITY, THE

REQUIREMENT ON THE BASEBAND SIGNAL TO INTERFERENCE,

AND ASSUME THAT THE RF IS GOING TO BE AT THE SAME

LEVEL AS IT WAS FOR THE ACOUSTIC SO THAT THE COMBINED

TWO NOISE SOURCES RESULT IN TOTAL {DLIFD} SIGNAL AT

THE ADEQUATE QUALITY -- DELIVERED SIGNAL AT THE

ADEQUATE QUALITY LEVEL.

THERE IS AT THE END OF THIS PROCESSION SEVERAL ISSUES

THAT REMAINED IN THE USEDDER'S CONTROL.

AS I MENTIONED, THERE ARE HOT SPOTS AND I'LL SHOW DATA

THAT SHOWS A TREMENDOUS AMOUNT OF VARIABILITY AND A

USER CAN HELP THEIR SITUATION BY MOVING THE PHONE IN

SUCH A WAY THAT THE ANTENNA IS AS FAR AS POSSIBLE FROM

THE HEARING AID.

ONE POINT, PARTICULARLY MIGHT BE, THAT ON HEARING AIDS

WHERE THE {MOK} {RO} PHONE IS NOT CENT HE REMEMBERED

IN THE EAR, REALLY WHAT YOU WANT IS THE PEAK ACOUSTIC

OUTPUT FROM THE RECEIVER CENT HE REMEMBERED OVER THE

MICROPHONE, NOT NECESSARILY OVER THE USER'S EAR WHICH

FOR A DPE MIGHT BE SLIGHTLY HIGH OR BEHIND THE EAR.

IF WE GO TO THE NEXT SLIDE WE CAN LOOK AT THAT.

THIS IS A GENERIC MODEL PHONE AND WE'LL LOOK AT THIS

FOR A MOMENT.

THIS IS A METHOD OF MOMENTS MODEL THAT WE DID ON A

COMPUTER AND IT SHOWS AN ANTENNA WITH A GENERIC SIZE

PHONE, UP AT THE TOP YOU SEE A FEED TO THE ANTENNA,

ABOUT A CENTIMETER DOWN FROM THE BOTTOM.

NOW, IF WE CAN GO TO THE NEXT ONE?

WHAT WE HAVE HERE IS THE RF FIELD GRADIENT OUT OF THAT

METHOD OF MOMENTS MODEL AND THIS IS A GENERIC

COMPUTATION, BUT IT MATCHES WELL WITH DATA THAT WE

HAVE TAKEN AS WELL AS OTHER COMPANIES.

YOU CAN SEE THE PEAK THERE IN THE PINKISH-RED, THAT IS

RIGHT OVER THE FEED OF THE ANTENNA.

AND NOTICE THAT WITHIN JUST A VERY FEW CENTIMETERS

THERE IS A GRADIENT, EVERY CHANGE OF COLOR IS A 10

PERCENT REDUCTION IN THE FIELD STRENGTH.

THE NEXT ONE WILL BE A TOP VIEW OF THAT.

HERE YOU CAN SEE THE PEAK FIELD AND YOU CAN SEE HOW IN

THE CENTER OF THE RECEIVER AREA, IF THE RECEIVER IS

THE TOP OF THE PHONE WHERE IT {FRL}LY IS, YOU'RE

ALREADY DOWN A GOOD 30-40 PERCENT FROM YOUR HOT SPOT

THERE.

BUT, IF YOU WERE TO HOLD THE PHONE JUST -- JUST A

CENTIMERE OR SO AWAY YOU COULD GET ANOTHER 20 TO 30

PERCENT REDUCTION.

THAT'S {MORN} IMPORTANT.

NOTICE THAT THE MAGNETIC FIELD IS UP SLIGHTLY HIGHER

BUT AGAIN VERY STEEP GRADIENT AND THIS IS TYPICAL OF

THE NEAR FIELD EFFECTS.

THESE NOW HAVE TO BE COORDINATED WITH THE DESIRED

SIGNAL, WHICH IS A FARFIELD HOMO GENEOUS RADIATED

PATTERN.

HERE YOU SEE, IF ANYTHING, THE GRADIENT IS EVEN

STEEPER IN THE MAGNETIC FIELD, AND I SHOULD SAY THAT

WITHIN HEARING AIDS THE RECEPTION MECHANISM CAN BE

MORE SENSITIVE TO THE E FIELD OR THE H FIELD.

THERE ARE DIFFERENT KINDS OF SURFACE AND DIFFERENT

KINDS OF HEARING AID TECHNOLOGIESIES, AND I KNOW FROM

TALKING TO SOME OF OUR FOLKS ON THE COMMITTEE AND FROM

THE HEARING AID SIDE, THAT SOME HEARING AIDS ARE MORE

SENSETIVE TO THE E FEEL AND SOME TO THE H.

SO THIS GRADIENT ISSUE, IF ANYTHING IS MORE PRONOUNCED

HERE.

AGAIN USER EDUCATION OF MOVING IT SLIGHTLY CAN DO

ALLOT TO HELP THE SITUATION.

IT'S THIS ISSUE THAT IS FOCUSED ON IDENTIFYING THE

EXACT TEST AREA, AND THE IMPORTANCE OF THAT.

SO THEN BACK TO THE STANDARD, WHAT DOES THE STANDARD

HAVE TO DO IN ORDER TO ISOLATE THE COMPONENTS OF THIS?

FIRST OF ALL, IT HAS TO MEASURE THE E AND H FIELDS OUT

OF THE PHONE.

AND BECAUSE WE'RE SO CLOSE THERE IS NO FIXED

RELATIONSHIP BETWEEN THESE TWO COMPONENTS SO THE

STANDARD HAS TO MEASURE BOTH.

LIKEWISE, IT HAS TO MEASURE FOR THE FIRST TIME, AS FAR

AS I'M AWARE OF, THE NEAR FIELD IMMUNITY OF THE

HEARING AID.

I DON'T KNOW OF ANOTHER NEAR FILED IMMUNITY TEST

STANDARD BUT THESE TWO PIECES OF EQUIPMENT ARE SO

CLOSE THAT WE HAVE TO EVALUATE BOTH AND SO IN THE

STANDARD THERE IS A RESIDENT DIPOLE, YOU MOVE THE

HEARING AID TO A HIGH H AND HIGH E AREA AND EVALUATE

THE IMMUNITY TO BOTH REALIZING I DIDN'T SAY GOING TO

BE EXPOSED TO BOTH.

THEN, WE HAVE TO LOOK AT THE BASEBAND H FIELD OF THE

PHONE AS FAR AS THE SIGNAL LEVEL, AND FINALLY FOR

BOTH, THE ACOUSTIC AND T COIL EQUIPPED HEARING AID, WE

NEED TO LOOK AT THE SIGNAL TO INTERFERENCE RATIO OF

THAT SIGNAL, SO THAT FOR THE T-COIL MODE WE HAVE

ADEQUATE SIGNAL LEVEL AND ADEQUATE SIGNAL TO

INTERFERENCE LEVEL.

AND THEN THE HEARING AID HAS TO TAKE AND DELIVER THAT

TO THE USER, BUT THAT IS ALREADY ENCOMPASSED IN

ANOTHER STANDARD AND SO BY REFERENCE, THAT PIECE OF

WORK IS DONE.

WE HAVE TALKED ABOUT PHONE EMISSIONS TESTING.

THERE IS SOME PARTICULAR ISSUES THAT WE HAVE TO BE

SENSITIVE TWO.

THE INTERFERENCE IS CENT HE REMEMBERED ON THE PEAK

FIELD LEVELS, AND MOST WORK THAT'S BEEN DONE TO THIS

DATE HAS BEEN MORE INTERESTED IN THE AVERAGE POWER

DELIVERED.

SO, WHATEVER MEASUREMENTS ARE DONE YOU HAVE TO BE

CAREFUL THAT THE PROBING SYSTEM USED IS ACTUALLY

MEASURING THE PEAK OF THE FIELD AND NOT JUST

AVERAGING.

ON THE IMMUNITY SIDE, LETS GO TO THE NEXT ONE AND I'LL

SHOW THAT DIPOLE.

WE HAVE A RESIDENT DIPOLE THAT WE'RE MOVING THE

HEARING AID THROUGH, HERE YOU CAN'T SEE IT, THIS IS

THE E FIELD, THE CENTER IS WHERE THE NULL IS, YOU CAN

SEE THE PACK OVER ONE TIP, AND THEN THE BEGINNING OF

THE PEAK GOING TO THE OTHER TIP, WHICH SENDS OUT OUT

OF THE GRAPH.

IT IS A VERY SHARP FIELD GRADIENT.

THIS IS ONE OF THE HARD BALANCES THAT THE STANDARD IS

HAVING TO MAKE.

THAT FIELD GRADIENT ACTUALLY CHANGES BY MORE THAN 10

PERCENT WITHIN A MILLIMETER IN ONE DIMENSION SO THAT

POSITIONING IS ACTUALLY CRUCIAL IN ORDER TO GET A

REPEATABLE TEST OF A HEARING AID.

IF YOU MOVE TO AN AIR WHERE THE FIELD DOESN'T CHANGE

RAPIDLY YOU START GETTING AWAY FROM THIS HIGHLY POLL

ARRESTIZED FIELD WHERE YOU HAVE HIGH DIFFERENTIATION

BETWEEN THE E AND H FIELD SO YOU HAVE THIS DIFFERENT

IF YOU WANT A HIGHLY DIFFERENTIATED FIELD YOU'RE IN AN

AREA WHERE SMALL POSITIONING DIFFERENCES MAKE A

TREMENDOUS AMOUNT OF DIFFERENCE IN EVERY DIMENSION.

I SAID IN ONE DIMENSION YOU CAN BE OUT OF THE 10

PERCENT RANGE WITHIN A MILLIMETER.

SO, IN THIS EVALUATION POSITION IS GOING TO BE

CRITICAL AND WE'RE GOING TO BE DOING SOME CAREFUL

EVALUATION OF THE REPEATABILITY OF THE TEST.

THE H FIELD IS SIMILAR EXCEPT THE PEAK MOVES INTO THE

CENTER WHERE YOU SEE THE TROUGH NOW, AND OUT AT THE

TIPS YOU GET A LOW H FIELD.

THE HEARING AID THEN IS MOVED ABOUT ONE CENTIMETER

FROM THE DIMOLE AND THEN OUT TO THE PEAK REGION AND

EVALUATED.

ANOTHER IMPORTANT ISSUE IS THE OPERATING STATE OF THE

HEARING AID THROUGH THIS, AND {UB}' HEARD SOME

DISCUSSION OF THAT IN THE REPORTS ON THE RESEARCH.

THE STANDARD HAS AN INPUT TO THE HEARING AID SO THAT

IT HAS A SOUND THAT IT'S RECEIVING AND PUTS IT A

NORMAL OPERATING STATE, AND THE AUDIO FREQUENCY

EMISSION IS THEN LOOKED AT TO BE WITHIN A RANGE THAT

WOULD BE IDENTICAL TO WHAT WOULD NORMALLY BE EXPECTED

IN A NORMAL PHONE SYSTEM.

NOW, THERE IS AN INTERESTING ISSUE HERE THAT WE'RE

LOOKING AT IN PARTS 68 TODAY FOR LANDLINE HEARING AID

COMPATIBILITY, THE WAY THAT YOU ESTABLISH THE AUDIO

LEVEL IS YOU PUT A VOLTAGE ON TIP AND RING AND YOU

ASSUME THERE IS A FIXED RELATIONSHIP AT WHAT IS {SEN}

SHAMLY THE OFFICE EQUIPMENT COMING OUT TO THE PHONE

AND THE SOUND COMING OUT OF THE PHONE, AND THAT'S

CURRENTLY IN THE TEST METHOD STANDARD 504.

HOWEVER, ETSI HAS RECENTLY FAVORED A DIFFERENT METHOD,

RATHER THAN ASSUMING A FIXED RELATIONSHIP FROM THE

VOLTAGE TO THE WIRESES AND THEREFORE THE MAGNETIC

FIELD OUT, ETSI PREFERS TO MEASURE THE LEVEL DIRECTLY

AND RELATE THAT TO THE MAGNETIC FIELD, AND THAT'S

WHERE WE'RE CENT HE REMEMBERED.

ALSO THE NEW REVISION OF 504 HAS THE DIRECT

MEASUREMENT OF THE ACOUSTIC OUTPUT OF THE PHONE AND

COMPARES IT DIRECTLY TO THE MAGNETIC FIELD OUTPUT.

TURNING TO SIGNAL TO INTERFERENCE, AS I'VE SAID THAT

IS A NEW DIMENSION TO THE LANDLINE COMPATIBILITY.

AS FAR AS THE HUMAN SIMULATION, ALL THE TESTS IN THE

STANDARD ARE LOOKING AT IN AIR TESTS, AND WE ARE

LOOKING TO THE RESEARCH TO GIVE US ANY KIND OF

ADJUSTMENT FACTORS WE NEED.

AND I SHOULD SAY A FINAL CONSIDERATION IN COMPLEXITY

IS TEST SIGNALS.

WITH PHONES WITH VOCODERS, YOU CANNOT PASS A NORMAL

TYPE OF SIGNAL TONE SIGNALS, BECAUSE THOSE TYPES OF

CIRCUITRY IS MEANT TO IDENTIFY NOISE AND NOISE

CANCELLATION IN ORDER TO SAVE ON TRANSMISSION

BANDWIDTH.

AND SO, IF YOU PASS A SINGLE TONE, SIP {CHRI} IT WILL

NOT BE PASSED OR IT WILL BE PASSED IN A WAY THAT THE

PHONE IS NOT IN THE SAME STATE IT WOULD BE FOR A VOICE

SIGNAL AT A COMFORTABLE LEVEL.

HEARING AIDS, INCREASINGLY, ARE GOING TO COMPLEX

SIGNAL PROCESSINGAL GO RHYTHMS AND ALSO HAVE SIMILAR

KINDS OF CHARACTERISTICS, SO, AS WE TEST THESE THINGS

IT'S IMPORTANT THAT THE PERSON TESTING UNDERSTAND

EXACTLY WHAT THE PROCESS IS GOING ON, BOTH IN THE

PHONE VOCODER AND IN THE HEARING AID, ESPECIALLY FOR

THE NEWER TECHNOLOGY HEARING AIDS AND PASS ADEQUATE

SIGNAL SO THAT THE UNIT IS NOT IN AN UNUSUAL STATE OF

OPERATION.

AND FINALLY, I WOULD JUST LIKE TO SAY, AS FAR AS THE

STANDARD, AS THE STANDARD'S BEEN GOING ALONG AND IT'S

BEEN FOCUSED PRIMARILY ON MEASUREMENTS, THERE HAS BEEN

A LOT OF DISCUSSION ABOUT SOLUTIONS, A LOT OF TRADING

OF IDEAS AND TECHNOLOGY, AND I WOULD SAY I'MITE

GRATIFIED TO SAY REALLY A SWELL OF EMERGING SOLUTIONS

{Z}.

WE CAN SEE ALREADY ON THE HEARING AID SIDE

MANUFACTURERS COMING OUT WITH INCORPORATION OF BY {KS}

PASS CAPACITORS WITHIN THE COMPONENTS AND YOU'VE HEARD

REPORTS ON GOOD RESPONSE FROM SOME OF THOSE.

WE SEE, I THINK ALMOST ALL OF PHONE MANUFACTURERS HAVE

OR HAVE IN PROCESS PAMPHLETS ADDRESSING THE ISSUE,

TALKING ABOUT THE DIFFERENT KINDS OF SOLUTIONS THEY

HAVE AVAILABLE TODAY, AND WE GET ALL KINDS OF

INDICATIONS OF OTHER WORK THAT'S GOING ON VERY

AGGRESSIVELY.

AND, FINALLY, AND PERHAPS FOR THIS AUDIENCE MOST

IMPORTANTLY, A NUMBER OF ITEMS HAVE BEEN IDENTIFIED

WITHIN THE USER CONTROL THAT CAN MATERIALLY HELP THEIR

SITUATION.

AS I'VE SAID, POSITIONING IS VITAL, ORIENTATION IS

VITAL.

IF YOU, A FEW DEGREES SHIFT CAN MAKE A BIG DIFFERENCE

IN THE AMOUNT OF INTERFERENCE JUST BECAUSE YOU

CROSS-POLARIZE CIRCUITS.

THERE ARE ALSO MATCHING ISSUES THAT ARE THERE TODAY

AND WILL BE THERE IN THE FUTURE AS WE CATEGORIZE THE

EQUIPMENT AND GIVE THE USER THE INFORMATION THEY NEED

TO MAKE INFORMED CHOICES.

AND SO, WITH THAT, I'LL CONCLUDE AND THANK YOU.

[APPLAUSE]

.

Q. A COMMENT ABOUT POSITIONING.

I'VE BEEN WEARING HEARING AID AND T-COIL FOR MANY

YEARS AND WHEN YOU GET THAT GOOD POINT IN A 15-20

MINUTE CONVERSATION YOU'RE MUSCLES ARE GOING TO GET

THAT RIGID.

IT'S NOT THAT EASY.

I WOULD LIKE TO SEE MYSELF A SOLUTION.

A. I THINK IN THE WHY WOULD WE HAVE OMNI-DIRECTIONAL

RECEIVERS IN THERE WHICH WE DON'T HAVE TODAY, AND

THAT'S A GOOD POINT.

Q. WHEN YOU TALK ABOUT THE VARIOUS U CATEGORIES FOR

PHONES AND HEARING AIDS, ARE YOU TALKING ABOUT SOME

TIME IN THE FUTURE SPECIFICATION SHEETS FOR HEARING

AIDS LISTING A U FACTOR FOR THE HEARING AID AND THEN

THE PACKAGING FOR THE CELLPHONE LISTING SOME KIND OF A

U FACTOR FOR THE CELLPHONE?

A. ANSI STANDARDS ARE CONSENSUS STANDARDS, AND WHAT

THEY DO IS THEY ESTABLISH WHERE THE CONSENSUS IS FOR

THE MEASUREMENT.

THEY DON'T HAVE ANY REGULATORY WEIGHT TO THEM AND,

HOWEVER, THERE IS A NOTICE OF PROPOSED RULE-MAKING

CURRENTLY OUT FROM THE ACCESS BOARD, AND THEY STATE

THEIR INTENT THAT WHEN THE STANDARD IS FINISHED,

THEY'RE GOING TO INCLUDE IT IN THEIR GUIDELINES.

AND SO, ANY KIND OF REPORTING REQUIREMENT WOULD COME

OUT OF SOMETHING LIKE THAT.

HOWEVER, I WOULD IMAGINE THAT MANUFACTURERS WOULD WANT

THEIR CUSTOMERS TO KNOW THIS INFORMATION, AND SO, I

WOULD EXPECT IT WOULD BE MADE AVAILABLE, BUT I

COULDN'T SPECK YOU LATE AS TO THE INVOLUNTARY SYSTEM

HOW THEY CHOOSE TO DO THAT OF.

SOMETHING I DIDN'T PUT UP FOR TIME BUT I WANT TO

CLARIFY.

THE U CATEGORIES ARE THERE FOR ACOUSTIC MODE, THAT IS

THE PRODUCTS MEET THE REQUIREMENT IN ACOUSTIC.

IF THEY ALSO MEET THE COMPLIMENTARY REQUIREMENTS IN T

COIL CODE.

SO, AND THAT WOULD BE IMPORTANT FOR HEARING AID USER

WITH T-COILS.