## Sheet1
EMI tests of Starkey hearing aids done at Environ Labs in Nov. 2008 using dipole antennas
This data is taken from Environ Engineering Report 39180-2, Rev 1 of 12-18-08.
Tests were done per ANSI C63.19-2007 using the primary method (dipole antennas).
Low frequency band = 800 - 950 MHz. High frequency band = 1600 - 2500 MHz.
|  |  | Field Strength |  | Field Strength |  | Equalize | Compare
HA tested | Freq. Band | RF field | for 55 dB IRIL |  | for 55 dB IRIL |  | H to E field | H - E field | M-rating
BTE1 | Low | E | 1587 | V/m | 64.0115 | dB V/m
" | " | H | 4.401 | A/m | 12.871 | dB A/m | 65.871 | 1.8595
" | High | E | 169 | V/m | 44.5577 | dB V/m |  |  | 3 - 4
" | " | H | 0.759 | A/m | -2.3952 | dB A/m | 50.6048 | 6.0471 | 4
BTE2 | Low | E | 1327 | V/m | 62.4574 | dB V/m
" | " | H | 3.715 | A/m | 11.3992 | dB A/m | 64.3992 | 1.9418
" | High | E | 872 | V/m | 58.8103 | dB V/m |  |  | 4
" | " | H | 2.345 | A/m | 7.4029 | dB A/m | 60.4029 | 1.5925 | 4
BTE3 | Low | E | 1406 | V/m | 62.9597 | dB V/m
" | " | H | 3.128 | A/m | 9.9053 | dB A/m | 62.9053 | -0.054372 | 4
" | High | E | 1054 | V/m | 60.4568 | dB V/m |  |  | 4
" | " | H | 3.449 | A/m | 10.7539 | dB A/m | 63.7539 | 3.2971
BTE4 | Low | E | 1513 | V/m | 63.5968 | dB V/m
" | " | H | 3.842 | A/m | 11.6911 | dB A/m | 64.6911 | 1.0944
" | High | E | 281 | V/m | 48.9741 | dB V/m |  |  | 4
" | " | H | 0.672 | A/m | -3.4526 | dB A/m | 49.5474 | 0.57326 | 4
BTE5 | Low | E | 374 | V/m | 51.4574 | dB V/m
" | " | H | 0.893 | A/m | -0.98297 | dB A/m | 52.017 | 0.5596
" | High | E | 178 | V/m | 45.0084 | dB V/m |  |  | 3 - 4
" | " | H | 0.677 | A/m | -3.3882 | dB A/m | 49.6118 | 4.6034 | 4
BTE6 | Low | E | 18.3 | V/m | 25.249 | dB V/m
" | " | H | 0.057 | A/m | -24.8825 | dB A/m | 28.1175 | 2.8685
" | High | E | 9.3 | V/m | 19.3697 | dB V/m |  |  | 1
" | " | H | 0.024 | A/m | -32.3958 | dB A/m | 20.6042 | 1.2346 | 1
ITE7 | Low | E | 1342 | V/m | 62.5551 | dB V/m
" | " | H | 3.108 | A/m | 9.8496 | dB A/m | 62.8496 | 0.29457
" | High | E | 463 | V/m | 53.3116 | dB V/m |  |  | 4
" | " | H | 1.335 | A/m | 2.5096 | dB A/m | 55.5096 | 2.198 | 4
ITE8 | Low | E | 1540 | V/m | 63.7504 | dB V/m
" | " | H | 3.589 | A/m | 11.0995 | dB A/m | 64.0995 | 0.34905
" | High | E | 351 | V/m | 50.9061 | dB V/m |  |  | 4
" | " | H | 1.675 | A/m | 4.4803 | dB A/m | 57.4803 | 6.5742 | 4
ITE9 | Low | E | 1623 | V/m | 64.2064 | dB V/m
" | " | H | 2.276 | A/m | 7.1434 | dB A/m | 60.1434 | -4.0629
" | High | E | 142 | V/m | 43.0458 | dB V/m |  |  | 3
" | " | H | 1.489 | A/m | 3.4579 | dB A/m | 56.4579 | 13.4121 | 4
[8 empty] | -4.0629 | Min
[8 empty] | 13.4121 | Max
[8 empty] | 2.4657 | Avg