Submission to the ANSI C63.19 Working Group

Comparison of Emissions and Immunity Data

Stephen Julstrom – January 21, 2010

As noted in a previous submission, the comparison of wireless device (WD) emissions measurements to hearing aid (HA) immunity measurements is complicated by the fact that the ratio of the emissions requirements (E/H field) is 50.4 dB (331 Ohms) while the ratio of the immunity requirements is 53.0 dB (447 Ohms).  Neither of these happens to be equal to the free-space (far-field) value of 51.5 dB (337 Ohms), the value theoretically present in a GTEM field.  In deciding when the E or H-field is dominant in a given WD-HA combination, these differences need to be considered.

From Tom Victorian’s recently submitted immunity measurements of nine BTE and ITE aids, the following histogram has been created to show the distribution of the resultant E/H field ratios, expressed as dB ratios, with the corresponding emissions histogram from a previous report presented immediately following for comparison.

Both graphs show a generally higher E/H ratio (immunity and emissions) for the low band in comparison to the high band.  If the E-field were to be the determining factor over the H-field in any given WD-HA combination, the WD E/H emissions ratio should be higher than the HA E/H immunity ratio.  A comparison of the two histograms shows that this is usually the case in each band, but not always.

A more specific tabulation of the percentage of WD-HA combinations where the H-field may be predicted to cause more interference than the E-field is presented next, comparing each of the nine measured hearing aids to the totality of the earlier GSM, CDMA, WCDMA, and iDEN devices measurements (more than 200, not including a couple of outlier measurements,and not including the Wi-Fi and Bluetooth devices).

| 9 Low Band Hearing Aid E/H ratios (dB) | Low Band % of H-dominant combinations | 9 High Band Hearing Aid E/H ratios (dB) | Low Band % of H-dominant combinations |
| --- | --- | --- | --- |
| 50.1 | 6.5% | 39.6 | 0.0% |
| 51.1 | 12.0% | 46.4 | 1.0% |
| 51.1 | 12.0% | 47.0 | 1.4% |
| 51.9 | 14.8% | 48.4 | 6.2% |
| 52.4 | 16.2% | 49.7 | 16.2% |
| 52.7 | 17.1% | 50.8 | 28.1% |
| 52.7 | 17.1% | 51.4 | 45.7% |
| 53.1 | 18.5% | 51.8 | 59.5% |
| 57.1 | 56.5% | 52.4 | 74.3% |
| total: | 19.0% | total: | 25.8% |

The percentage of H-dominant combinations is not insignificant, particularly for some hearing aids with low relative H-field immunity (or, equivalently, high relative E-field immunity).  To further complicate things, a GTEM HA immunity measurement does not specifically separate the E and H-field measurements.  If the HA were rotated in the GTEM to the worst-case orientation, then the effective E/H testing ratio is the free-space ratio of 51.5 dB.  Relative to this ratio, 13.4% of the low band and 50% of the high band combinations could potentially be H-dominant in this E/H-blind measurement.

Ignoring the actual testing E/H ratios for the moment, many of the above tabulated combinations are H-dominant by only a dB or two.  The following table breaks down the percentage of WD/HA combinations that are H-dominant by >0 dB, >1 dB, etc.:

| % of WD/HA combinations H-field dominant by: | GSM, CDMA, WCDMA, iDEN low band | GSM, CDMA, WCDMA high band |
| --- | --- | --- |
| > 0 dB | 19.0% | 25.8% |
| > 1 dB | 14.4% | 13.9% |
| > 2 dB | 10.9% | 7.5% |
| > 3 dB | 7.1% | 4.1% |
| > 4 dB | 4.6% | 2.1% |
| > 5 dB | 3.1% | 0.9% |

From this data, it may be further estimated that 9.3% of the low band and 7.7% of the high band device combinations would likely result in a change of one total M-category summation if H-field measurements were dropped on both WD and HA (in comparison to including both emissions and immunity measurement ratings at matching E/H ratios).

It is evident that both E and H-field M-rating tables are not needed for GTEM testing of hearing aid immunity (and are, in fact, confusing), but other implications regarding H-field testing are not unequivocal. The interference potential of perhaps 1 in 5 WD/HA combinations could be underestimated to at least some degree by E-field only measurements, although perhaps less than 1 in 10 WD/HA combinations by an amount resulting total M-category summation change.

SDJ – 1/21/2010