| Table E.2-Hearing aid immunity measurements Hearing aid near-field immunity measurement uncertainty estimation |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Contribution | Data dB | Data type | Prob. dist. | Weight | Uncertainty dB | Notes/comments |
| RF reflections | ±0.8 | Spec | Rect | 2/√3 | ±0.92 | Reflections < -20 dB |
| Power meter (forward) | ±0.06 | Spec | Rect | 2/√3 | ±0.069 | VSWR ≤ 1.08, Γ ≤ 0.04 |
| Power meter (reverse) | ±0.06 | Spec | Rect | 2/√3 | ±0.069 | VSWR ≤ 1.08, Γ ≤ 0.04 |
| Directional coupler | ± 1.0 | Spec. | Rect | 2/√3 | ±1.15 | VSWR ≤ 1.15, Γ ≤ 0.07 |
| Cable loss | ± 1.0 | Uncert'y | Norm. | 2/2 | ±1.00 |  |
| Hearing aid loading of ant. | -- | -- | -- | -- |  | VSWR ≤ 1.9, Γ ≤ 0.31 |
| Mismatch | ± 0.19 | Spec. | U-shaped | 2/√2 | ±0.27 | 20Log(1 ± Γ1Γ2) |
| Positioning accuracy | ± 1.62 | Spec. | Rect. | 2√3 | ±1.87 | E.2.3 |
| Acoustic transmisson line | -- | -- | -- | -- |  | TBD |
| Microphone | ± 1.0 | Spec. | Rect. | 1/√3 | ±0.58 |  |
| 2 cc coupler | - | - | - | - |  | TBD |
| Pre-amplifier | ± 1.0 | Spec. | Rect. | 1/√3 | ±0.58 |  |
| Frequency analyzer | ±0.5 | Spec. | Rect. | 1/√3 | ±0.29 |  |
| System repeatability | ±0.5 | Std.Dev. | Norm. | 1 | ±0.50 |  |
| EUT repeatability | -- |  |  | -- | -- | TBD |
| Combined standard uncertainty, uc(y) |  |  | Norm. | 1 | 2.79 |  |
| Expanded uncertainty, U |  |  | Norm. | 2 | 5.57 |  |