TELECON AGENDA

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Monday, October 6, 2008 at 1:00PM CT

Location: 	UL Headquarters

Northbrook, IL

- Call to order and introductions

- Approval of agenda

- Review and approval of past minutes

- Patent Policy, Individual Participation, Open Meetings and Anti-Trust Slides

- Technology Evaluation

- Separate New Technology Test in Clause 4?

- Review of draft and status of items listed on website.

- Review of 700 MHz Task Group Report

- Workplan & schedule
Restructuring Task Groups and Review of Action Items

- Next meeting: TBD

- Adjourn

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Project Status

| 1 | Status - Resolved |
| --- | --- |
|  | To expand the frequency range of the standard to cover 698 MHz to 6 GHz. The extension of the lower frequency range is to include the wireless mobile technologies in 698 MHz band as mandated by recent FCC Report and Order. The extension of the upper frequency range up is to include newer technologies in the bands between 3 - 6 GHz, including the 4.9 GHz band as referenced in the latest FCC Report and Order |
|  | 700 MHz task group has conveyed its recommendations to the WG. Its changes are reflected in the current draft. |
| 1a | Status - TBD |
|  | a) Design/Testing and Simulation of Dipoles for the frequency extensions. |
|  | Some test has been proposed. Some measurements in process. Need additional volunteers to confirm measurements are generic to design, confirmed by multiple labs and document measurement uncertainty and lab-to-lab variability. |
| 1b | Status - Resolved |
|  | b) Extend calibration range of test probes from 698 MHz to 6 GHz. |
|  | Issue referred to WG for IEEE 1309 and NIST Fields Measurement Group. |
| 2a | Status - In progress |
|  | To provide AWF factors for systems operating in the 698 MHz to 6 GHz frequency range. |
|  | WG has agreed to ratain continuity with current AWF and testing methods using slow response probes as direct measurement methods are developed. 2 proposals received for an approach to a generalized AWF factor, one to use A-Weighting and a new weighting function has been proposed by Steve Julstrom. Issue needs further discussion and resolution by the WG. Work is needed to map a direct measurement to the categories in the current standard and bridge direct measurements to slow probe measurements. |
| 2b | Status - In progress |
|  | If possible to develop a generalized treatment based on A-Weighting or some other appropriate weighting to replace the AWF table. |
|  | This issue has been linked with item #2. See comments under that item. |
| 3 | Status - In progress |
|  | Evaluate the requirement in Clause 4, 7 and other appropriate sections for RF H-Field measurements. |
|  | Discussion has tended toward retaining the H-Field measurment but the WG has not made a final resolution on the issue. Propose E-Mail ballot following 10/6/08 WG meeting. |
| 4 | Status - In progress |
|  | To evaluate the correlation of hearing aid immunity test results using the dipole and WB TEM methods and consider the advisability of different field strength requirements for each of the methods to bring the test methods into closer agreement. |
|  | Proposal from HIA received. WG needs to have further discussion and make a final determination. Data is being taken and is anticipated in November. A proposal is anticipated from HIA once the new data is available. |
| 5 | Status - TDB |
|  | Review and update test equipment specifications generally, specifically Annex D.3, D.4 and D.7. |
|  | Item has been discussed. A volunteer is needed to take the action to review and propose wording/editing to the current text. Propose removing D.3, audio signal generator, which is no longer used in a measurement. Remove D.4, bandpass filter, for clause 6 cite standard to specify 1/3 octave filter and A-Weighting filter. Bandpass filter is not used. Remover D.7, Frequency Generator, 600 Ohm frequency generator is not used. |
| 6 | Status - TBD |
|  | To evaluate and possibly improve the specification of the band pass filter referenced in sub-clauses 6.1.1.1 and 2. |
|  | Item has been discussed. A volunteer is needed to take the action to review and propose wording/editing to the current text. Need standard citation for 1/3 Octave Bandpass Filter. |
| 7 | Status - TBD |
|  | Determine and specify the power measurement that is most closely linked to user experience, peak, RMS or other parameter of power. |
|  | Item has been discussed. A volunteer is needed to take the action to review and propose wording/editing to the current text. Addressed in generalized method. |
| 8 | Status - Resolved |
|  | Reconsider the need for 3 T-Coil positions relative to user needs. |
|  | On March 4, 2008 Telecon WG agreed to use 2 positions, described as transverse and perpendicular. Modify draft to reflect decision. |
| 9 | Status - Resolved |
|  | The RF probe separation distance is 1.5 cm and the calibration values in table 4.2 were made at 1.0 cm. These calculations need to be redone to the 1.5 cm distance. |
|  | Values for 1.5 cm have been presented and placed in the current draft. See issue 1a. Some measurements in process. Need additional volunteers to confirm measurements are generic to design, confirmed by multiple labs and document measurement uncertainty and lab-to-lab variability. |
| 10 | Status - TBD |
|  | To combine sub-clauses "4.4.1.2.1 Manual scanning method" and "4.4.1.2.2 Automated scanning method" into a single section requiring a positioning resolution and accuracy. |
| 11 | Status - TBD |
|  | To add a 'no RF test' option for low power transmitters that are almost certain to achieve category 4 due to their low transmission power. Resolved per motion at October 6 WG3 meeting. |
| 12 | Status - In progress |
|  | To review and adopt any editorial improvements proposed, including review and update of references and bibliography. |
|  | Changes are either brought to WG for consideration or reflected in draft when proposed. Changes made as items are identified to WG. |

| † Based upon A-weighted sensitivity to frequencies with -20 dB gain or greater Identification of new technology (any stakeholder can bring a request forward for preliminary evaluation) Available industry / open standards defining modulation Preliminary Technology Evaluation: No regular repeating bursts or modulation components above 100 Hz and below 20 kHz† Carrier freq < 2.7 GHz Transmit pwr < 200 mW Preliminary Device Evaluation: (handsets / prototypes available) E- and H- field acoustic coupling > M4 Baseband H-field inductive coupling > T4 ANSI C63.19 test protocol, sufficient sample (n > 10 models), and AWF = 0 FCC spectrum rules defining carrier frequency / channels Resources to pursue experimental data ANSI ASC C63® decision re: proceed with HAC limits (via official PINS) Preliminary data provided to the FCC and member groups Purpose of preliminary data to provide stakeholders with information to make their own interim Regulatory and Policy decisions: May request handsets be disqualified until ANSI C63.19 limits developed May request handsets left “open” until ANSI C63.19 limits developed May request no need for further ANSI C63.19 evaluation or limits Experimental Evaluation: May include lab testing, additional correlation with other technologies, etc New Limits ANSI C63.19 Revision Regular reports to the FCC and stakeholders re: preliminary data |  |
| --- | --- |

## (D.3)

Figure D.17 – Spectral weighting response

| Frequency (Hz) | 100 | 125 | 160 | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Weighting (dB) | -14.63 | -11.80 | -9.06 | -6.92 | -5.10 | -3.56 | -2.29 | -1.40 | -0.73 | -0.26 | 0 |

| Frequency (Hz) | 1000 | 1250 | 1600 | 2000 | 2500 | 3150 | 4000 | 5000 | 6300 | 8000 | 10000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Spectral Weighting (dB) | 0 | 0.10 | -0.01 | -0.45 | -1.45 | -3.36 | -6.40 | -10.06 | -14.46 | -19.52 | -24.70 |

## Table D.11 – Spectral weighting response at third-octave frequencies