Telecon Minutes

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Friday, February 13, 2008 at 11:30AM EST

Participants

Stephen Berger

Jim Turner

Steve Julstrom

Daoud Attiya

Tom Knipple

Dave Case

Dave Preves

Stephen Liu

David Dzumba

Bill Hurst

Joe Morissey

Steve Coston

Minutes

- The chair called the meeting to order at 10:30 AM CT.   Participants were identified and self-introductions were made.

- The agenda was reviewed.

- The minutes for the Dec. 15, 2008 and January 21, 2009 were approved.

- Draft Revision 1.7 was posted to the committee webpage before the meeting.  A review of the decisions from the December 15, 2008 meeting was conducted and discussion made of further changes to the draft, as follows:

  - A motion was made to change C.4 Calibration of dipoles, to only an E-Field calibration.  The motion passed unanimously.

  - A motion was made to revise the report form in Annex L to reflect new limits, elimination of H-Field testing and frequency break at 960 MHz.  The motion passed unanimously.

  - A motion was made to remove Annex K, Measurement of peak power across multiple airlink technologies.  The motion passed unanimously.

  - A motion was made to change the measurement distance on dipole calibration and hearing aid immunity measurement to 1.5 cm from 1.0 cm.  The motion passed with one dissenting vote.

  - A motion was made to implement 2 test procedures in Clause 5, one for fast probe and one for slow probe.  The motion passed unanimously.

  - A motion was made to separate the probe calibration from MIF, creating a new section to treat the development of MIF.  The motion passed unanimously.

- Telecons were set for:

  - Monday March 9, 2009, 11:30 AM ET

- A discussion was held of the possible implication of the changes made on the T-Coil testing.  No decisions were made in this regard on this call.

The discussion transitioned into a brief discussion of the no-test threshold.

- In-Person meeting: 

Mon. April 20, 2009 at TBD 
IEEE Headquarters
Piscataway, NJ

- Adjourn

TELECON AGENDA

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Friday, February 13, 2009 at 11:30AM EST

Phone Number: 	(218) 936-4700

Meeting ID: 	626642

- Call to order and introductions

- Approval of agenda

- Review and approval of minutes for Dec. 15 meeting & Jan. 21 telecon.

- Review of draft 1.7

  - C.4 Calibration of dipoles, can only an E-Field calibration be used?

  - Report form in Annex L needs to be update to reflect new limits, elimination of H-Field testing and frequency break at 960 MHz.

  - Removal of Annex I, Measurement of peak power across multiple airlink technologies

- Next telecon:

Monday March 9, 2009, 11:30 AM ET
same call information as for Feb. 13 telecon

- In-Person meeting: 

Mon. April 20, 2009 at TBD 
IEEE Headquarters
Piscataway, NJ

- Adjourn

Summary of Action Items

| # | Item | Status |
| --- | --- | --- |
|  | Action Items from January 26, 2009 WG Meeting |  |
| 1 | Action item: Steve Julstrom, Tom Knipple and Stephen Berger will be reviewing the draft and incorporate contribution from Steve Julstrom and Tom Knipple. | February 19 |
| 2 | Action item: Companies and groups that identified they will be performing comparative and verification testing are to mature plans and report plan on the Feb. 12 telecon. | Due March 9 |
| 3 | Action item: |  |

| Conference Summary |  |
| --- | --- |
| Conference Type: | Web-Scheduled Standard |
| Conference Date: | Monday, January 26, 2009 |
| Reservation Start Time: | 10:30 AM Central Std Time |
| Reservation End Time: | 11:55 AM Central Std Time |
| Dial-in Number: | 1-218-936-4700 |
| Access Code: | 626xxx |
| Conference Details |  |
| #; Start Time; End Time; Caller Number; Mins.. 1; 01/26/2009 10:28 AM; 11:44 AM; 954-723-xxxx; 77. 2; 01/26/2009 10:28 AM; 11:44 AM; 76. 3; 01/26/2009 10:29 AM; 11:44 AM; 952-828-xxxx; 76. 4; 01/26/2009 10:29 AM; 11:44 AM; 301-596-xxxx; 76. 5; 01/26/2009 10:29 AM; 11:44 AM; 410-290-xxxx; 75. 6; 01/26/2009 10:30 AM; 11:44 AM; 630-452-xxxx; 75. 7; 01/26/2009 10:31 AM; 11:28 AM; 773-706-xxxx; 57. 8; 01/26/2009 10:32 AM; 11:44 AM; 519-888-xxxx; 73. 9; 01/26/2009 10:32 AM; 11:49 AM; 941-201-xxxx; 77. 10; 01/26/2009 10:33 AM; 11:44 AM; 512-869-xxxx; 72. 11; 01/26/2009 10:35 AM; 11:44 AM; 919-536-xxxx; 70. 12; 01/26/2009 10:36 AM; 11:44 AM; 301-362-xxxx; 69. 13; 01/26/2009 10:41 AM; 10:45 AM; 469-231-xxxx; 5. 14; 01/26/2009 10:45 AM; 11:02 AM; 469-231-xxxx; 17. 15; 01/26/2009 10:56 AM; 11:44 AM; 202-651-xxxx; 49. 16; 01/26/2009 11:28 AM; 11:44 AM; 773-706-xxxx; 17. |
| Total Calls: | 16 |
| Peak number of active lines: | 14 |
| Lines reserved: | 8 |