Telecon Minutes

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Monday, November 30, 2009 at 2:30 PM EST

Participants

Stephen Berger

Steve Julstrom

Tom Knipple

Jim Turner

Dave Preves

Dave Chapman

Daoud Attayi

Minutes

- The chair called the meeting to order.   Participants were identified and self-introductions were made.

- The agenda was approved without change.

- A discussion was held of the logistics and test plan for the round robin testing (see draft test plan on the WG website).  Testing is planned at RIM the 1st 2 weeks in December.

  - The following configurations were proposed for the round robin testing:

    - GSM: Enhanced Full Rate

    - WCDMA: 12.2 kbps AMR Mode

    - CDMA: RC1/SO3 mode

    - iDEN (still needed)

  - The suggested MIF values are:

    - GSM +3.98 dB

    - CDMA Full Rate -19.2dB

    - CDMA 1/8 rate +3.4dB

    - WCDMA -20 dB

    - iDEN (TBD)

- Next telecon meeting: 

Thursday December 17, 2009 at 2:30-4:00 PM ET

- Adjourn

TELECON AGENDA

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Monday, November 30, 2009 at 2:30 AM EST

Phone Number: 	(218) 936-4700

Meeting ID: 	626642

- Call to order and introductions

- Approval of agenda

- Round Robin test plan and logistics

- Next telecon meeting: 

TBD

- Adjourn

| Conference Summary |  |
| --- | --- |
| Conference Type: | Web-Scheduled Standard |
| Conference Date: | Monday, November 30, 2009 |
| Reservation Start Time: | 01:30 PM Central Std Time |
| Reservation End Time: | 02:55 PM Central Std Time |
| Dial-in Number: | 1-218-936-4700 |
| Access Code: | 626xxx |
| Conference Details |  |
| #; Start Time; End Time; Caller Number; Mins.. 1; 11/30/2009 01:27 PM; 02:05 PM; 512-869-xxxx; 39. 2; 11/30/2009 01:30 PM; 02:05 PM; 773-706-xxxx; 36. 3; 11/30/2009 01:31 PM; 02:05 PM; 519-888-xxxx; 35. 4; 11/30/2009 01:31 PM; 02:05 PM; 512-372-xxxx; 34. 5; 11/30/2009 01:32 PM; 02:05 PM; 941-201-xxxx; 34. 6; 11/30/2009 01:33 PM; 02:05 PM; ano-nym-xxxx; 33. 7; 11/30/2009 01:37 PM; 02:05 PM; 309-659-xxxx; 29. 8; 11/30/2009 01:49 PM; 02:05 PM; 410-290-xxxx; 17. |
| Total Calls: | 8 |
| Peak number of active lines: | 8 |
| Lines reserved: | 10 |

Validation Test Plan

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

This test plan is being developed to validate the test method changes being considered for the next revision of the ANSI C63.19 standard.

Overview

This test method intends to achieve the following:

- Validate the clarity of the draft by performing the required tests.

  - The intention is to test phones using the following RF modulations:

    - GSM

    - CDMA

    - UMTS

    - LTE (if available)

  - The measurements will provide target values for MIF and final readings, which will be useful for laboratories to use in validating their own implementation of the test.

  - Validate that the test methods work for both fast and slow probes.

- Validate that the correlation of the new test method and new categories to the current test method and categories.

- Validate the correlation of the new test method to field performance of the phones tested

Tasks

The following tasks are necessary to achieve the objectives of this test plan:

| # | Task | Assignment / Status |
| --- | --- | --- |
| 1 | Identify Test Labs & their capbility - Test equipment (fast and/or slow probe) - Base stations (GSM, CDMA, etc.) | Daoud Attiya (RIM) Tom Knipple (Motorola) |
| 2 | Secure phones to be tested | ?? |
| 3. | Draft base station configuration for each protocol to be tested: - GSM - CDMA - UMTS - LTE (if available) | Justin Chao (PCTest) Dave Chapman (ATT) |
| 4. | Provide MIF test equipment: - Hardware implementation - Software implementation | Steve Julstrom Stephen Berger |
| 5. | Schedule tests, manage shipment of test objects & gathering of test data. | Stephen Berger |
| 6. | Have phones tested: - Using current standard - Using draft revision - Measure phone for MIF on live network in 5 different locations/circumstances | See item 1 |
| 7. | Compile and analyze data and provide it to the WG. | Stephen Berger |

My assumption is the RR can proceed with the straw man RR proposal I sent out, since no one had any comments on the objectives and level of detail.

Here are some other assumptions that I have made –

That testing will begin the first week in December and I had the manufacturers send their phones directly to RIM.

Participating Labs – Slow Probe

RIM

PCTEST

Motorola

Possibly Samsung* or Nokia* using overseas labs

Participating Labs – Fast Probe*

ETS Lingren or TEM

AT&T

Cisco

No formal commitment yet