Telecon Minutes

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Thursday, March 11, 2009 at 9:30 AM EST

Participants

Stephen  Berger

Daoud Attayi

Dave Preves

Jim Turner

Steve Coston

David Case

Fin Bomholt

Minutes

- The chair called the meeting to order.   Participants were identified and self-introductions were made.

- The agenda was reviewed and unanimously approved.

- Jim Turner reported that 3 labs had completed the 1st round-robin and the equipment was going to the 4th lab.  A 5th has volunteered to participate.

A discussion was held of the timing to get data from the round-robin to the WG to support decision making at the April 19 meeting.

  - It was agreed that Steve Julstrom and Jim Turner will prepare the data analysis and inform the committee as data was available for release.

  - The chair is to check with the FCC regarding the timeframe in which they need the revised standard.

- A review of the draft was conducted.

  - A motion was made to remove the H-Field from Clause 5, only requiring E-Field testing of WD’s.  The motion passed with 2 abstentions.

  - A motion was made to make the indirect method the preferred method in clause 5.1.  The motion passed with 2 abstentions.

  - A motion was made to accept the proposed wording for Clause 4 with a low power threshold of 21 dBm.  The motion passed with 2 abstentions.

  - An action item was accepted by Dave Case to provide missing test frequencies for Table B.2.

- Adjourn

TELECON AGENDA

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

Time: 	Thursday, March 11, 2009 at 9:30 AM EST

Phone Number: 	(218) 936-4700

Meeting ID: 	626642

- Call to order and introductions

- Approval of agenda

- Round Robin test status and logistics

- Discussion and decision on application of new weighting function to ABM2 measurement

- General review of draft and project status

- Next telecon meeting: 

March 25, 2010 – 9:30 AM ET
April 1, 2010 – 9:30 AM ET

- Adjourn

Conference Summary Conference Type:; Web-Scheduled Standard. Conference Date:; Thursday, March 11, 2010. Reservation Start Time:; 08:30 AM Central Std Time. Reservation End Time:; 09:55 AM Central Std Time. Dial-in Number:; 1-218-936-4700. Access Code:; 626xxx. Conference Details. #; Start Time; End Time; Caller Number; Mins.. 1; 03/11/2010 08:28 AM; 09:30 AM; 144-245-xxxx; 62. 2; 03/11/2010 08:29 AM; 09:30 AM; 630-452-xxxx; 62. 3; 03/11/2010 08:29 AM; 09:26 AM; 512-869-xxxx; 58. 4; 03/11/2010 08:29 AM; 09:30 AM; 330-896-xxxx; 61. 5; 03/11/2010 08:31 AM; 08:35 AM; 919-274-xxxx; 4. 6; 03/11/2010 08:32 AM; 09:30 AM; 519-888-xxxx; 59. 7; 03/11/2010 08:33 AM; 09:30 AM; 301-362-xxxx; 57. 8; 03/11/2010 08:35 AM; 09:30 AM; 919-274-xxxx; 55. 9; 03/11/2010 08:45 AM; 09:30 AM; 941-201-xxxx; 46. 10; 03/11/2010 09:26 AM; 09:30 AM; 512-466-xxxx; 5. 11; 03/11/2010 09:36 AM; 09:36 AM; 512-372-xxxx; 1.. Total Calls:; 11. Peak number of active lines:; 9. Lines reserved:; 10.

Validation Test Plan

ANSI C63.19

Method of Measurements for Compatibility between

Wireless Communications Devices

and

Hearing Aids

This test plan is being developed to validate the test method changes being considered for the next revision of the ANSI C63.19 standard.

Overview

This test method intends to achieve the following:

- Validate the clarity of the draft by performing the required tests.

  - The intention is to test phones using the following RF modulations:

    - GSM

    - CDMA

    - UMTS

    - LTE (if available)

  - The measurements will provide target values for MIF and final readings, which will be useful for laboratories to use in validating their own implementation of the test.

  - Validate that the test methods work for both fast and slow probes.

- Validate that the correlation of the new test method and new categories to the current test method and categories.

- Validate the correlation of the new test method to field performance of the phones tested

Tasks

The following tasks are necessary to achieve the objectives of this test plan:

| # | Task | Assignment / Status |
| --- | --- | --- |
| 1 | Identify Test Labs & their capbility - Test equipment (fast and/or slow probe) - Base stations (GSM, CDMA, etc.) | Daoud Attiya (RIM) Tom Knipple (Motorola) |
| 2 | Secure phones to be tested | ?? |
| 3. | Draft base station configuration for each protocol to be tested: - GSM - CDMA - UMTS - LTE (if available) | Justin Chao (PCTest) Dave Chapman (ATT) |
| 4. | Provide MIF test equipment: - Hardware implementation - Software implementation | Steve Julstrom Stephen Berger |
| 5. | Schedule tests, manage shipment of test objects & gathering of test data. | Stephen Berger |
| 6. | Have phones tested: - Using current standard - Using draft revision - Measure phone for MIF on live network in 5 different locations/circumstances | See item 1 |
| 7. | Compile and analyze data and provide it to the WG. | Stephen Berger |

My assumption is the RR can proceed with the straw man RR proposal I sent out, since no one had any comments on the objectives and level of detail.

Here are some other assumptions that I have made –

That testing will begin the first week in December and I had the manufacturers send their phones directly to RIM.

Participating Labs – Slow Probe

RIM

PCTEST

Motorola

Possibly Samsung* or Nokia* using overseas labs

Participating Labs – Fast Probe*

ETS Lingren or TEM

AT&T

Cisco

No formal commitment yet